BIDDING DOCUMENTS

FOR SUPPLY, INSTALLATION & COMMISSIONING OF MEDICAL EQUIPMENT ON TURNKEY BASIS FOR GOVT. MEDICAL COLLEGE & HOSPITAL, MAHUA (VAISHALI) AND GOVT. MEDICAL COLLEGE & HOSPITAL, MADHUBANI, BIHAR

Bid Reference No.: BMSIC/2024-25/ME-382

Bihar Medical Services & Infrastructure Corporation Limited

2nd & 3rd Floor, Swasthya Bhawan, Behind IGIMS, Sheikhpura,

Adjacent to State Health Society, Patna- 800014, Bihar

Bihar Medical Services & Infrastructure Corporation Limited

2nd & 3rd Floor, Swasthya Bhawan, Behind IGIMS, Sheikhpura,

Adjacent to State Health Society, Patna- 800014, Bihar

Telephones: 0612-2219634                                                                e-mail: md-bmsicl-bih@nic.in

INVITATION FOR E-BIDS FOR

SUPPLY, INSTALLATION & COMMISSIONING OF MEDICAL EQUIPMENT ON TURNKEY BASIS

[Modify as appropriate to indicate general description of items under procurement]

| Tender Reference No. | BMSICL/2024-25/ME-382 |
| --- | --- |
| Date of Pre- Bid Meeting and Submission of Pre-bid queries | 29th November 2024 at 15:00 Hrs in Conference hall of BMSICL, 3rd Floor, Swasthya Bhawan, Behind IGIMS, Sheikhpura, Adjacent to State Health Society, Bihar. All pre-bid queries to be submitted through e-mail on bmsicltenderequipment@gmail.com upto 29th November 2024 till 17:00 Hrs, further to this date no more query will be entertained |
| Last date and time of submission of online bids | 16th December 2024 up to 17:00 Hrs. |
| Last date and time for submission of original documents of EMD, and technical bid. | 17th December 2024 till 14:00 Hrs. |
| Date, Time and Place of opening of Technical Bid | 17th December 2024 (at 15:00 Hrs.) on the website of https://eproc2.bihar.gov.in in the office of BMSICL |
| Date and time of opening of financial Bids | To be announced later on https://eproc2.bihar.gov.in |
| Validity of Tender | 180 Days |
| Cost of the tender document | Rs.11,800 /- (Eleven Thousand Eight Hundred only) Non- refundable. |
| Tender Processing Fee | Rs 590/- (on the website of https://eproc2.bihar.gov.in) |

1.   To participate in e-Tendering the tenderer will have to be registered with e-Tendering service provider. For this, help desk – mjunction Services Limited, RJ Complex, 2nd Floor, Canara Bank, Campus, Khajpura,  Ashiana  Road,  P.S. – Shastri  Nagar,  Patna-800014,  Toll  Free  No. - 18005726571, Email-ID: eproc2support@bihar.gov.in  can be approached.

2. The cost of tender document is acceptable only in Online mode (on the website: https://eproc2.bihar.gov.in) and it is non-refundable.

3.   The required amount of Earnest Money is acceptable in the form of Bank Guarantee issued by nationalized / schedule bank in favour of Managing Director, Bihar Medical Services & Infrastructure Corporation Limited, Patna and payable at Patna (Only Offline mode). The Earnest Money deposited in any other form shall not be acceptable.

4.   The Tender Inviting Authority reserves the right to extend the schedule of tender or to reject the tender without assigning any reason.

5.   The fee of bid processing is to be deposited by the tenderer through net banking i.e. RTGS / NEFT / Debit Card. The tenderer must ensure the payment before schedule time otherwise the corporation will not be responsible for any delay.

6.   It is essential to deposit hard copy of EMD Fee in the form of BG (Offline mode), of quoted item in sealed envelope at Bihar Medical Services & Infrastructure Corporation Limited, Patna by 17th December 2024 at 14.00 Hrs.

7.   Note: Please number the documents with serial number on each and every page and do mention the total number of pages of bidding document. In technical Bid parallelly assign the corresponding page numbers of supporting documents. Any discrepancy or misrepresentation in this aspect will not be entertained.

8.   Any queries and questions regarding the tender should be addressed to MD BMSICL (either

through letter or through e-mail:- md-bmsicl-bih@nic.in and / or bmsicltenderequipment@gmail.com or contact no. 0612-2219634/35) up to 7 days before of closing of online bid registration.

9.   All communication, addendum / corrigendum related to this tender will be issued on the website of http://eproc2bihar.gov.in

10. Managing Director, BMSICL reserves the right to reject any or all the applications without assigning any reason.

Note- Changes suggested by prospective bidder in Pre-Bid meeting may be incorporated in the tender document at the sole discretion of the tender inviting authority and for which corrigendum will be issued separately and uploaded on the website https://eproc2.bihar.gov.in & www.bmsicl.gov.in

Managing Director

BMSICL, Patna

TABLE OF CONTENT

Invitation for e- bids (IFB)........................................................................................................ 5

Section I- instruction to bidders (ITB) ...................................................................................... 7

A.         Introduction .................................................................................................................. 8

1.         Scope of Bid ................................................................................................................. 8

2.         Fraud and Corruption .................................................................................................... 8

3.         Eligible Bidders ............................................................................................................ 9

4.         One Bid Per Bidder....................................................................................................... 9

5.         Cost of Bidding ............................................................................................................. 9

6.         Alternative Tender....................................................................................................... 10

B.        The Bidding Documents .............................................................................................. 11

7.         Contents of Bidding Documents.................................................................................. 11

8.         Clarification of Bid Documents................................................................................... 11

9.         Pre-bid Meeting ........................................................................................................... 11

10.      Amendment of Bidding Documents ............................................................................ 12

C.        Preparation of Bids .................................................................................................... 13

11.      Language of Bid.......................................................................................................... 13

12.      Documents Constituting the Bid ................................................................................. 13

13.      Bid Form ..................................................................................................................... 13

14.      Bid Prices ................................................................................................................... 13

15.      Documents Required to be Submitted......................................................................... 14

16.      Documents Establishing Bidder’s Qualification ......................................................... 14

17.      Documents Establishing Goods Conformity to Bidding Documents......................... 15

18.      Earnest Money Deposit (EMD) ................................................................................... 15

19.      Period of Validity of Bids ........................................................................................... 16

20.      Preparation of Bid ....................................................................................................... 17

D. Submission of Tenders ...................................................................................................... 19

21.      Method of Bids Submission ....................................................................................... 19

22.      Deadline for Submission of Bids ................................................................................ 20

23.      Late Bids ..................................................................................................................... 20

24.      Modification and Withdrawal of Bids ........................................................................ 20

E. Bid Opening and Evaluation .............................................................................................. 21

25.      Opening of Bids by Purchaser..................................................................................... 21

26.      Clarification of Bids.................................................................................................... 21

27.      Preliminary Evaluation ............................................................................................... 21

28.      Evaluation and Comparison of Substantially Responsive Bids .................................. 22

29.      Contacting the Purchaser ............................................................................................ 24

F. Award of Contract .............................................................................................................. 25

30.      Post-qualification......................................................................................................... 25

31.      Award Criteria............................................................................................................. 25

32.      Purchaser’s Right to Vary Quantities .......................................................................... 25

33.      Purchaser’s Right to Accept any Bid and to Reject Any or all Bids ........................... 25

34.      Issue of Notification of Award .................................................................................... 25

35.      Signing of Contract ..................................................................................................... 26

36.      Performance Security .................................................................................................. 26

37.      General Guidelines for the Submission of e-tender .................................................... 26

INVITATION FOR E- BIDS (IFB)

For Supply, Installation & Commissioning of Medical Equipment on Turnkey Basis for

Govt. Medical College & Hospital, Mahua (Vaishali) and Govt. Medical College & Hospital, Madhubani, Bihar

By Managing Director,

Bihar Medical Services & Infrastructure Corporation Limited

2nd & 3rd Floor, Swasthya Bhawan, Behind IGIMS, Sheikhpura,

Adjacent to State Health Society, Patna- 800014, Bihar

Bid Reference No.: BMSIC/2024-25/ME-382                                 Date:       November 2024

1.  The Bihar Medical Services and Infrastructure Corporation Limited, Patna (name of purchaser) on behalf of the Government of Bihar, invites e-bids from sole bidder as defined in section-II of this tender document / OEM of any one of the equipment listed in the schedule of requirement for Supply, testing, Demonstration, Installation and Commissioning of Medical Equipment and related services as listed below on Turnkey basis for Govt. Medical College & Hospital, Mahua (vaishali) and Govt. Medical College & Hospital, Madhubani, Bihar

| Schedule No. | Brief Description of Goods and Services on TURNKEY basis | Qty./No. | Delivery Schedule (in days) | Earnest Money Deposit (EMD) in Indian Rupees |
| --- | --- | --- | --- | --- |
| 1 | Supply, Installation & Commissioning of Medical Equipment on Turnkey Basis for Govt. Medical College & Hospital, Mahua (Vaishali) and Govt. Medical College & Hospital, Madhubani, Bihar | (As mentioned in section –IV) of the bid document | As per GCC Clause -8 | 1,26,00,000/- (One Crores Twenty-Six Lakhs only) |

2.  The Qualification Criteria, Detailed Technical Specifications, Scope of Work, Cost of Tender Document, Earnest Money Deposit and other conditions can be seen in the tender document to be downloaded from the website of  http://eproc2bihar.gov.in.

3.  The bids must be uploaded (e-mode/ online) at the address given at page 2 on or before 17:00 hrs. of 16th December 2024. All bids must be accompanied by an Earnest Money Deposit (EMD) as specified in the bidding document. Bids submitted after 17:00 hrs. of 16thDecember 2024 shall be rejected.

4.  The Pre-bid meeting shall be organized at the purchaser’s office on 29th November 2024 at 15:00

Hrs. In the Pre-bid meeting, the prospective bidders may seek clarification on any issues related to the terms, conditions and technical specifications given in the bidding documents.

5. Technical bids will be opened (in e- mode) at Bihar Medical Services & Infrastructure Corporation Ltd., , 3rd Floor, Swasthya Bhawan, Behind IGIMS, Sheikhpura, Adjacent to State Health Society, Bihar 17th December 2024  at 15.00 Hrs. on the website of http://eproc2bihar.gov.in. The bidder’s representatives may attend the bid opening meeting.

6.  The Purchaser reserves the right to cancel / annul the bidding process without assigning any reason thereof.

7.  In the event, the dates specified for the bid receipt and opening are declared as holidays for purchaser’s office, the due date for submission of bids and opening of bids shall be the immediate following working day at the appointed time.

Managing Director

BMSICL, Patna

SECTION I- INSTRUCTION TO BIDDERS (ITB)

INSTRUCTIONS TO BIDDERS

A         INTRODUCTION

1.  SCOPE OF BID

The Bihar Medical Services & Infrastructure Corporation Limited, Patna (name of purchaser) on behalf of Government of Bihar, invites e-bids from sole bidder / OEM (Original Equipment Manufacturer) of any one of the equipment listed in schedule of requirement for Supply, Testing, Demonstration, Installation and Commissioning of Medical Equipment on turnkey basis for Govt. Medical College & Hospital, Mahua (Vaishali) and Govt. Medical College & Hospital, Madhubani Bihar as specified in the Schedule of Requirements.

2.  FRAUD AND CORRUPTION

2.1	It is required that the Purchasers as well as Bidders / Suppliers / Contractors observe the highest standard of ethics during the procurement and execution of Contracts. In pursuance of this policy, the Purchaser:

(a) defines, for the purposes of this provision, the terms set forth below as follows:

(i)   “corrupt practice” means the offering, giving, receiving, or soliciting of anything of value to influence the action of a public official in the procurement process or in Contract execution; and

(ii)   “fraudulent practice” means a misrepresentation of facts and / or concealment of fact in order to influence a procurement process or the execution of a Contract to the detriment of the Purchaser; it includes collusive practices among bidders (prior to or after bid submission) designed to establish bid prices at artificial, noncompetitive levels and to deprive the Purchaser of the benefits of free and open competition

(b) will declare a firm ineligible and debar the firm, either indefinitely or for a stated period of time, to be awarded a Contract if it at any time determines that the firm has engaged in corrupt or fraudulent practices in competing for, or in executing, a contract. In such cases, appropriate legal action as per court of law shall be initiated for which the concerned bidder shall be solely responsible.

2.2	Furthermore, bidders shall be aware of the provision stated in Sub-Clauses 19.4 and 22.1.d. of the General Conditions of Contract

3.  ELIGIBLE BIDDERS

3.1(A) The eligible bidder should be a legal entity and have satisfactory supply of similar nature in the last 05 (Five) years till the date of bid opening according to any one of the following requirements. Three similar supply, installation/Commissioning work of each costing not less than the amount equal to 40% of the estimated cost (Rs 73 Crores) put to tender Or Two similar supply, installation/Commissioning work of each costing not less than the amount equal to 60% of the estimated cost (Rs 73 Crores) put to tender Or One similar supply, installation/Commissioning work of costing not less than the amount equal to 80% of the estimated cost (Rs 73 Crores) put to tender. - “Similar Work shall mean supply, installation and commissioning of medical equipment/ Medical Furniture/ Medical Devices on turnkey basis as mentioned in schedule of requirement in any Government Hospital/ Government Medical College/ Government Medical Institutions in India.” Bidder must have to submit completion certificate of the same at the time of bid submission and the completion certificate shall be issued by Competent Authority (procurement authority/ end users).

OR

3.1 (B)

- The bidder must have experience in supplying, operating, and maintaining medical services in any of the Government hospital/ Medical college/ Medical institution of India for at least last five preceding financial years.

- The value of the single medical service contract should be at least 80% of the bid value (Rs. 73 crores).

- The bidder must have to submit same satisfactory performance certificate issued by authority/ user by mentioning the value of running/ completed medical service contract supporting with the respective work order/ contract/ MoU.

3.2  A firm declared ineligible by the Purchaser in accordance with ITB Sub-Clause 2.1 (b) and GCC Sub-Clause 19.4 shall be ineligible to bid for a contract during the period of time determined by the Purchaser.

Pursuant to ITB Clause 12, the Bidder shall furnish, as part of its bid, documents establishing,

to the Purchaser’s satisfaction, the Bidder’s eligibility to bid.

3.3 The sole bidder must have average annual turnover of Rs 100 Crores (average of any three years during the last 5 years). Bidder must submit copies of audited balance sheet and profit and loss statement detailing the same. Joint ventures and consortiums are not allowed for bidding.

4.  ONE BID PER BIDDER

A firm shall submit only one bid.  The firm to quote for all equipment as listed in the schedule of requirement as amended.  Part Bids will not be considered.  Joint ventures will not be considered.

5.  COST OF BIDDING

The bidder shall bear all costs associated with the preparation and submission of the bid. The Purchaser will, in no case, be responsible or liable for these costs, regardless of the conduct or outcome of the biding process.

6.  ALTERNATIVE TENDER

Alternative Tenders are not permitted. However, the Tenderers can quote T w o alternate Models/Make meeting the tender specifications with single EMD.

B. THE BIDDING DOCUMENTS

7.  CONTENTS OF BIDDING DOCUMENTS

7.1	The goods required to be supplied; bidding procedures and contract terms and conditions are prescribed in the Bidding Documents. The Bidding Document include the following:

Section I 	Instructions to Bidders (ITB)

Section II 	General Conditions of Contract (GCC) Section III	Special Conditions of Contract (SCC) Section IV	Schedule of Requirements (SOR) Section V	Technical Specifications

Section VI	Sample Forms

7.2	The “Invitation for Bids” does not form part of the Bidding Documents and is included as a reference only. In case of discrepancies between the Invitation for Bid and the Bidding Documents listed in 7.1 above, said Bidding Documents will take precedence.

7.3	The Bidder is expected to examine all instructions, forms, terms and specifications in the Bid Documents. Failure to furnish all information required as per the Bid Documents or submission of the bids not substantially responsive to the Bid Documents in every respect will be at the bidder’s risk and may result in rejection of the bid.

8.  CLARIFICATION OF BID DOCUMENTS

8.1	A prospective bidder, requiring any clarification on the Bid Documents shall notify the Purchaser in writing or by FAX/e-mail at the Purchaser’s mailing address indicated in the invitation of Bid.

8.2	Any clarification issued by the Purchaser in response to query raised by prospective bidders shall be published on the website of BMSICL and shall form an integral part of bid documents and it may amount to an amendment of relevant clauses of the bid documents.

9.  Pre-bid Meeting

9.1	The bidder or his representative is invited to attend a pre-bid meeting, which will take place in the office of BMSICL Patna on 29th November 2024  up to 15:00 Hrs.

9.2	The purpose of the meeting will be to clarify issues and to answer questions on any matter that may be raised at that stage.

9.3	The bidder may submit any query in writing or by FAX/ e-mail to reach the purchaser well before the time to be considered during the pre-bid meeting and not after 48 hours since the pre-bid meeting.

9.4	Any modification of the bidding document listed in ITB Clause 7.1 which may become necessary as a result of the pre-bid meeting shall be made exclusively through the issue of an Addendum pursuant to ITB Clause 10 and not through the minutes of the pre-bid meeting & the same will be uploaded at http://eproc2.bihar.gov.in  & or www.bmsicl.gov.in.

9.5       Non-attendance at the pre-bid meeting will not be a cause for disqualification of a bidder.

10.      AMENDMENT OF BIDDING DOCUMENTS

10.1	At any time, prior to the date of submission of Bids, the Purchaser may, for any reason, whether at its own initiative or in response to a clarification requested by a prospective bidder, or pursuant to ITB Clause 9, modify bid documents by amendments.

10.2	The amendments shall be notified by uploading the same at http://eproc2.bihar.gov.in and/or website of BMSICL i.e.  www.bmsicl.gov.in

10.3	In order to afford prospective bidders a reasonable time to take the amendment into account in preparing their bids, the purchaser may, at its discretion, extend the deadline for the submission of bids suitably.

C. PREPARATION OF BIDS

11.      LANGUAGE OF BID

The bid, as well as all correspondence and documents relating to the bid exchanged by the Bidder and the Purchaser, shall be written in English language. However, the purchaser as well as bidder may correspond in Hindi language also.

12.      DOCUMENTS CONSTITUTING THE BID

The bid prepared by the bidder shall comprise the following components:

(a)    A Bid Form and a Price Schedule completed in accordance with ITB Clauses 13 and 14;

(b)    Documentary evidence established in accordance with ITB Clause 15 and 16 that the

Bidder is eligible and qualified to perform the contract if its bid is accepted;

(c)    Documentary evidence established in accordance with ITB Clause 17 that the goods and ancillary services to be supplied by the Bidder conform to the bidding documents.

(d)    Earnest Money Deposit (EMD) furnished in accordance with ITB Clause 18.

(e)  Tender Document fee is acceptable only in Online mode (On the website:

https://eproc2.bihar.gov.in 2) and it is non-refundable.

(f)     Tender Document fee in the form of Demand Draft in favour of Managing Director, Bihar

Medical services and Infrastructure Corporation Ltd. Payable at Patna.

13.      BID FORM

The bidder shall complete the Bid Form and appropriate Price Schedule, indicating the goods to be supplied, brief description of the goods, quantity and prices as per section VI.

14.      BID PRICES

14.1	The bidder shall give the total composite price along with price of each equipment exclusive of GST but inclusive of Customs duty, packing, forwarding, freight, octroi/entry tax and insurance etc. Price should be as per the price schedule given in Section VI. No Foreign exchange will be made available by the purchaser.

14.2     Break-up of the prices indicated in the Price Schedule shall be entered in the following manner:

(i)  The Unit price should be inclusive of, Excise duty, Sales Tax, Freight, octroi/entry tax

Forwarding, Packing, Insurance and any other Levies/Charges etc.

(ii)  The supplier shall quote as per price schedule given in section VI for all the items given in schedule of requirement.

(iii)	The bidder can quote for both Medical College & Hospitals as well as for single Medical College & Hospital also. L1 will be decided on the basis rate quoted per Medical College-wise.

(iv) Bidder has to quote individual rates for individual Medical College & Hospital.

14.3	The price quoted by the bidder shall remain fixed during the entire period of contract and shall not be subject to variation on any account. A bid submitted with an adjustable price quotation will be treated as non - responsive and rejected.

14.4	The prices quoted by the bidder shall be in sufficient detail to enable the Purchaser to arrive at the price of material/goods offered.

14.5	“DISCOUNT, if any, offered by the bidders shall not be considered unless specifically indicated in the price schedule. Bidders desiring to offer discount shall therefore modify their offers suitably while quoting and shall quote clearly net price taking all such factors like Discount, free supply, etc, into account”.

14.6	The price approved by the Purchaser for procurement will be FOR destination which will be inclusive of all Taxes, Levies, packing, forwarding, freight and insurance as mentioned in Para 14.1, above.

15.      DOCUMENTS REQUIRED TO BE SUBMITTED

15.1     The bidder shall furnish, as part of the bid documents, the documents as called for in the

Check List (Annexure – 13).

15.2	The offered product may be required to be type approved / demonstrated at the time and place of installation. The supplied product would be evaluated for its ability to meet the technical specifications, manufacturability, reliability, testability, ease of installation, maintainability etc. Necessary documents to substantiate these attributes will have to be submitted at the time of application for approval by the supplier for obtaining type approval.

Or

Products can be demonstrated only at Site during installation and commissioning at Govt. Medical College & Hospital, Mahua (vaishali) and Govt. Medical College & Hospital, Madhubani, Bihar Or

In  case  goods  offered  have  already  been  type  approved/  validated  by  the  Purchaser, documentary evidence to this effect shall be submitted by the bidder.

16.      DOCUMENTS ESTABLISHING BIDDER’S QUALIFICATION

16.1    Pursuant to ITB Clause 12, the bidder shall furnish, as part of its bid, documents establishing the Bidder’s qualification to perform the Contract if its bid is accepted.

16.2    The documentary evidence of the Bidder’s qualifications as per eligibility criteria to perform the Contract shall establish to the Purchaser’s satisfaction that:

a)  The bidder should furnish the information on past works and satisfactory performance for being eligible in accordance with eligibility for the bidder in the Performa given under Section VI, Form No.6 and provide self-attested copies of Orders for the works for which performance certificate in form no 6 is being provided.

b) The Bidders shall invariably furnish documentary evidence in support of the supply, installation/commissioning certificate in the form of performance certificates issued by those end users whose purchase orders are referred in the performance statement as submitted in Annexure- VI by the bidders “. The  supply, installation/commissioning certificate shall be issued by Competent Authority (procurement authority/ end users).

c)  The bidder should furnish authorization in the prescribed format given at Section VI, assuring full guarantee and warranty obligations as per GCC Clause 14 for the equipment offered.

d)  The bidder must have average turnover of Rs 100 Crores (average of any three years during the last 5 years). Bidder must submit copies of audited balance sheet and profit and loss statement detailing the same in support of its financial standing.

16.3	Submission of Manufacturer’s authorization and end of life certificate for the equipment manufactured by other than the bidder in Annexure-V is mandatory at the time of supply.   However, this may not be compulsory in case of low cost (FOC) equipment.

17.	DOCUMENTS     ESTABLISHING     GOODS     CONFORMITY     TO     BIDDING DOCUMENTS

17.1	Pursuant to ITB Clause 12, the Bidder shall furnish, as part of its bid, documents establishing the conformity to the bidding documents of all goods and services which the Bidder proposes to supply under the contract.

17.2	The documentary evidence of conformity of the goods and services to the bidding documents may be in the form of literature, drawings and data, and shall consist of :

(a)   A detailed description of the essential technical and performance characteristics of the goods;

(b)	An item-by-item commentary on the Purchaser's Technical Specifications, demonstrating substantial responsiveness of the goods and services to those specifications or a statement of deviations and exceptions to the provisions of the Technical Specifications.

17.3	For purposes of the commentary to be furnished pursuant to ITB Clause 17.2 (b) above, the Bidder shall note that standards for workmanship, material, and references to brand names or catalogue numbers designated by the Purchaser in its Technical Specifications are intended to be descriptive only and not restrictive. The Bidder may substitute alternative standards, brand names and/or catalogue numbers in its bid, provided that it demonstrates to the Purchaser's satisfaction that the substitutions ensure substantial equivalence to those designated in the Technical Specifications.

18.      EARNEST MONEY DEPOSIT (EMD)

18.1	Pursuant to ITB Clause 12, the bidder shall furnish, as part of his bid, an Earnest Money Deposit (EMD) for an amount of mentioned in Section IV – Schedule of Requirements in the form of Bank Guarantee (Only Offline mode). No other mode of payment is acceptable.

18.2	The Earnest Money Deposit (EMD) is required to protect the purchaser against the risk of bidder’s conduct, which would warrant the forfeiture of Earnest Money Deposit (EMD) pursuant to ITB Clause 18.7.

18.3   The Earnest Money Deposit (EMD) shall be in the form of Bank Guarantee issued by a Nationalized/ Scheduled Bank in the favour of Purchaser here it is Managing Director, BMSICL Patna.

(i) The bank guarantee of adequate amount covering the requirement of EMD should be valid for a period of 30 days beyond the validity of Bid.

(ii)  BG issued to cover the requirement of EMD that should be issued from Nationalized

Bank/ Scheduled Bank.

(iii)  The BG should be submitted in the technical bids in a separate cover. The cover should be subscribed as “EMD Fee in the form of BG for tender no. BMSICL/2024-25/ME-382”.

(iv) In case where the document of Earnest Money Deposit (EMD) is not submitted in the manner prescribed above, the commercial, technical offers SHALL NOT BE OPENED AND THE BID SHALL BE REJECTED.

Note: - Bank Guarantee to be provided in the format provided in Annexure- 4.

18.4	A bid not secured in accordance with para 17.1, and 17.3 shall be rejected by the Purchaser being non-responsive at the bid opening stage and returned to the bidder unopened.

18.5	The Earnest Money Deposit (EMD) of the unsuccessful bidder will be discharged/returned as promptly as possible, but after finalization of tender. No interest will be paid against EMD and or performance security deposited by the bidders and no presentation will be allowed in this case.

18.6	The successful bidder’s Earnest Money Deposit (EMD) will be discharged upon the bidder’s acceptance of the advance purchase order satisfactorily in accordance with GCC Clause 5 and furnishing the performance security.

18.7     The Earnest Money Deposit (EMD) may be forfeited:

(a)    If the bidder withdraws his bid during the period of bid validity as specified in this bidding document

(b)   In the case of successful bidder if the bidder fails:

(i)     To sign the contract in accordance with ITB Clause 30 or

(ii)    To furnish performance security in accordance with GCC Clause 5.

19.      PERIOD OF VALIDITY OF BIDS

19.1     Bid shall remain valid for 180 days from the date of opening of bids prescribed by the purchaser pursuant to ITB Clause 25.1. A bid valid for a shorter period shall be rejected by the purchaser being non-responsive.

19.2	In exceptional circumstances, the purchaser may request the consent of the bidder for an extension to the period of bid validity. The request and the response there to shall be made in writing. The Earnest Money Deposit (EMD) provided under ITB Clause 18 shall also be suitably extended. The bidder may refuse the request without forfeiting his Earnest Money Deposit (EMD). A bidder accepting the request and granting extension will not be permitted to modify his bid.

20.      PREPARATION OF BID

20.1     The Bid shall be submitted online and in physical form in parts / covers as mentioned below:

- Tender Fee (Online only) & EMD (only in the form of Bank Guarantee which to be submitted   offline).

(ii)       Tender Processing Fee (Only Online) (iii)  	Technical Bid (Only Online)

(iv) 	Price Bid (Only Online).

Bidders are Bidders are requested not to requested to submit EMD of quoted item of tender (Offline in the form of BG). Submit hard copy of tender document. In case the hard copy of financial bid is submitted the tender shall be straightway rejected.

Also, uploading of the price bid in prequalification bid or technical bid will result in rejection of the tender. The entire Document which will become part of the tender (Online, Physical) should be either typed or written in indelible ink and the same shall be signed (& with official seal) by the tenderer or by a person(s) who has been duly authorized to bind the tenderer to the contract. The tender shall be duly signed at the appropriate places as indicated in the TE documents and all other pages of the tender including printed literature, if any shall be initiated by the same person(s) signing the tender.

The tender shall not contain any erasure or overwriting, except as necessary to correct any error made by the tenderer and, if there is any such correction; the same shall be initialed by the person(s) signing the tender. The entire document being part of tender document should be page numbered. A person signing (manually or digitally) the tender form or any documents forming part of   the contract on behalf of another shall be deemed to warranty that he has authority to bind such other persons and if, on enquiry, it appears that the persons so signing had no authority  to do so, the purchaser may, without prejudice to other civil and criminal remedies, cancel the contract and hold the signatory liable for all cost and damages.

20.2	In case Bidder is clamming for exemption from payment of Earnest Money, in accordance with SCC clause, then documentary evidence must be submitted in both Physical and in Online Mode.

- Prices are to be quoted in the attached

- Price Bid format online as per the directions on the official website.

Note: - It is the responsibility of tenderer to go through the TE document to ensure furnishing all required documents in addition to above, if any.

D. SUBMISSION OF TENDERS

21.      METHOD OF BIDS SUBMISSION

21.1

a)  The tender shall be submitted in online and in physical form as mentioned in ITB

Clause 20.

b) 	Technical bid should contain the clause-by-clause compliance statement for the quoted goods vis-à-vis the technical specifications in the tender enquiry in addition to

other required document as mentioned in TE Document.

c)  Technical bid should contain the brochure, catalogue of offered/ quoted items which should reasonably explain in detail about the quoted items & it should also confirm the clause –by-clause compliance of technical specification as asked in TE Document and other technical details incorporated by the purchaser in the TE documents to establish technical responsiveness of the goods and services offered in its tender.

d)  In case there is any variation and/or deviation between the goods & services prescribed by the purchaser and that offered by the tenderer, the tenderer shall list out the same

in a chart form without ambiguity and provide the same along with its tender.

e)  If a tenderer furnishes wrong and/or misguiding data, statement(s) etc. about technical acceptability of the goods and services offered by it, its tender will be liable to be ignored and rejected in addition to other remedies available to the purchaser in this regard.

f)	Failure in complying above mentioned clause 21.1 (a) – (e)., may lead to rejection of tender.

g)  Bidders are requested not to submit the hard copy of Financial Bid, along with the physical documentary evidence of submission of EMD of tender, Technical Bid. In

case the hard copy of financial bid is submitted in physical form, the tender shall be straightway rejected.

h)  Also, uploading of the price bid in prequalification bid or technical bid will result in rejection of the tender Unless otherwise specified, the tenderers are to submit its tender

online and deposit the physical form of tenders EMD, Technical bid and if applicable documentary support for seeking exemptions of EMD as per SCC clause are to be

submitted in physical form, no other documents are required to be submitted in physical form) in sealed envelope to the purchaser address.

| 21.2. | a) | The envelopes shall be addressed to the purchaser at the following address: Bihar Medical Services and Infrastructure Corporation Limited, 3rd Floor, Swasthya Bhawan, Behind IGIMS, Sheikhpura, Adjacent to State Health Society, Patna- 800014, Bihar |
| --- | --- | --- |
|  | b) | The envelope shall bear (the name and address of the Purchaser), the tender number |
|  |  | and the words ‘DO NOT OPEN BEFORE’ (due date & time) & may be sent by registered post or delivered in person on above mentioned address (address is given |
|  |  | in Clause 21.2 (a) above). The responsibility for ensuring that the Sealed envelope containing documentary evidence of EMD and / documentary support for seeking |
|  |  | exemptions of, EMD as per SCC clause are delivered in time would vest with the bidder and The purchaser shall not be responsible for any delay. In the event of the |
|  |  | specified date for physical submission of tender falls on /is subsequently declared a holiday or closed day for the purchaser, the tenders will be received up to the appointed time on the next working day. c) The Physical form of tender shall be delivered upto ..................,2024 by 14:00 Hrs BMSICL, 3rd Floor, Swasthya Bhawan, Behind IGIMS, Sheikhpura, Adjacent to State Health Society, Bihar India if delivered elsewhere will be rejected. |

c)  The Physical form of tender shall be delivered upto 17th December 2024 by

14:00Hrs BMSICL, 3rd Floor, Swasthya Bhawan, Behind IGIMS, Sheikhpura,

Adjacent to State Health Society, Bihar India if delivered elsewhere will be rejected.

d)  Venue of bid opening:-17th December 2024  by 15:00 Hrs on the website of https://eproc2.bihar.gov.in  at BMSICL, Patna, If due to administrative reason, the venue of Bid opening is changed, it will be displayed prominently on the notice board of the Purchaser’s office/at the Website address https://eproc2.bihar.gov.in

Note: - If the envelopes are not sealed and marked as required at ITB Clause 21.1 and 21.2, the bid shall be rejected.

22.      DEADLINE FOR SUBMISSION OF BIDS

22.1	Bids must be received by the Purchaser at the address and up to the due date and time specified under ITB Clause 21.2.

22.2	The Purchaser may, at its discretion, extend this deadline for the submission of bids by amending the Bid Documents in accordance with ITB clause 10 in which case all rights and obligations of the purchaser and bidders previously subject to the deadline will thereafter be subjected to the deadline as extended.

23.      LATE BIDS

Any bid received by the purchaser after the deadline for submission of bids prescribed by the purchaser pursuant to ITB clause 22, shall be rejected and the physical form of technical bid will be returned unopened to the bidder.

24.      MODIFICATION AND WITHDRAWAL OF BIDS

24.1	No bid may be modified subsequent to the deadline for submission of bids. The bidder may modify or withdraw its bid after submission, provided that written notice of the modification or withdrawal is received by the purchaser prior to the deadline prescribed for submission of bids along with a written power of attorney authorizing the signatory of the withdrawal.

24.2	The bidder’s modification or withdrawal notice shall be prepared, sealed, marked and dispatched as required in the case of bid submission in accordance with the provision of ITB

Clause 21. A withdrawal notice may also be sent by FAX/ e-mail but followed by a signed confirmation copy by post not later than the deadline for submission/ uploading of bids.

24.3	Bids requested to be withdrawn in accordance with ITB Clause 24.1 above, shall be returned unopened to the Bidders.

24.4	No bid may be withdrawn in the interval between the bid submission deadline and the expiration of the bid validity period specified in ITB Clause 19. Withdrawal of a bid during this interval may result in the forfeiture of the Bidder’s Earnest Money Deposit (EMD), pursuant to ITB Clause 18.7

E. BID OPENING AND EVALUATION

25.      OPENING OF BIDS BY PURCHASER

25.1	The purchaser shall open the technical bids in the presence of bidders or their authorized representatives who chose to attend, at the due date and time of bid opening. The bidder’s representatives, who are present, shall sign in an attendance register. Authority letter to this effect shall be submitted by the bidders before they are allowed to participate in bid opening (A Format is given in Section VI).

25.2	A maximum of two representatives of any bidder shall be authorized and permitted to attend the bid opening.

25.3	The bidder’s names, modifications, bid withdrawals, requisite Earnest Money Deposit (EMD) and such other details as the purchaser, at its discretion, may consider appropriate will be announced at the time of opening. No bid shall be rejected at the time if bid opening, except for late bids, bids without EMD (except in case where exemption of EMD has       been

requested in pursuant to Special condition of Contract) & for such rejected bid no further evaluation will be done.

25.4	The price bids of bidders whose Technical bids are found technically responsive and comply with the bid documents will only be considered for financial evaluation. The date of opening of financial bids shall be communicated to such bidders, whose Technical bids are found technically responsive. The bidder’s representative may be present at the time of opening of price bid at the pre-appointed time, date and venue.

25.5	The date fixed for opening of bids, if subsequently declared as holiday by the Government, the revised date of schedule will be notified. However, in absence of such notification, the bids will be opened on next working day, time and venue remaining unaltered.

26.      CLARIFICATION OF BIDS

To assist in the examination, evaluation and comparison of bids, the purchaser may, at its discretion ask the bidder for the clarification of its bid. The request for the clarification and the response shall be in writing. Unless the purchaser asks for change in price due to clarifications sought, the bidder is not permitted to alter the price Bid (online submission only).

27.      PRELIMINARY EVALUATION

27.1	Purchaser shall evaluate the bids to determine whether they are complete, whether any computational errors have been made, whether required sureties have been furnished, whether the documents have been properly signed and whether the bids are generally in order. Bids from representatives, without proper Authorization from the manufacturer as per Section VI, may be treated as non-responsive. However, in case of an equipment without such authorization at the time of submission of bid, the undertaking from the bidder to comply with the conditions laid down in clause 16.3 supra as well as to cover full warranty/CMC of such equipment shall be taken into account for the purpose of  technical responsiveness under this clause.

27.2	Purchaser shall evaluate the bids to determine whether they are complete, whether any computational errors have been made, whether required sureties have been furnished, whether the documents have been properly signed and whether the bids are generally in order. Bids from representatives, without proper Authorization from the manufacturer as per Section VI, shall be treated as non-responsive except in cases of those equipment for which the bidder

gives undertaking to comply with the provisions of Clause 16.3 supra.

27.3	Arithmetical errors shall be rectified on the following basis. If there is a discrepancy between the unit price and total price that is obtained by multiplying the unit price and quantity, the unit price shall prevail and the total price shall be corrected by the purchaser.  If there is a discrepancy between words and figures, the amount in words shall prevail. If the supplier does not accept the correction of the errors, his bid shall be rejected.

27.4	Prior to the detailed evaluation pursuant to ITB Clause 28, the Purchaser will determine the substantial responsiveness of each bid to the Bid Document. For purposes of these clauses, a substantially responsive bid is one which confirms to all the terms and conditions of the Bid Documents without material deviations. Deviations from or objections or reservations to critical provisions such as those concerning Performance Security (GCC clause 5), Warranty (GCC clause 14), Force Majeure (GCC clause 21), Applicable Law (GCC clause 28) and Taxes and duties (GCC clause 30) along with deviation in Technical Specifications will be deemed as material deviation. The purchaser’s determination of bid’s responsiveness shall be based on the contents of the bid itself without recourse to extrinsic evidence.

27.5	A bid determined as substantially non-responsive will be rejected by the purchaser and shall not be subsequent to the bid opening be made responsive by the bidder by correction of the non-conformity.

27.6	The Purchaser may waive any minor infirmity or non-conformity or irregularity in a bid which doesn’t constitute a material deviation, provided such waiver doesn’t prejudice or affect the relative ranking of any bidder.

28.      EVALUATION AND COMPARISON OF SUBSTANTIALLY RESPONSIVE BIDS

28.1     Financial bid shall be considered for opening of whose technical bid which has been found eligible by the technical evaluation committee. The Purchaser shall evaluate in detail and compare the bids previously determined to be substantially responsive pursuant to ITB

Clause 27.

28.2	The purchasers evaluation of bid will take into account, in addition to the bid price (ex- factory/ex-warehouse/off-the-shelf price of goods offered from India, such price to include all costs as well as duties and taxes paid or payable on components and raw materials incorporated or to be incorporated in the goods, and excise duty on finished goods if payable) and price of incidental services, the following factors, in the manner and to the extent indicated in ITB clause 28.3 and in the Technical Specifications:

(a)       cost of inland transportation, insurance and other costs within India incidental to the delivery of goods to their final destination.

28.3     Pursuant to ITB clause 28.2 the following evaluation methods will be applied:

(a)     Inland transportation, ex-factory/ from port-of-entry, insurance and incidentals. (i)  Inland transportation, insurance and other incidentals, for delivery of goods to the

Project site as stated in ITB clause 14.2. These costs will be added to bid price

(b)    Deviation in Payment Schedule:

(c)     The General Conditions of Contract clause 15 indicate the payment schedule offered

by the Purchaser. If a bid deviates from the schedule and if such deviation is considered acceptable to the Purchaser, the bid will be evaluated by calculating interest earned for any earlier payments involved in the terms outlined in the bid as compared to those stipulated in this invitation at a rate of 12% per annum.

(d)    Spare parts and after sales service facilities in India:

The cost of the Purchaser of establishing the minimum service facilities and parts inventories, as outlined elsewhere in the bid invitation, if quoted separately, shall be added to the bid price.

(e)     Compressive Annual Maintenance Contract (CMC):

(i)  The Purchaser desires to have separately comprehensive maintenance charges for all equipment costing above Rs. 50,000/- in BOQ , for a period of 7 years after the expiry of the warranty period, clearly indicating  year-wise comprehensive maintenance charges, which shall not be considered for determining L1 Criteria to the bid price. In any case the CMC price should not be more that 6.20 % (Inc GST) of the quoted unit price of equipment. Bids without CMC charge will be considered non-responsive. Withdrawal or non-compliance of agreed terms and conditions after the execution of the contract will lead to invoking of penal provisions and may also lead to blacklisting of the successful bidder for a period of three years and forfeiture of Security deposit. In future, if the situation warrants, BMSICL reserves the right to sign the CMC agreement with the L1 bidder/OEM at the rate quoted by the respective firm in the price bid.

(ii) Any major repair pointed out by the Purchaser shall be rectified by the Supplier from the date of intimation within a period of 3 calendar days and commission the equipment to the satisfaction of the Purchaser, failing which the purchaser has write to levy a penalty on the Supplier a sum of Rs.2,500/- per day or part thereof for each equipment until the equipment are repaired and commission to the satisfaction of the Purchaser. Failure to repair may also lead to forfeit of Security deposit.

(f)      Spares:

(i)  The supplier shall be required to provide a list and rates of consumables required for an equipment which is a closed system.

(ii) The cost of spares quoted by bidder will not be used to arrive at final price.

(iii)In the event of termination of production of the equipment/ spare parts, the supplier shall notify the purchaser at least two years in advance of the impending termination to enable the purchaser to procure lifetime spares.  The supplier shall also provide at his own cost to the purchaser, the blueprint drawings and specifications of spare parts if and when

(g)     Repair of faulty equipment and setting up of Repair Facilities:

(i)  The supplier shall establish adequate repair facilities for repair of faulty equipment in

India within a period six months from the date of purchase order.

28.4     Technical evaluation:

(i)  Technical evaluation of the Bid will be done on the basis of technical qualification

criteria and documents mentioned (TECHNICAL BID- COVER ‘A’) in Mandatory Documents Link present in the web portal of the  https://eproc2.bihar.gov.in  .Failing which the bid will not be considered for technical evaluation.

(ii) Hard copy of tender documents uploaded shall not be submitted along with the EMD

as on or before the last day of submission of tender for purely evaluation purposes.

(iii)The  technical  evaluation  shall  be  done  only  on  the  basis  of  documents/papers submitted by the bidder on  https://eproc2.bihar.gov.in

29.      CONTACTING THE PURCHASER

29.1	Subject to ITB Clause 26, no bidder shall try to influence the Purchaser on any matter relating to its bid, from the time of the bid opening till the time the contract is awarded.

29.2     Any effort by a bidder to modify his bid or influence the purchaser in the purchaser’s bid

evaluation, bid comparison or contract award decision shall result in the rejection of the bid.

F. AWARD OF CONTRACT

30.      POST-QUALIFICATION

30.1	The Purchaser will determine to its satisfaction whether the Bidder that is selected as having submitted the lowest evaluated responsive bid is qualified to perform the Contract satisfactorily, in accordance with the criteria listed in ITB Clause 16 & 17 and the information submitted by the bidder in the proforma for performance statement for the period of last three years given in Sec VI as well as other information the Purchaser deems necessary and appropriate.

30.2	An affirmative post-qualification determination will be a prerequisite for award of the contract to the lowest evaluated Bidder. A negative determination will result in rejection of the Bidder’s bid, in which event the Purchaser will proceed to the next-lowest evaluated Bidder to make a similar determination of that Bidder’s capabilities to perform satisfactorily.

31.      AWARD CRITERIA

Subject to ITB Clause 33, the Purchaser shall award the Contract to the Bidder whose bid has been determined to be substantially responsive and has been determined to be the lowest evaluated bid and whose goods have been type approved/validated by the purchaser.

32.      PURCHASER’S RIGHT TO VARY QUANTITIES

The Purchaser reserves the right at the time of Contract award or within the stipulated last date of delivery, to increase or decrease, by 25%, the quantity of goods and services beyond that originally specified in the Schedule of Requirements without any change in unit price or other terms and conditions.

33.      PURCHASER’S RIGHT TO ACCEPT ANY BID AND TO REJECT ANY OR ALL

BIDS

The Purchaser reserves the right to accept or reject any bid, and to annul the bidding process and reject all bids, at any time prior to award of contract without assigning any reason whatsoever and without thereby incurring any liability to the affected bidder or bidders on the grounds of purchaser’s action.

34.      ISSUE OF NOTIFICATION OF AWARD

34.1	The issue of Notification of Award shall constitute the intention of the Purchaser to enter into contract with the bidder.

34.2	Prior to the expiration of the period of bid validity, the Purchaser will notify the successful Bidder in writing by registered letter or by cable, to be subsequently confirmed in writing by registered letter, that its bid has been accepted.

34.3	The bidder shall within 15 days of issue of the Notification of Award, to give their acceptance along with performance security in conformity with Section VI provided with the bid document.

35.      SIGNING OF CONTRACT

35.1     The issue of Notification of Award shall constitute the award of contract on the bidder.

35.2	Promptly after the Purchaser notifies the successful Bidder that its bid has been accepted, the Purchaser will send the Bidder the Contract Form provided in the Bidding Documents, incorporating all agreements between the parties.

35.3	Within fifteen (15) days of receipt of the Contract Form, the successful Bidder shall sign and date the Contract Form and return it to the Purchaser.

36.      PERFORMANCE SECURITY

36.1	Within fifteen (15) days of the receipt of notification of award from the Purchaser, the Contract, using the Performance Security Form provided in the Bidding Documents or in another form acceptable successful Bidder shall furnish the performance security in accordance with the Conditions of to the Purchaser.

36.2	Failure of the successful Bidder to comply with the requirement of ITB Clause 35 and ITB Sub Clause 36.1 shall constitute sufficient grounds for the annulment of the award and forfeiture of the Earnest Money Deposit (EMD), in which event the Purchaser may make the award to the next-lowest evaluated bid submitted by a qualified Bidder or call for new bids.

37.      GENERAL GUIDELINES FOR THE SUBMISSION OF E-TENDER

Instructions/ Guidelines for tenders for electronic submission of the tenders online have been annexed for assisting the prospective Tenderers to participate in e- Tendering.

a)  Registration of Tenderers: Any tenderer willing to take part in the process of e-Tendering will have to be enrolled & registered with the Government e- Procurement system, through logging on to https://eprocbihar.gov.in. The prospective Tenderer is to click on the link for

e-Tendering site as given on the web portal.

b)  Digital Signature certificate (DSC): Each Tenderer is required to obtain a class-II or Class- III Digital Signature Certificate (DSC) from NIC for submission of tenders, from the approved service provider of the National Information’s Centre (NIC) on payment of requisite amount.

c)  The Tenderer can search & download NIT & Tender Documents electronically from computer once he logs on to the website using the Digital Signature Certificate. This is the only mode of collection of Tender Documents.

d)  Submission of Tenders: General process of submission, Tenders are to be submitted through online to the website at a time for each work, one in technical Proposal & the other is Financial Proposal before the prescribed date &time using the Digital Signature Certificate (DSC) the documents are to be uploaded virus scanned copy duly Digitally Signed. The documents will get encrypted (transformed into non-readable formats).

TABLE OF CONTENT

Section II- General Conditions of Contract ......................................................................................................29

1.        Definitions ...............................................................................................................................................30

2.           Standards .................................................................................................................................................31

3.           Use of Contract Documents and Information; Inspection and Audit by the Purchaser .............................31

4.           Patent Rights ...........................................................................................................................................32

5.           Performance Security ..............................................................................................................................32

6.           Inspection and Tests................................................................................................................................32

7.           Packing ...................................................................................................................................................33

8.           Delivery and Documents .........................................................................................................................33

9.           Training ..................................................................................................................................................34

12.         Insurance ................................................................................................................................................35

13.         Transportation.........................................................................................................................................35

14.         Warranty / Shelf Life ...............................................................................................................................35

15.         Payment Terms .......................................................................................................................................37

16.         Prices ......................................................................................................................................................37

17.         Change Orders........................................................................................................................................38

18.         Subcontracts ...........................................................................................................................................38

19.         Delays in the Supplier’s Performance .....................................................................................................38

20.         Liquidated Damages ................................................................................................................................39

21.         Force Majeure .........................................................................................................................................39

22.         Termination for Default...........................................................................................................................40

23.         Termination for Insolvency ....................................................................................................................40

24.         Termination for Convenience ..................................................................................................................40

25.         Settlement of Disputes.............................................................................................................................41

26.         Limitation of Liability..............................................................................................................................41

27.         Governing Language................................................................................................................................42

28.         Applicable law........................................................................................................................................42

29.         Notices....................................................................................................................................................42

30.         Taxes and duties .....................................................................................................................................42

Section III- Special Conditions of Contract .....................................................................................................43

Special conditions of contract...............................................................................................................................44

Section IV- Schedule of Requirements .............................................................................................................45

Section-vi (sample forms) ...................................................................................................................................398

1.	Bid Form .................................................................................................................................................399

2.	Price Schedule .........................................................................................................................................401

3.	Form – 3 Form of Contract Agreement ....................................................................................................402

4. 	Performance Security Bank Guarantee .....................................................................................................404

5. 	Manufacturer’s Authorization Form .........................................................................................................405

6. 	Proforma for Performance Statement .......................................................................................................406

7. 	Consignee Receipt Certificate / Installation Report / Certificate ..............................................................407

8. 	Statement for Technical Deviation ...........................................................................................................408

SECTION II- GENERAL CONDITIONS OF CONTRACT

1.  DEFINITIONS

In this Contract, the following terms shall be interpreted as indicated:

(a) “The  Purchaser”  means  the  Bihar  Medical  Services  and  Infrastructure  Corporation

Limited (BMSICL), the organization purchasing the Goods.

(b) “The Bidder” means the individual or firm who participates in the tender and submits its bid.

(c) “Days” means calendar days.

(d) “Sole Bidder” means Bidder who will be a sole provider of all the equipments (Company including OPC /LLP/Partnership etc.) The Sole Provider should be registered as a legal entity such as company registered under Companies Act, Partnership Act or an equivalent law applicable in the region/state/ country.

(e) “GCC” means Conditions of Contract.

(f)  “The Supplier” means the individual or firm supplying the goods and Services under the contract.

(g) “The Goods” means all equipment, machinery, and/or other materials which the Supplier is required to supply to the Purchaser under the contract.

(h) “Services” means services ancillary to the supply of the Goods, such as transportation and insurance, and any other incidental services, such as installation, commissioning, provision of technical assistance, training and other obligations of the Supplier covered under the Contract.

(i)  “End User” means the consignees stated in the Schedule of Requirements.

(j)  “The Notification of Award” means the intention of the Purchaser to place the Purchase order on the bidder or to enter in to contract with the bidder.

(k) “The Contract” means the agreement entered into between the Purchaser and the Supplier, as recorded in the Contract Form signed by the parties, including all the attachments and the appendices thereto and all documents incorporated by reference therein.

(l)  “The Contract Price” means the price payable to the Supplier under the contract for the full and proper performance of its contractual obligations.

(m) “Validation” is a process of testing the equipment as per the specifications including requirements for use in hospital is carried out in simulated field environment.

1.1  Application: The General Conditions shall apply to the extent that they are not superseded by provisions in other parts of the contract.

2.  STANDARDS

The bidder shall procure, demonstrate, install & commission medical equipment for entire departments of said  medical college/s in accordance with the requirement  of the client as  per Specifications mentioned in bid document. The quality assurance for low cost (FOC) item must be undertaken by the bidder himself , with a warranty cover of at least one year. The bidder shall furnish, at the time of  delivery of  equipment,  the documents as called for in the  Check List  (Annexure  – 13): Manufacturer's Authorization, End of Life certificate  (if quoted by bidder other than manufacturer) as per Annexure 5. Quality Standard Certification (USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO) in accordance with technical specification of the specific equipment in this bid document. These quality certificates of the offered makes and models should be attached in technical bid documents. Bidder can offer TWO alternate makes & Models of an equipment matching with technical specification in all respect.

Before the equipment are procured and the orders are placed, the supplier shall get specifications, make

& model of the equipment and approved from the specification committee. No change or deviation in the broad specifications will be permitted in the schedule of requirements. The bidder will submit authorization certificate of the respective equipment to quote, execute the agreement, supply and provide the maintenance against this order.

The supplier shall provide OEM warranty for each medical equipment procured, installed & Commissioned under this contract. The minimum warranty period shall be 3 years.

End of life (EOL) Product-Bidder must quote for the model with latest hardware and versions and make

sure that no quoted equipment including hardware/software should come to an end of life within next five years from date of handing over to the Medical college & spare supports for 10 years. The bidder must submit a certificate from the respective company on their letter head in support of Non-EOL of the equipment at the time of delivery.

The goods supplied under this contract shall conform to the standards prescribed in the Technical

Specifications mentioned in section V.

3.  USE OF CONTRACT DOCUMENTS AND INFORMATION; INSPECTION AND AUDIT BY THE PURCHASER

3.1	The Supplier shall not, without the Purchaser’s prior written consent, disclose the Contract, or any provision thereof, or any specification, plan, drawing, pattern, sample, or information furnished by or on behalf of the Purchaser in connection therewith, to any person other than a person employed by the Supplier in the performance of the Contract. Disclosure to any such employed person shall be made in confidence and shall extend only as far as may be necessary for purposes of such performance.

3.2        The Supplier shall not, without the Purchaser’s prior written consent, make use of any

document except for purposes of performing the Contract.

3.3	Any document, other than the Contract itself, enumerated in GCC Sub-Clause 3.1 shall remain the property of the Purchaser and shall be returned (all copies) to the Purchaser on completion of the Supplier’s performance under the Contract if so required by the Purchaser.

3.4        The Supplier shall permit the Purchaser to inspect the Supplier’s accounts and records relating to the performance of the Contract and to have them audited by auditors appointed by the Purchaser, if so required.

4.  PATENT RIGHTS

The supplier shall indemnify the purchaser against all third-party claims of infringement of patent, trademark or industrial design rights arising from use of the goods or any part thereof in India.

5.  PERFORMANCE SECURITY

5.1	The supplier shall furnish performance security to the purchaser for an amount equal to 05 % of the value of purchase order within 15 days from the date of issue of Notification of Award by the Purchaser.

5.2        The proceeds of the performance security shall be payable to the Purchaser as compensation for

any loss resulting from the supplier’s failure to complete its obligations under the contract.

5.3	The performance security denominate in Indian Rupees shall be in the form of Bank Guarantee issued by a Scheduled/Nationalized Bank or demand draft. The performance security should be valid for the period beyond one hundred eighty (180) days following the date of completion of the Supplier’s performance obligations under the Contract, including any warranty/ shelf-Life Duration obligations

5.4	The performance security will be discharged by the Purchaser and returned to the Supplier not later than thirty (30) days following the date of completion of the Supplier’s performance obligations under the Contract, including any warranty and CMC (if executed by BMSICL) obligations.

6.  INSPECTION AND TESTS

6.1	The Purchaser or his representative shall have the right to inspect and test the goods as per prescribed test schedules for their conformity to the specifications. Where the Purchaser decides to conduct such tests on the premises of the supplier or its subcontractor(s), all reasonable facilities and assistance like Testing instruments and other test gadgets including access to drawings and production data shall be furnished to the inspectors at no charge to the purchaser. The supply will be accepted only after quality assurance tests are carried out by the Purchaser as per prescribed schedule and material passing the test successfully.

6.2	Should any inspected or tested goods fail to conform to the specifications the purchaser may reject them and the supplier shall either replace the rejected goods or make all alterations necessary to meet Specification requirements free of cost to the purchaser.

6.3	Notwithstanding the pre-supply tests and inspections prescribed in GCC Clause 6.1 & 6.2 above, the items /goods and accessories (if any) on receipt in the Purchaser’s premises will also be tested during actual but before “take over” and if any equipment/ items /goods or part thereof is found defective, the same shall be replaced free of all cost to the purchaser as laid down in GCC Clause 6.4 below.

6.4	If any goods/ material or any part thereof, before it is taken over under GCC Clause 6.5, is found defective or fails to fulfill the requirements of the contract, the inspector shall give the Supplier notice setting forth details of such defects or failure and the supplier shall make the defective item good, or alter the same to make it comply with the requirements of the contract forthwith and in any case within a period not exceeding three months of the initial report. These replacements shall be made by the supplier free of all charges at site. Should it fail to do so within this time, the purchaser reserves the discretion to reject and replace at the cost of the supplier the whole or any portion of items/ goods as the case may be, which is defective or fails to fulfill the requirements of the contract. The cost of any such replacement made by the purchaser shall be deducted from the amount payable to the supplier.

6.5	When the performance tests called for have been successfully carried out, the inspector / ultimate consignee will forthwith issue a Taking over Certificate. The inspector /ultimate consignee shall not delay the issue of any “taking Over Certificate” contemplated by this clause on account of minor defects in the items /goods which do not materially affect the commercial

/ actual/intended use thereof provided that the supplier shall undertake to make good the same in a time period not exceeding two months. The Taking Over Certificate shall be issued by the ultimate consignee within six weeks of successful completion of tests. In this case, a Consignee Receipt Certificate issued by the consignee as per the Format given in Section VI shall be equivalent to “Taking Over Certificate”, issuance of which shall certify receipt of goods in safe and sound condition. However, they shall not discharge the supplier of their warranty/ Shelf life obligation. The Consignee Receipt Certificate in respect of last consignment against the Contract will be equivalent to “Taking Over Certificate”.

6.6	Nothing in GCC Clause 6 shall in any way release the Supplier from any warranty or other obligations under this contract.

7.        PACKING

7.1	The Supplier shall provide such packing of the Goods as is required to prevent their damage or deterioration during transit to their final destination, as indicated in the Contract. The packing shall be sufficient to withstand, without limitation, rough handling during transit and exposure to extreme temperatures, salt, and precipitation during transit and open storage. Packing case size and weights shall take into consideration, where appropriate, the remoteness of the Goods’ final destination and the absence of heavy handling facilities at all points in transit.

7.2	The packing, marking and documentation within and outside the packages shall comply strictly with such special requirements as shall be provided for in the Contract including additional requirements, if any, specified in SCC and in any subsequent instructions ordered by the purchaser.

7.3        Packing Instruction: The supplier will be required to mark separate packages for each consignee.

Each package will be marked on three sides with proper paint/indelible ink, the following:

(i). Purchaser:

(ii). Contract No. (iii).  Supplier Name

(iv). Packing List reference Number

8.  DELIVERY AND DOCUMENTS

8.1	Upon or before delivery of the Goods, the Supplier shall notify the Purchaser in writing and deliver the following documents to the Purchaser:

(i)	One original and three copies of the contractor’s invoice, showing Purchaser, the Contract number, Goods’ description, quantity, unit price, and total amount. Invoices must be signed in original and stamped or sealed with the company stamp/seal;

(ii)    Three copies of delivery note, railway consignment notes, road consignment note, truck or air waybill, or multi-modal transport document showing Purchaser as Bihar Medical Services and Infrastructure Corporation Limited [ enter correct name of Purchaser for excise purposes] and delivery through to final destination as stated in the Contract;

(iii)   Copy of the Insurance Certificate, showing the Purchaser as the beneficiary; (iv)    Three copies of the packing list identifying contents of each package;

(v)     One original of the manufacturer’s or Supplier’s Warranty certificate covering all items supplied should be submitted along with the installation certificate. Warranty starts from the date it is successfully installed and warranty certificates should mention clearly date of installation and from that date warranty is for a period of 3 years.

(vi)    Original copy of the Certificate of Inspection furnished to Supplier by the nominated inspection agency if any.

The above certificate shall be received by the Purchaser upon arrival Warranty and installation certificate alone upon installation.

Note:   In the event that the documents presented by the Supplier are not in accordance with the Contract, then payment will be made against issue of the ‘Consignee Receipt Certificate’, to be issued in accordance with GCC Clause 6 above and after installation verification by BMSICL. The exact method of installation verification will be decided by Tender Inviting Authority.

8.2       The actual delivery schedule will be given in Schedule of Requirement and / Notification of Award/ supply order. The delivery of the goods and documents shall be completed within 90 days and for imported products the delivery schedule will be 120 days from the date of issue of supply order. However, the Delivery schedule may be extended further depending on the quantity to be delivered. The extension in delivery schedule of high-end equipment shall be decided as per the recommendations of TSC in the pre bid meeting.

8.3 	All Technical assistance for installation, commissioning and monitoring of the equipment shall be provided by the Supplier at no extra cost during laboratory evaluation, validation/ type approval and field trial, if any.

8.4 	The delivery period should include supply of items at the consignee place and there after successfully installation, demonstration of equipment at consignee place wherever required it should also include trial, run and commissioning.

9   TRAINING

9.1 Bidder shall demonstrate and provide training on use and proper application of    equipment to the consignees persons at Govt. Medical College & Hospital, Mahua (vaishali) and Govt. Medical College & Hospital, Madhubani, Bihar

10. INCIDENTAL SERVICES

10.1 The supplier may be required to provide any or all of the following services

(a)     Performance or supervision of on-site assembly and/or start-up of the supplied Goods; (b)     Furnishing of tools required for assembly and/or maintenance of supplied Goods;

(c)     Performance of supervision or maintenance and/or repair of the supplied Goods, for a

period of time agreed by the parties provided that this service shall not relieve the supplier of any warranty obligations under this contract.

(d)     Furnish detailed operations and maintenance manual for each appropriate unit of supplied goods.

11. Spares

11.1	The supplier shall be required to provide a list of the following material and notifications pertaining to spare parts manufactured or distributed by the supplier of spares including cost and quantity considered for arriving at the price of spares.

Such spare parts as the purchaser may elect to purchase from the supplier provided that such purchase shall not relieve the supplier of any warranty obligation under the contract.

In the event of termination of production of the spare parts, the supplier shall:

give advance notification to the purchaser pending termination (not less than 2 years), in sufficient time to enable the purchaser to procure life time spare; and

Following such advance intimation of termination, furnish at no cost to the purchaser, the blueprints, drawings and specifications of spare parts, if and when requested.

12. INSURANCE

12.1 The Goods supplied under the Contract shall be insured in an amount equal to 110% of the EXW value of the Goods from “warehouse to warehouse” on “all risks” basis including war risks and strikes.

13. TRANSPORTATION

Where the Supplier is required under the Contact to transport the Goods to a specified place of destination, defined in Consignee list, transport to such place of destination, including insurance and storage, as shall be specified in the Contract, shall be arranged by the Supplier, and related costs shall be included in the Contract Price.

14. WARRANTY/ SHELF LIFE

14.1     The supplier shall warrant that the goods to be supplied shall be new and free from all defects and faults in materials used, workmanship and manufacture and shall be of  the highest grade and consistent with the established and generally accepted standards for materials of the type ordered and shall perform in full conformity with the specifications and drawings. The supplier shall be responsible for any defect that may develop under the conditions provided by the contract and under proper use, arising from faulty material, design or workmanship such as corrosion of   the equipment, inadequate quantity of  material  to  meet  equipment requirements, inadequate contact protection, deficiencies in circuit design and/or otherwise and shall remedy such defects at his own cost when called upon to do so by the Purchaser who shall state in writing in what respect the stores are faulty. This warranty shall survive inspection or payment for / and acceptance of goods, but shall expire (except in respect of complaints notified prior to such date) three years after the goods have been taken over under GCC Clause 6.5

above.

14.2	This warranty shall remain valid for three years after the goods or any portion thereof as the case may be, have been delivered to and accepted at the final destination indicated in the contract.

14.3	If it becomes necessary for the Supplier to replace or renew any defective portion(s) of the equipment under this clause, the provisions of the GCC Clause 14.1 shall apply to the portion(s) of the equipment so replaced or renewed or until the end of the above- mentioned period of three years, whichever may be later. If any defect is not remedied by the supplier within a reasonable time, the Purchaser may proceed to get the defects remedied from other supplier etc., at the supplier’s risk and expenses, but without prejudice to any other rights which the purchaser may have against the supplier in respect of such defects.

14.4	Replacement under warranty clause shall be made by the supplier free of all charges at site including freight, insurance and other incidental charges.

14.5     A.        No conditional warranty will be acceptable.

B.         Warranty as well as Comprehensive Maintenance contract will be inclusive of all Accessories and Turnkey work if any and it will also cover the following wherever applicable: -

(i).             Any kind of motor.

(ii).            Plastic & Glass Parts against any manufacturing defects.

(iii).           All kind of sensors.

(iv).           All kind of coils, probes and transducers.

(v).            Printers and imagers including laser and thermal printers with all parts.

(vi).           Air-conditioners

C.         Replacement and repair will be undertaken for the defective goods.

D.         Proper marking has to be made for all spares for identification like  printing  of installation and repair dates.

14.6	In case of any claim arising out of this warranty, the Purchaser/Consignee shall promptly notify the same in writing and /e- mail to the supplier.

14.7	Upon receipt of such notice, the supplier shall, within 72 hours on a 24(hrs) X 7 (days) X 365 (days) basis respond to take action to repair or replace the defective goods or parts thereof, free of cost, at the ultimate destination. The supplier shall take over the replaced parts/goods after providing their replacements and no claim, whatsoever shall lie on the purchaser for such replaced parts/goods thereafter. The penalty clause for non- rectification will be applicable as per tender conditions, mentioned under ITB clause

28.3 (e) (ii) under Comprehensive Annual maintenance contract. It may include but not limited to forfeiting of performance security & taking legal proceeding deemed fit as per applicable Indian Law.

14.8     The Purchaser/Consignee reserve the rights to enter into Annual Maintenance Contract

/ Comprehensive Maintenance Contract between Consignee and the Supplier for the period as mentioned in TE document.

14.9  The supplier along with its Indian Agent and the Service Provider/CMC provider (if applicable)  shall  ensure  continued  supply  of  the  spare  parts  for  the  machines  and equipments supplied by them to the purchaser for 10 years from the  date of installation and handing over.

14.10 The Supplier along with its Indian Agent and the Service Provider/CMC Provider (if applicable) shall always accord most favored client status to the Purchaser vis-à-vis it’s other Clients/Purchasers of its equipment/machines/goods etc. and shall always give the most competitive price for its machines /equipment supplied to the Purchaser/Consignee.

15. PAYMENT TERMS

15.1	The method and conditions of payment to be made to the supplier under the contract may

15.2     be specified in the Special Conditions of Contract &/Notification of Award.

15.3	The Supplier’s request(s) for payment shall be made to the Purchaser in writing, accompanied by an invoice describing, as appropriate, the Goods delivered and Services performed, and by documents submitted pursuant to GCC Clause 8, and upon fulfillment of other obligations stipulated in the Contract.

15.4     Payments shall be made promptly by the Purchaser, but in no case later than sixty (60)

days after submission of an invoice or claim by the Supplier.

[Hint:   The actual payment conditions for new products or procurements having installation and CMC/AMC services may be decided on case to case  basis  and  incorporated in  special conditions of the contract if required]

15.5     No payment will be made for goods rejected at the site on testing

15.6     Payment for goods shall be made in Indian Rupees as follows:

a) No advance payment is payable.

b) 50% payment will be made against supply of items at the respective sites and after submission  of  satisfactory inspection report from the consignee  and 50 % payment will be made against  certification  from  the  consignee  in  the  format  provided  in schedule VI and  after verification of installation / supply  by  purchaser  (BMSICL, Patna.) or its nominated agency/person.

16. PRICES

16.1	Prices charged by the supplier for goods delivered and services performed under the contract shall not be higher than the prices quoted by the Supplier in his Bid.

16.2	In the case of revision of Statutory Levies/Taxes during the finalization period of tender, the Purchaser reserves the right to ask for reduction in the prices.

a)  Prices once fixed will remain valid during the schedule delivery period. In case of Increase and decrease of Taxes and other statutory duties the affect in price will be decided by BMSICL. The decision of Tender Inviting Authority will be final for the same.

b)  Any   increase in taxes and other statutory  duties /levies  from  the  date  of submission  of  bid  till  the  end  of  the delivery period will be paid extra

however. Any increase in the taxes and other statutory duties / levis after the expiry of the delivery date shall be to the supplier’s account. However, benefit or any decrease in the taxes/duties shall be passed on to the purchaser by the supplier, after the expiry of the delivery date shall be to the suppliers’s account. However, benefit or any decrease in the taxes/duties shall be passed on to the purchaser by the supplier.

17. CHANGE ORDERS

17.1	The purchaser may, at any time, by a written order given to a supplier, make changes within the general scope of the contract in any one or more of the following:

(a)   drawings, designs or specifications, where Goods to be supplied under the contract are to be specifically manufactured for the Purchaser;

(b)   the method of transportation or packing; (c)    the place of delivery; or

(d)   the services to be provided by the supplier.

17.2	If any such change causes an increase or decrease in the cost of, or the time required for the execution of the contract an equitable adjustment shall be made in the contract price or delivery schedule, or both, and the contract shall accordingly be amended. Any proposal by the supplier for adjustment under this clause must be made within thirty days from the date of the receipt of the change in order.

18. SUBCONTRACTS

The Supplier shall notify the Purchaser in writing of all subcontracts awarded under this contract if not already specified in his bid. Such notification, in his original bid or later shall not relieve the supplier from any liability or obligation under the Contract.

19. DELAYS IN THE SUPPLIER’S PERFORMANCE

19.1	Delivery of the Goods and performance of the services shall be made by the Supplier in accordance with the time schedule specified by the purchaser in its purchase order. In case the supply is not completed in the stipulated delivery period, as indicated in the Purchase Order, purchaser reserves the right either to short close/cancel this purchase order and/or recover liquidated damage charges. The cancellation/short closing of the order shall be at the risk and responsibility of the supplier and purchaser reserves the right to purchase balance unsupplied item at the risk and cost of the defaulting vendors.

19.2	Delay by the Supplier in the performance of its delivery obligations shall render the Supplier liable to any or all of the following sanctions: forfeiture of its performance security, imposition of liquidated damages and/or termination of the contract for default.

19.3 	If at any time during the performance of the contract, the supplier encounters condition impending timely delivery of the goods and performance of service, the Supplier shall promptly notify to the Purchaser in writing the fact of the delay, it’s likely duration and its cause(s). As soon as practicable after receipt of the supplier’s notice, the Purchaser shall evaluate the situation and may at its discretion extend the period for performance of the contract (by not more than 6 weeks) subject to furnishing of additional performance security by the supplier @ 5% of the total value of the Purchase Order or a part thereof, in case of part execution for the period of up to four weeks after such extended schedule of delivery. Non-performance of the contract even during this extended period may lead to forfeiture of such additional Performance Guarantee.

[Hint: Each case of delivery extension shall have to be examined a fresh vis-à-vis the prevailing market prices]

19.4	If supplier fails to perform its contractual obligations, pursuant to GCC Clause 19.3 above, the purchaser may consider debarring the firm for the period of 1-5 years for participation in future invitation of bids. The period of debar, as stated above, shall be at the sole discretion of the Purchaser.

20. LIQUIDATED DAMAGES

20.1	The date of delivery of the goods stipulated in the acceptance of the tender should be deemed to be the essence of the contract and delivery must be completed no later than the dates specified therein. Extension will not be given except in exceptional circumstances. Should, however, deliveries be made after expiry of the contracted delivery period, without prior concurrence of the purchaser and be accepted by the consignee, such delivery will not deprive the purchaser of his right to recover liquidated damage under GCC Clause 20.2 below.

20.2	Should the supplier fails to deliver the store or any consignment thereof within the period prescribed for delivery, the purchaser shall be entitled to recover 0.5 % of the value of the delayed supply for each week of delay or part thereof. In the case of package supply where the delayed portion of the supply materially hampers installation and commissioning of the systems, L/D charges shall be levied as above on the total value of the concerned package of the Purchase Order. Quantum of liquidated damages assessed and levied by the purchaser shall be final and not challengeable by the supplier. However, when supply is made within 21 days of QA clearance in the extended delivery period, the consignee may accept the stores and in such cases the LD shall be levied up to the date of QA clearance.

21. FORCE MAJEURE

21.1	If, at any time, during the continuance of this contract, the performance in whole or in part by either party of any obligation under this contract is prevented or delayed by reasons of any war or hostility, acts of the public enemy, civil commotion, sabotage , fires, floods, explosions, epidemics, quarantine restrictions, strikes, lockouts or act of God (hereinafter referred to as events) provided notice of happenings of any such eventuality is given by either party to the other within 21 days from the date of occurrence thereof, neither party shall by reason of such event be entitled to terminate this contract nor shall either party have any claim for damages against other in respect of such non-performance or delay in performance, and deliveries under the contract shall be resumed as soon as practicable after such an event come to an end or cease to exist, and the decision of the Purchaser as to whether the deliveries have been so resumed or not shall be final and conclusive. Further that if the performance in whole or part of any obligation under this contract is prevented or delayed by reasons of any such event for a period exceeding 60 days, either party may, at its option, terminate the contract.

21.2	Provided, also that if the contract is terminated under this clause, the Purchaser shall be at liberty to take over from the Supplier at a price to be fixed by the purchaser, which shall be final, all unused, undamaged and acceptable materials, bought out components and stores in course of manufacture which may be in possession of the Supplier at the time of such termination or such portion thereof as the purchaser may deem fit, except such materials, bought out components and stores as the Supplier may with the concurrence of the purchaser elect to retain.

22. TERMINATION FOR DEFAULT

22.1	The Purchaser may, without prejudice to any other remedy for breach of contract, by written notice of default, sent to the supplier, terminate this contract in whole or in part

a)    If the supplier fails to deliver any or all of the goods within the time period(s) specified in the  contract,  or  any extension  thereof  granted  by the  purchaser  pursuant  to  GCC Clause19;

b)      if the supplier fails to perform any other obligation(s) under the Contract; and

c)	if the supplier, in either of the above circumstances, does not remedy his failure within a period of 15 days (or such longer period as the purchaser may authorize in writing) after receipt of the default notice from the purchaser.

d)	If the Supplier, in the judgment of the Purchaser, has engaged in corrupt and fraudulent practices in competing for executing the Contract, pursuant to ITB Clause 2.

22.2     In the event the purchaser terminates the contract in whole or in part pursuant to GCC

Clause22.1. the purchaser may procure, upon such terms and in such manner as it deems appropriate, goods similar to those undelivered and the supplier shall be liable to the Purchaser for any excess cost for such similar goods. However, the supplier shall continue the performance of the contract to the extent not terminated.

22.3	In the event, any sums found due to the Purchaser / Government under or by virtue of the fulfillment of contractual obligations, these shall be recoverable from the Supplier and his / its properties, movable and immovable, under the provisions of the Revenue Recovery Act, for the time being in force as tough as they are arrears of land revenue or in any manner and within such time as the Purchaser / Government may deem fit. Any sum of money due and payable to the Supplier from Government / Purchaser may be adjusted against sum of money due to the Supplier under any other contract.

23. TERMINATION FOR INSOLVENCY

The Purchaser may at any time terminate the Contract by giving written notice to the Supplier,

without compensation to the supplier. If the supplier becomes bankrupt or otherwise insolvent as declared by the competent court provided that such termination will not prejudice or affect any right of action or remedy which has accrued or will accrue thereafter to the purchaser.

24. TERMINATION FOR CONVENIENCE

24.1	The Purchaser, by written notice sent to the Supplier, may terminate the Contract, in whole or in part, at any time for its convenience. The notice of termination shall specify that termination is for the Purchaser’s convenience, the extent to which performance of the Supplier under the Contract is terminated, and the date upon which such termination becomes effective.

24.2	The Goods that are complete and ready for shipment within thirty (30) days after the Supplier’s receipt of notice of termination shall be accepted by the Purchaser at the Contract terms and prices. For the remaining Goods, the Purchaser may elect:

(a)  to have any portion completed and delivered at the Contract terms and prices; and/or.

(b) to cancel  the  remainder  and  pay  to  the  Supplier  an  agreed  amount  for partially completed Goods and Services and for materials and parts previously procured by the Supplier.

25. SETTLEMENT OF DISPUTES

25.1	If any dispute or difference of any kind whatsoever shall arise between the Purchaser and the Supplier in connection with or arising out of the Contract, the parties shall make every effort to resolve amicably such dispute or difference by mutual consultation.

25.2	If the parties even after 30 days fail to resolve their dispute or difference by such mutual consultation, then either the Purchaser or the Supplier may give notice to the other party of its intention to commence arbitration, with respect to such disputes or differences arising out of the terms and conditions of the contract In accordance with the procedure prescribed under Bihar Public Works Contracts Disputed Arbitration Tribunal Act 2008.

25.3       The arbitration shall be in accordance with the procedure prescribed under the Bihar

Public Works Contracts Disputed Arbitration Tribunal Act 2008.

25.4       Notwithstanding any reference to arbitration herein,

(a)  the parties shall continue to perform their respective obligations under the Contract unless they otherwise agree; and

(b) the Purchaser shall pay the Supplier any monies due the Supplier.

25.5	The contract shall be governed by and interpreted in accordance with the laws of India from the time being in force. All disputes arising out of this tender will be subject to jurisdiction of courts of law in Patna

26. LIMITATION OF LIABILITY

26.1	Except in cases of criminal negligence or willful misconduct, and in the case of infringement pursuant to GCC Clause 4,

- the Supplier shall not be liable to the Purchaser, whether in contract, or otherwise, for any indirect or consequential loss or damage, loss of use, loss of production, or loss of profits or interest costs, provided that this exclusion shall not apply to any obligation of the Supplier to pay liquidated damages to the Purchaser and

- the aggregate liability of the Supplier to the Purchaser, whether under the Contract, in tort or otherwise, shall not exceed the total Contract Price, provided that this limitation shall not apply to the cost of repairing or replacing defective items/goods

27. GOVERNING LANGUAGE

27.1	The Contract shall be written in English language. All correspondence and other documents pertaining to the Contract that are exchanged by the parties shall be written in the Hindi / English language.

28. APPLICABLE LAW

28.1       The Contract shall be interpreted in accordance with the laws of Union of India.

29. NOTICES

29.1	Any notice given by one party to the other pursuant to this Contract shall be sent to the other party in writing or by cable, telex, or facsimile and confirmed in writing to the other party’s address.

29.2	A notice shall be effective when delivered or on the notice’s effective date, whichever is later

30. Taxes and Duties

30.1	The Supplier shall be entirely responsible for all taxes, duties, octroy, road permits, license fees, etc., incurred until delivery of the contracted Goods to the Purchaser.

SECTION III- SPECIAL CONDITIONS OF CONTRACT

SPECIAL CONDITIONS OF CONTRACT

1.   The special conditions of contract shall supplement the ‘Instructions to the Bidders’ as contained in Section I & “General Conditions of the Contract” as contained in Section II and wherever there is a conflict, the provisions herein shall prevail over those in Section I and Section II.

2.   No Exemption from payment of EMD is permitted except in case of NSIC (National Small Industries Corporation rule as per Sankalp 675 (1), Dated 09/09/2013 of Govt. of Bihar) registered small scale industries. No exemption in tender fee will be allowed in any case.

3.   The tender will be processed through E –tender mode only, So tender should be submitted in following manner:-

i)  Tender Fee (Online only) & EMD (only in the form of Bank Guarantee which to be submitted offline).

(ii) Tender Processing Fee (Only Online) (iii) Technical Bid (Only Online)

(iv) Price Bid (Only Online).

Bidders are requested to submit EMD of quoted item of tender (Offline in the form of BG) and sample of quoted item. Bidders are requested not to submit hard copy of tender document. In case the hard copy of financial bid is submitted the tender shall be straightway rejected.

4.   If asked by the purchaser, then the tenderer must do demonstration of equipment. In the demonstration tenderer has to show the clause-by-clause compliance of technical specification. Demonstration of quoted items will be at Govt. Medical College & Hospital, Mahua (vaishali) and Govt. Medical College & Hospital, Madhubani, Bihar

5.   L1 shall be decided on the basis of the overall lowest rate to be arrived at by summing up of all the individual rates quoted for each main item as detailed in the BOQ/Price bid/Financial bid. It is pertinent to mention here that the individual rate of any main item shall be computed on the basis of unit price of the said item (equipment) along with the cost of consumables/reagents required for conducting a total number of one thousand tests by the said item (equipment).

6.   The rate quoted by the bidder for the consumables/reagents shall be valid for the period of

03 (Three) years.

Section IV- SCHEDULE OF REQUIREMENTS

SCHEDULE OF REQUIREMENTS

Note:  The successful firm agrees to the proposal of BMSICL (The purchaser) that such equipment for which rate contract already exits with BMSICL and is lower than the quoted rate of successful firm (The supplier) on the date of issue of Letter of Intent (LOI) to Successful firm L1 bidder, will not be supplied by the Firm and such equipment will be procured by BMSICL

| Schedule No. | Brief Description of Goods and Services on TURNKEY basis | Qty./No. | Delivery Schedule (in days) | Earnest Money Deposit (EMD) in Indian Rupees |
| --- | --- | --- | --- | --- |
| 1 | Supply, Installation & Commissioning of Medical Equipment on Turnkey basis for Govt. Medical College & Hospital, Mahua (vaishali) and Govt. Medical College & Hospital, Madhubani, Bihar | (As mentioned in section –IV) of the bid document | As per GCC Clause -8 | 1,26,00,000/- (One Crores Twenty-Six Lakhs only) |

List of Main Item

| Sr No. | Departments | NAME OF THE ITEMS | Qty |
| --- | --- | --- | --- |
| 1 | ANAESTHESIOLOGY | Anesthesia machine and accessories like laryngoscope, all size endotracheal tubes, nasal and oral airways, Magill’s forceps, | 8 |
| 2 | ANAESTHESIOLOGY | Electrical Suction apparatus | 8 |
| 3 | ANAESTHESIOLOGY | Fibre optic bronchoscope | 1 |
| 4 | ANAESTHESIOLOGY | USG machine (Desirable) | 1 |
| 5 | ANAESTHESIOLOGY | PNS (Desirable) | 1 |
| 6 | ANAESTHESIOLOGY | a) Adult Manikin | 1 |
| 7 | ANAESTHESIOLOGY | b) Pediatric Manikin | 1 |
| 8 | ANAESTHESIOLOGY | Oxygen therapy unit | 5 |
| 9 | ANAESTHESIOLOGY | Provision for resuscitation equipment and CPR Algorithms | 1 |
| 10 | ANAESTHESIOLOGY | Side lab for emergency investigations (Optional) | 1 |
| 11 | ANAESTHESIOLOGY | TOF monitor (PNS) (Optional) | 1 |
| 12 | ANAESTHESIOLOGY | Patients controlled analgesia system (portable) | 1 |
| 13 | ANAESTHESIOLOGY | OT Table fluoroscopy compatible | 1 |
| 14 | ANAESTHESIOLOGY | Nerve locator | 1 |
| 15 | ANAESTHESIOLOGY | Anodyne Machine (Optional) (For diabetic foot care) | 1 |
| 16 | ANAESTHESIOLOGY | Transcutaneous Electric Nerve Stimulating Machine (optional) | 1 |
| 17 | ANAESTHESIOLOGY | Anesthetic machine with resuscitation equipment | 1 |
| 18 | ANAESTHESIOLOGY | Nerve stimulator | 1 |
| 19 | ANAESTHESIOLOGY | Patient Examination Table | 4 |
| 20 | ANAESTHESIOLOGY | Suction Machine | 5 |
| 21 | ANATOMY | Drill machine | 3 |
| 22 | ANATOMY | Hand saw, preferably metal | 2 |
| 23 | ANATOMY | Band saw for sectioning body and limbs | 2 |
| 24 | ANATOMY | Storage tank to hold 10 cadavers, static/movable, durable tank with input and output facility with lid | 3 |
| 25 | ANATOMY | Dissecting instruments for cadaveric dissection | 10 |
| 26 | ANATOMY | Meat cutting machine for thin body sections (trans and vertical) for gross anatomy sectional study | 2 |
| 27 | ANATOMY | Embalming Machine | 1 |
| 28 | ANATOMY | Dissection microscope | 2 |
| 29 | ANATOMY | Microtomes, rotary | 2 |
| 30 | ANATOMY | Microtomes, Sledge, large cutting | 1 |
| 31 | ANATOMY | Paraffin Embedding bath | 2 |
| 32 | ANATOMY | Hot plate for Flattening Section | 2 |
| 33 | BIOCHEMISTRY | Analytical Balance : upto 200g/1gm increment | 1 |
| 34 | BIOCHEMISTRY | Semi autoanalyser | 3 |
| 35 | BIOCHEMISTRY | Constant temperature water bath Tank Capacity: (Temperature range 5 to 80 degree Celsius) | 2 |
| 36 | BIOCHEMISTRY | Complete Chromatographic Unit for paper & TLC | 2 |
| 37 | BIOCHEMISTRY | Centrifuge ≥ 8 tubes | 4 |
| 38 | BIOCHEMISTRY | Complete Electrophoresis apparatus with power supply (Paper, PAGE, agarose) | 1 |
| 39 | BIOCHEMISTRY | Densitometer with computer | 2 |
| 40 | BIOCHEMISTRY | Fume cupboard | 1 |
| 41 | BIOCHEMISTRY | Digital Analytical Balance | 2 |
| 42 | BIOCHEMISTRY | Balance Micro | 1 |
| 43 | BIOCHEMISTRY | Spectrophotometer | 1 |
| 44 | BIOCHEMISTRY | ELISA (Demonstration) | 1 |
| 45 | BIOCHEMISTRY | PCR Machine | 1 |
| 46 | BIOCHEMISTRY | Auto Analyzer (Either in the institution or elsewhere on a visit) | 1 |
| 47 | BIOCHEMISTRY | Hot air oven (More than 200 litres) | 3 |
| 48 | BIOCHEMISTRY | Water baths | 3 |
| 49 | BIOCHEMISTRY | Incubator 37oC | 2 |
| 50 | COMMUNITY MEDICINE | Centrifuge clinical | 1 |
| 51 | COMMUNITY MEDICINE | Baby weighing machine | 2 |
| 52 | COMMUNITY MEDICINE | Salters Baby weighing machine | 2 |
| 53 | COMMUNITY MEDICINE | Otoscope | 1 |
| 54 | COMMUNITY MEDICINE | Ophthalmoscope | 1 |
| 55 | COMMUNITY MEDICINE | Incubator, electric (can be procured from Microbiology) | 1 |
| 56 | Dentistry | DENTAL X-RAY-INTRA ORAL WITH RVG | 2 |
| 57 | Dentistry | DENTAL X-RAY –EXTRA ORAL (O.P.G) –DIGITAL | 1 |
| 58 | Dentistry | MINI AUTOCLAVE (VACUUM TYPE) | 1 |
| 59 | Dentistry | DENTAL INSTRUMENTS | 2 |
| 60 | DERMATOLOGY, VENEROLOGY & LEPROSY | Hyfrecator/ Electrosurgical instrument | 1 |
| 61 | DERMATOLOGY, VENEROLOGY & LEPROSY | Cryotherapy with liquid Nitrogen | 1 |
| 62 | DERMATOLOGY, VENEROLOGY & LEPROSY | Iontophoresis machine | 1 |
| 63 | DERMATOLOGY, VENEROLOGY & LEPROSY | Wood’s lamp | 2 |
| 64 | EMERGENCY MEDICINE | Emergency trolley-cum-beds (Should have facilitics for propping up the patient along with railings and wheels) | 24 |
| 65 | EMERGENCY MEDICINE | ICU ventilators (shuould be universal ventilators with facility to ventilate neonates and children also) | 3 |
| 66 | EMERGENCY MEDICINE | HHHFNC (heated humidified high-flow nasal cannula) with circuits and interfaces for all age groups (neonates, infants, children, adolescents and adults) | 2 |
| 67 | EMERGENCY MEDICINE | Artificial self-inflating bag (Ambu bag) ( adult, paediatric, infant and neonatal) | 2 |
| 68 | EMERGENCY MEDICINE | Point -of care laboratory for quantitative estimatin of cardiac enzymes, ABG and electrolytes | 1 |
| 69 | EMERGENCY MEDICINE | Ultrasonic nebulizers | 4 |
| 70 | EMERGENCY MEDICINE | ACLS, BLS and Airways mannequins (child and adult) | 1 |
| 71 | EMERGENCY MEDICINE | Suturing mannequin | 1 |
| 72 | EMERGENCY MEDICINE | Blood and Fluid warmer | 1 |
| 73 | EMERGENCY MEDICINE | Electric Warming blanket | 1 |
| 74 | EMERGENCY MEDICINE | Rapid autoclave machine | 1 |
| 75 | EMERGENCY MEDICINE | Labor cot | 1 |
| 76 | EMERGENCY MEDICINE | Portable suction machines | 4 |
| 77 | EMERGENCY MEDICINE | Oxygen Concentrator | 5 |
| 78 | FORENSIC MEDICINE & TOXICOLOGY | Digital Spectrophotometer | 1 |
| 79 | FORENSIC MEDICINE & TOXICOLOGY | Chemical Balance | 1 |
| 80 | FORENSIC MEDICINE & TOXICOLOGY | Distillation Plant | 1 |
| 81 | FORENSIC MEDICINE & TOXICOLOGY | Centrifuge | 1 |
| 82 | FORENSIC MEDICINE & TOXICOLOGY | Slide Warming Table | 1 |
| 83 | FORENSIC MEDICINE & TOXICOLOGY | Hot Plate | 1 |
| 84 | FORENSIC MEDICINE & TOXICOLOGY | Spectroscopic Lens With Adjustable Slit | 2 |
| 85 | FORENSIC MEDICINE & TOXICOLOGY | Cold Storage For Dead Bodies | 2 |
| 86 | FORENSIC MEDICINE & TOXICOLOGY | Weighing Machine For Dead Bodies | 1 |
| 87 | FORENSIC MEDICINE & TOXICOLOGY | Autopsy Tables | 2 |
| 88 | FORENSIC MEDICINE & TOXICOLOGY | Autopsy Saw With Accessories | 2 |
| 89 | FORENSIC MEDICINE & TOXICOLOGY | Weighing Machine For Organs | 2 |
| 90 | FORENSIC MEDICINE & TOXICOLOGY | Weighing Machine For Fetus | 2 |
| 91 | GENERAL MEDICINE | Ophthalmoscope | 5 |
| 92 | GENERAL MEDICINE | Flexible Video End viewing Oesophago-Gastroduodenoscope (optional) | 1 |
| 93 | GENERAL MEDICINE | Flexible Video Colonoscope (optional) | 1 |
| 94 | GENERAL MEDICINE | Proctoscope | 2 |
| 95 | GENERAL MEDICINE | Fiber optic bronchoscope (May be shared with TB&CD) | 1 |
| 96 | GENERAL MEDICINE | Spirometer (Ordinary) | 4 |
| 97 | GENERAL MEDICINE | Equipment for Cardiac pacing | 1 |
| 98 | GENERAL MEDICINE | Echocardiography machine | 1 |
| 99 | GENERAL MEDICINE | Hemodialysis machine | 3 |
| 100 | GENERAL MEDICINE | RO Plant for haemodialysis | 1 |
| 101 | GENERAL MEDICINE | Arterial blood gas analyzer | 1 |
| 102 | GENERAL MEDICINE | EMG and nerve conduction velocity machine | 1 |
| 103 | GENERAL MEDICINE | Invasive Mechanical Ventilator | 8 |
| 104 | GENERAL MEDICINE | Non-Invasive mechanical ventilator (Maybe shared with TB and CD) | 5 |
| 105 | GENERAL MEDICINE | Nebulizer | 12 |
| 106 | GENERAL MEDICINE | Upper GI endoscope (Optional) | 1 |
| 107 | GENERAL MEDICINE | Sigmoidoscope (optional) | 1 |
| 108 | GENERAL MEDICINE | Colonoscope (optional) | 1 |
| 109 | GENERAL MEDICINE | *Haemocytometer | 2 |
| 110 | GENERAL MEDICINE | *Light Microscope | 1 |
| 111 | GENERAL MEDICINE | Tread mill test machine | 1 |
| 112 | GENERAL MEDICINE | Portable Suction Machine | 5 |
| 113 | GENERAL SURGERY | Proctoscope | 2 |
| 114 | GENERAL SURGERY | Operation Theatre Table | 1 |
| 115 | GENERAL SURGERY | Pedestal lights | 2 |
| 116 | GENERAL SURGERY | Anesthesia Equipment | 1 |
| 117 | GENERAL SURGERY | Autoclave | 2 |
| 118 | GENERAL SURGERY | Non invasive Multi Para Monitors | 4 |
| 119 | GENERAL SURGERY | Cystoscope&Resectoscope | 1 |
| 120 | GENERAL SURGERY | Flexible Video Colonoscope | 1 |
| 121 | GENERAL SURGERY | Flexible Video Side viewing Gastroduodenoscope for ERCP Video Colonoscope VIDEO PROCESSOR WITH LIGHT SOURCE COMPATIBLE WITH THE ABOVE MENTIONED VIDEOSCOPE. High Definition Lcd Colour Monitor ( Flat Panel) | 1 |
| 122 | GENERAL SURGERY | Flexible Video End viewing Oesophago-Gastroduodenoscope | 1 |
| 123 | GENERAL SURGERY | Flexible Video Sigmoidoscope | 1 |
| 124 | GENERAL SURGERY | Flexible Video Bronchoscope | 1 |
| 125 | GENERAL SURGERY | Harmonic Scalpel | 1 |
| 126 | GENERAL SURGERY | Suction machine | 7 |
| 127 | MICROBIOLOGY | Microscopes | 100 |
| 128 | MICROBIOLOGY | Balance Electronic Digital | 1 |
| 129 | MICROBIOLOGY | BOD Incubator | 1 |
| 130 | MICROBIOLOGY | Centrifuge | 3 |
| 131 | MICROBIOLOGY | CO2 Incubator/Candle Jar | 2 |
| 132 | MICROBIOLOGY | Deep Freeze -20° C & Deep Freezer | 1 |
| 133 | MICROBIOLOGY | Distilled water Plant | 1 |
| 134 | MICROBIOLOGY | Elisa Reader | 1 |
| 135 | MICROBIOLOGY | Elisa Washer | 1 |
| 136 | MICROBIOLOGY | Lab Refrigerator (minimum 400 litres) | 3 |
| 137 | MICROBIOLOGY | Laminar flow | 1 |
| 138 | MICROBIOLOGY | Microscope with universal condenser containing oil immersion, Bright field, Phase Contrast & Dark ground | 1 |
| 139 | MICROBIOLOGY | Automated Blood Culture System | 1 |
| 140 | MICROBIOLOGY | Colony Counter | 1 |
| 141 | MICROBIOLOGY | Antibiotic Discs for Antibiotic susceptibility testing | 250 |
| 142 | MICROBIOLOGY | Antibiotic zone scale | 5 |
| 143 | MICROBIOLOGY | Antisera-Salmonella | 5 |
| 144 | MICROBIOLOGY | Antisera-Shigelladysenteriae | 5 |
| 145 | MICROBIOLOGY | Antisera-Shigellaflexnari | 5 |
| 146 | MICROBIOLOGY | Antisera-Shiegellasonnie | 5 |
| 147 | MICROBIOLOGY | Antisera-Vibrio cholerae | 5 |
| 148 | MICROBIOLOGY | Real-time PCR machine calibrated for the fluorophore dyes with UPS | 1 |
| 149 | MICROBIOLOGY | Refrigerated Microcentrifuge | 1 |
| 150 | MICROBIOLOGY | All other consumbles and kits as required for Virology studes. | 5 |
| 151 | MICROBIOLOGY | Autoclave | 2 |
| 152 | MICROBIOLOGY | Biosafety Cabinet Type - 2A | 3 |
| 153 | MICROBIOLOGY | Hot Air Oven | 2 |
| 154 | MICROBIOLOGY | Incubator | 2 |
| 155 | MICROBIOLOGY | Water bath with variable temperature | 2 |
| 156 | MICROBIOLOGY | minus 80 degree deep freezer with UPS | 1 |
| 157 | OBSTETRICS & GYNAECOLOGY | Cryo/electro cautery apparatus | 1 |
| 158 | OBSTETRICS & GYNAECOLOGY | NST machine | 2 |
| 159 | OBSTETRICS & GYNAECOLOGY | Cystoscope & Resectoscope | 1 |
| 160 | OBSTETRICS & GYNAECOLOGY | High End suction machine | 6 |
| 161 | OBSTETRICS & GYNAECOLOGY | Vacuum Extractor and suction machine | 1 |
| 162 | OBSTETRICS & GYNAECOLOGY | Fetal Monitor for Antepartum Surveillance | 1 |
| 163 | OBSTETRICS & GYNAECOLOGY | D & C Set with MTP Set | 5 |
| 164 | OBSTETRICS & GYNAECOLOGY | Caessarian Set | 5 |
| 165 | OBSTETRICS & GYNAECOLOGY | Hysterectomy Set (Abdominal & Vaginal) | 5 |
| 166 | OBSTETRICS & GYNAECOLOGY | Suction machine | 4 |
| 167 | OPHTHALMOLOGY | Trial set with trial frame both for adult and children | 2 |
| 168 | OPHTHALMOLOGY | Automated Perimeter | 1 |
| 169 | OPHTHALMOLOGY | Ophthalmoscope (Direct) | 5 |
| 170 | OPHTHALMOLOGY | Streak Retinoscope | 4 |
| 171 | OPHTHALMOLOGY | Indirect Ophthalmoscope | 5 |
| 172 | OPHTHALMOLOGY | Keratometer | 1 |
| 173 | OPHTHALMOLOGY | Gonioscope | 2 |
| 174 | OPHTHALMOLOGY | Schiotz’s tonometer | 4 |
| 175 | OPHTHALMOLOGY | Cryo Unit | 1 |
| 176 | OPHTHALMOLOGY | Operation Theatre Table | 2 |
| 177 | OPHTHALMOLOGY | Operation theatres should have resuscitation equipment like Anesthesia machine and accessories like laryngoscope, all size endotracheal tubes, nasal and oral airways, Magill’s forceps ,Mechanical ventilator - separate or with anesthesia machine, LMA / PLMA of all sizes, Electrical suction apparatus | 1 |
| 178 | ORTHOPAEDICS | Set for Hip Replacement | 1 |
| 179 | ORTHOPAEDICS | Set for Knee replacement | 1 |
| 180 | ORTHOPAEDICS | Interlock nailing sets | 1 |
| 181 | ORTHOPAEDICS | fracture reduction OT table | 2 |
| 182 | ORTHOPAEDICS | Simple OT tables | 1 |
| 183 | ORTHOPAEDICS | Pneumatic drill and reamer | 1 |
| 184 | ORTHOPAEDICS | Electrical drill and reamer set | 1 |
| 185 | ORTHOPAEDICS | Arthoscope | 1 |
| 186 | ORTHOPAEDICS | Physiotherapy and occupational Therapy equipment sets | 1 |
| 187 | ORTHOPAEDICS | Electrical Suction apparatus | 2 |
| 188 | OTORHINOLARYNGOLOGY | Suction apparatus | 1 |
| 189 | OTORHINOLARYNGOLOGY | Otoscope with halogen bulb, rechargeable battery and Siegle attachment | 1 |
| 190 | OTORHINOLARYNGOLOGY | Head Light With LED/Halogen lamp (Battery Operated) | 1 |
| 191 | OTORHINOLARYNGOLOGY | Rigid nasal endoscope 0 degree,4 mm and 2.7 mm size | 1 |
| 192 | OTORHINOLARYNGOLOGY | Rigid nasal endoscope 30 degree, 4 mm and 2.7 mm size | 1 |
| 193 | OTORHINOLARYNGOLOGY | Light source and light cable | 1 |
| 194 | OTORHINOLARYNGOLOGY | Endoscopic Camera with suitable display with recording & archiving facility. (with High Definition Camera) | 1 |
| 195 | OTORHINOLARYNGOLOGY | Light source | 1 |
| 196 | OTORHINOLARYNGOLOGY | Light cable | 1 |
| 197 | OTORHINOLARYNGOLOGY | Bronchoscopes different sizes for different age groups | 3 |
| 198 | OTORHINOLARYNGOLOGY | Light source and cable | 1 |
| 199 | OTORHINOLARYNGOLOGY | (d) Nasal endoscopy trolley with 0 degree & 30 degree 4mm endoscope with light source, cable , monitor & camera | 1 |
| 200 | OTORHINOLARYNGOLOGY | Sterilizer | 1 |
| 201 | OTORHINOLARYNGOLOGY | Spot Mobile Light | 1 |
| 202 | OTORHINOLARYNGOLOGY | Basic OT table | 1 |
| 203 | OTORHINOLARYNGOLOGY | Puretone audiometer | 1 |
| 204 | OTORHINOLARYNGOLOGY | Brainstem evoked response audiometer with ASSR | 1 |
| 205 | OTORHINOLARYNGOLOGY | OAE Impedance audiometer (With sound treated air-conditioned room for audiometry) | 1 |
| 206 | OTORHINOLARYNGOLOGY | Flexible nasopharyngolaryngoscope | 1 |
| 207 | OTORHINOLARYNGOLOGY | Electronystagmograph(Optional) | 1 |
| 208 | OTORHINOLARYNGOLOGY | Brainstem evoked response audiometer | 1 |
| 209 | PAEDIATRICS | Digital Weighing machine(for New born) | 2 |
| 210 | PAEDIATRICS | Ophthalmoscope | 2 |
| 211 | PAEDIATRICS | Fumigation Machine | 1 |
| 212 | PAEDIATRICS | Ambu bag | 6 |
| 213 | PAEDIATRICS | Mechanical Ventilator (neonatal and child) | 2 |
| 214 | PAEDIATRICS | Suction machine | 2 |
| 215 | PAEDIATRICS | Transport Incubator | 1 |
| 216 | PATHOLOGY | Manual Rotary Microtome | 1 |
| 217 | PATHOLOGY | Automated Rotary Microtome | 1 |
| 218 | PATHOLOGY | Cryostat | 1 |
| 219 | PATHOLOGY | Hot plate | 2 |
| 220 | PATHOLOGY | Paraffin embedding bath | 1 |
| 221 | PATHOLOGY | Heated Paraffin Embedding Module | 1 |
| 222 | PATHOLOGY | Cold Plate for Modular Tissue Embedding System | 1 |
| 223 | PATHOLOGY | Automated Tissue Processor –Histokinette | 1 |
| 224 | PATHOLOGY | Ultrapure water solutions - Distilled water plant | 1 |
| 225 | PATHOLOGY | Centrifuge machine | 5 |
| 226 | PATHOLOGY | Fully Automated high throughput Multi-Stainer Workstation | 1 |
| 227 | PATHOLOGY | Fully Automated Embedding System (Heated embedding module & cold plate) | 1 |
| 228 | PATHOLOGY | Fully Automated Flexible Coverslipping Workstation | 1 |
| 229 | PATHOLOGY | Standalone paraffin dispensing module cold plate holding more than 100 cassettes | 1 |
| 230 | PATHOLOGY | Stand alone cold plate | 1 |
| 231 | PATHOLOGY | Single Pan Digital Balance, Chemical | 2 |
| 232 | PATHOLOGY | Balance, chemical with weights | 2 |
| 233 | PATHOLOGY | For Diagnostic & Research Work - Trinocular head Microscope with Bright field, Dark field, Fluorescent & Polarizing Facility, high end Apochromatic lenses with Camera with HDMI Multi output camera Minimum 5MP with Projector & Ultra HD TV > 52 inches & Screen including Software Capable of Brightfield& Immunofluorescence Photography with connectivity to projector & LED TV (At least 55 inches Ultra HD) | 1 |
| 234 | PATHOLOGY | Penta Head Microscope with High end Optics with HDMI Multi output Photographic camera (> 5 MP) including Software | 1 |
| 235 | PATHOLOGY | Deca Head Microscope with High end Optics with HDMI Multi output Photographic camera (> 5 MP) including Software | 1 |
| 236 | PATHOLOGY | Grossing Station - Stainless steel, with Control panel, air filtration system, Track mounted adjustable computer arm with articulation, LED lights that are color and intensity, Dedicated USB ports for camera control and data transfer adjustable, Integrated pathology camera system, Instrument Set (High quality) Height Adjustable Stainless Steel Chairs With Split AC of appropriate capacity. | 1 |
| 237 | PATHOLOGY | Fully Automated Immuno-histo-chemistry Setup with Continuous supply of Important Antibodies, Lymphoma Panel etc. | 1 |
| 238 | PATHOLOGY | Five part Fully Automated Cell Counter | 1 |
| 239 | PATHOLOGY | Three Part Fully Automated Cell Counter | 1 |
| 240 | PATHOLOGY | Coagulometer (Fully automated) | 1 |
| 241 | PATHOLOGY | Autoclave | 2 |
| 242 | PATHOLOGY | Water bath | 2 |
| 243 | PATHOLOGY | Water bath (Tissue Floatation) | 2 |
| 244 | PHARMACOLOGY | Manikins for demonstration of intravenous injection, enema, local, intramuscular injections, intracardiac injection and other routes of drug administration | 10 |
| 245 | Physical Medicine & Rehabilitation | Skeleton Traction set | 2 |
| 246 | Physical Medicine & Rehabilitation | Inter ferential therapy unit | 2 |
| 247 | Physical Medicine & Rehabilitation | Short Wave Diathermy | 2 |
| 248 | Physical Medicine & Rehabilitation | Hotpacks & Hydrocollator | 2 |
| 249 | Physical Medicine & Rehabilitation | Exercise table | 2 |
| 250 | Physical Medicine & Rehabilitation | Static Cycle | 2 |
| 251 | Physical Medicine & Rehabilitation | Medicine ball | 2 |
| 252 | Physical Medicine & Rehabilitation | Quadricaps exerciser | 2 |
| 253 | Physical Medicine & Rehabilitation | Coordination Board | 1 |
| 254 | Physical Medicine & Rehabilitation | Handgrip strength measurement Board* | 1 |
| 255 | Physical Medicine & Rehabilitation | ADL Kit & hand exerciser | 1 |
| 256 | Physical Medicine & Rehabilitation | CPM Machine for knee joint | 1 |
| 257 | Physical Medicine & Rehabilitation | CPM Machine for Shoulder joint | 1 |
| 258 | Physical Medicine & Rehabilitation | Wall bar(Wall mounting) | 1 |
| 259 | Physical Medicine & Rehabilitation | Shoulder wheel(Wall mounted) | 1 |
| 260 | Physical Medicine & Rehabilitation | Laser | 1 |
| 261 | Physical Medicine & Rehabilitation | Pelvic Traction bed | 1 |
| 262 | Physical Medicine & Rehabilitation | Survical Traction bed(Wall mounted) | 1 |
| 263 | Physical Medicine & Rehabilitation | Survical Traction bed | 1 |
| 264 | Physical Medicine & Rehabilitation | Parallel Bar with mirror | 1 |
| 265 | Physical Medicine & Rehabilitation | Staircase up & Down (Wooden) | 1 |
| 266 | Physical Medicine & Rehabilitation | Suspension Bed with pulley | 2 |
| 267 | Physical Medicine & Rehabilitation | Gym Ball (Paediatric & Adult) | 2 |
| 268 | Physical Medicine & Rehabilitation | Ultrasound Therapy | 1 |
| 269 | Physical Medicine & Rehabilitation | TENS E | 1 |
| 270 | Physical Medicine & Rehabilitation | Electrical Stimulator | 1 |
| 271 | Physical Medicine & Rehabilitation | Infrared Lamp | 4 |
| 272 | PHYSIOLOGY | Polygraphs | 1 |
| 273 | PHYSIOLOGY | Bicycle Ergometer | 1 |
| 274 | PHYSIOLOGY | Ophthalmoscope | 5 |
| 275 | PHYSIOLOGY | Otoscope | 1 |
| 276 | PHYSIOLOGY | Multi channel Physiograph, 3 channels, complete with accessories | 1 |
| 277 | PHYSIOLOGY | Student physiograph, (single channel) with accessories | 2 |
| 278 | PHYSIOLOGY | Centrifuge, high speed with tachometer | 1 |
| 279 | PHYSIOLOGY | Digital Physiograph | 1 |
| 280 | PSYCHIATRY | Electro Convulsive Therapy (E.C.T.) machine preferably with ECG & EEG monitoring | 1 |
| 281 | PSYCHIATRY | ECT machine without monitor | 1 |
| 282 | PSYCHIATRY | Lithium analyzer (may be shared with clinical pathology department) | 1 |
| 283 | PSYCHIATRY | Thin layer chromatography (for drug dependence treatment) (may be shared with other departments or in Central Research lab) | 1 |
| 284 | PSYCHIATRY | Ventilating circuit, monitors, ECG/ETCO2 and paO2, noninvasive BP monitoring equipment for resuscitation, intubation, ventilation and suction. | 1 |
| 285 | Respiratory Medicine | Fiber optic Bronchoscope | 1 |
| 286 | Respiratory Medicine | Rigid Bronchoscope | 1 |
| 287 | Respiratory Medicine | Spirometer | 1 |
| 288 | Respiratory Medicine | Ventilator | 6 |
| 289 | SKILL LABS | First aid, Bandaging, spliinting | 1 |
| 290 | SKILL LABS | Basic Life Support (BLS), CPR (cardio Pulmonary Resuscitation)mannequin | 1 |
| 291 | SKILL LABS | Various Types of injections- Subcutaneous, Intra-muscular, Intra-venous | 1 |
| 292 | SKILL LABS | Urine Catheter insertion | 1 |
| 293 | SKILL LABS | Skin & Fascia Suturing | 1 |
| 294 | SKILL LABS | Breast examination model/mannequin | 1 |
| 295 | SKILL LABS | Gynecological examination mode/mannequin including ICUD (Intra Uterine Contraceptive Device)Training model | 1 |
| 296 | SKILL LABS | Obstetrics mannequins including obstetric examination, conduct and management of vaginal delivery. | 1 |
| 297 | SKILL LABS | Neonatial & Pediatric resuscitaion mannequins | 1 |
| 298 | SKILL LABS | Whole body mannequins | 1 |
| 299 | SKILL LABS | Trauma mannequin | 1 |

List of FOC Item (Low Cost Equipment)

| Sr. No. : | Department : | Name of the Items : | Qty. : | Type : |
| --- | --- | --- | --- | --- |
| 1 | ANAESTHESIOLOGY | Blood Pressure Instrument (Non-invasive) | 4 | Free of Cost |
| 2 | ANAESTHESIOLOGY | Height Measurement Scale | 4 | Free of Cost |
| 3 | ANAESTHESIOLOGY | Weighing machine | 4 | Free of Cost |
| 4 | ANAESTHESIOLOGY | LMA / PLMA of all sizes | 8 | Free of Cost |
| 5 | ANAESTHESIOLOGY | Resuscitation equipment (CPR)- Ambu bag with face mask | 1 | Free of Cost |
| 6 | ANAESTHESIOLOGY | Blood Pressure Monitor (digital/Electronic) | 5 | Free of Cost |
| 7 | ANATOMY | Brain knife | 2 | Free of Cost |
| 8 | ANATOMY | Plastic tanks for storing soft and dissected parts | 4 | Free of Cost |
| 9 | ANATOMY | Articulated Skeleton set | 2 | Free of Cost |
| 10 | ANATOMY | Bones (Dis-articulated) Set | 10 | Free of Cost |
| 11 | BIOCHEMISTRY | Urinometers calibrated (Mercury based instruments to be replaced with other alternatives) | 5 | Free of Cost |
| 12 | BIOCHEMISTRY | Digital Colorimeters | 3 | Free of Cost |
| 13 | BIOCHEMISTRY | Glucometer with strips ( For POCT ) | 5 | Free of Cost |
| 14 | BIOCHEMISTRY | Thermometer 0 – 250 degree Celsius | 5 | Free of Cost |
| 15 | BIOCHEMISTRY | Digital pH meters | 5 | Free of Cost |
| 16 | BIOCHEMISTRY | Fixed volume pipettes -- 1ml,0.5ml,0.2ml,0.1ml and 0.02ml (Qty = 10 of each volume i.e. total 50 volume) | 50 | Free of Cost |
| 17 | BIOCHEMISTRY | Bottle dispensers | 10 | Free of Cost |
| 18 | BIOCHEMISTRY | Variable and fixed volume micro auto pipettes (Qty = 2 of each volume i.e. total 4 volume) | 4 | Free of Cost |
| 19 | BIOCHEMISTRY | Vortex mixers | 2 | Free of Cost |
| 20 | BIOCHEMISTRY | Vacutainer Tube | 1 | Free of Cost |
| 21 | COMMUNITY MEDICINE | Hydrometres, milk | 2 | Free of Cost |
| 22 | COMMUNITY MEDICINE | Balance for weighing food stuff(Capacity 2 Kg). | 2 | Free of Cost |
| 23 | COMMUNITY MEDICINE | Weighing machine adult | 6 | Free of Cost |

| 24 | COMMUNITY MEDICINE | Harpender Calipers (for skinfold thickness) | 2 | Free of Cost |
| --- | --- | --- | --- | --- |
| 25 | COMMUNITY MEDICINE | Height measuring stand | 3 | Free of Cost |
| 26 | COMMUNITY MEDICINE | Refrigerator 9 cu.ft. | 3 | Free of Cost |
| 27 | COMMUNITY MEDICINE | Chloroscope | 10 | Free of Cost |
| 28 | COMMUNITY MEDICINE | Horrock’s Apparatus | 3 | Free of Cost |
| 29 | COMMUNITY MEDICINE | MUAC tapes | 10 | Free of Cost |
| 30 | COMMUNITY MEDICINE | Haemoglobinometer | 5 | Free of Cost |
| 31 | COMMUNITY MEDICINE | BP Apparatus (Digital) | 10 | Free of Cost |
| 32 | COMMUNITY MEDICINE | Stethoscope | 10 | Free of Cost |
| 33 | COMMUNITY MEDICINE | Sound level meter | 3 | Free of Cost |
| 34 | COMMUNITY MEDICINE | Water sampling bottle from any depth | 1 | Free of Cost |
| 35 | COMMUNITY MEDICINE | Needle Shredder | 3 | Free of Cost |
| 36 | COMMUNITY MEDICINE | Vaccine carrier | 5 | Free of Cost |
| 37 | COMMUNITY MEDICINE | Craft water testing kit | 1 | Free of Cost |
| 38 | COMMUNITY MEDICINE | Treatment kits as per national health programs | 3 | Free of Cost |
| 39 | COMMUNITY MEDICINE | Iodine testing kit | 10 | Free of Cost |
| 40 | COMMUNITY MEDICINE | Glucometer | 10 | Free of Cost |
| 41 | COMMUNITY MEDICINE | Mosquito catching kit | 1 | Free of Cost |
| 42 | COMMUNITY MEDICINE | Clinical Thermometer | 10 | Free of Cost |
| 43 | COMMUNITY MEDICINE | First Aid Kit | 1 | Free of Cost |
| 44 | Dentistry | Oxygen Cylinder with Set | 1 | Free of Cost |
| 45 | Dentistry | BP Instrument | 1 | Free of Cost |
| 46 | Dentistry | Emergency Drugs kit | 1 | Free of Cost |
| 47 | DERMATOLOGY, VENEROLOGY & LEPROSY | (b) Giemsa stain | 1 | Free of Cost |
| 48 | DERMATOLOGY, VENEROLOGY & LEPROSY | (c) KOH smear | 1 | Free of Cost |
| 49 | EMERGENCY MEDICINE | Laryngoscope with all sized blades (straight and vurved) including sizes for neonates, infants and children | 4 | Free of Cost |
| 50 | EMERGENCY MEDICINE | Oxygen cylinders | 4 | Free of Cost |
| 51 | EMERGENCY MEDICINE | Spine boards with slings and scotch tape of all sizes | 2 | Free of Cost |
| 52 | EMERGENCY MEDICINE | Splints for all types of fractures | 3 | Free of Cost |
| 53 | EMERGENCY MEDICINE | Basic equipment for suturing and wound care | 3 | Free of Cost |
| 54 | EMERGENCY MEDICINE | Glucometer | 2 | Free of Cost |
| 55 | FORENSIC MEDICINE & TOXICOLOGY | Anthropometric Set including | 2 | Free of Cost |
| 56 | FORENSIC MEDICINE & TOXICOLOGY | A) Folding Metal Rod Upto 7 Ft | 1 | Free of Cost |

| 57 | FORENSIC MEDICINE & TOXICOLOGY | B) Osteometric Board | 1 | Free of Cost |
| --- | --- | --- | --- | --- |
| 58 | FORENSIC MEDICINE & TOXICOLOGY | C) Craniometer | 1 | Free of Cost |
| 59 | FORENSIC MEDICINE & TOXICOLOGY | D) Mandibulometer | 1 | Free of Cost |
| 60 | FORENSIC MEDICINE & TOXICOLOGY | E) Goniometer | 1 | Free of Cost |
| 61 | FORENSIC MEDICINE & TOXICOLOGY | F) Vernier Calipers | 2 | Free of Cost |
| 62 | FORENSIC MEDICINE & TOXICOLOGY | G) Equipment for Reporting Height | 1 | Free of Cost |
| 63 | FORENSIC MEDICINE & TOXICOLOGY | H) Weighing Machine Dial Type Human | 1 | Free of Cost |
| 64 | FORENSIC MEDICINE & TOXICOLOGY | Digital pH Meter | 1 | Free of Cost |
| 65 | FORENSIC MEDICINE & TOXICOLOGY | Refrigerator | 1 | Free of Cost |
| 66 | FORENSIC MEDICINE & TOXICOLOGY | Glass Cutting Pencil | 2 | Free of Cost |
| 67 | FORENSIC MEDICINE & TOXICOLOGY | Dissection Set Complete | 2 | Free of Cost |
| 68 | FORENSIC MEDICINE & TOXICOLOGY | Digital BP Instrument | 2 | Free of Cost |
| 69 | FORENSIC MEDICINE & TOXICOLOGY | Stethoscope | 2 | Free of Cost |
| 70 | FORENSIC MEDICINE & TOXICOLOGY | Dissection Set Complete | 4 | Free of Cost |
| 71 | FORENSIC MEDICINE & TOXICOLOGY | Brain Knife | 2 | Free of Cost |
| 72 | FORENSIC MEDICINE & TOXICOLOGY | Hack Saw | 2 | Free of Cost |
| 73 | FORENSIC MEDICINE & TOXICOLOGY | Rib Shear Left & Right | 2 | Free of Cost |
| 74 | FORENSIC MEDICINE & TOXICOLOGY | Measuring Tape( Steel Tape Roll) | 2 | Free of Cost |
| 75 | FORENSIC MEDICINE & TOXICOLOGY | Magnifying Lens | 2 | Free of Cost |
| 76 | FORENSIC MEDICINE & TOXICOLOGY | X- Ray View Box (4 In 1) | 1 | Free of Cost |
| 77 | FORENSIC MEDICINE & TOXICOLOGY | Tooth Extractor Left & Right | 1 | Free of Cost |
| 78 | FORENSIC MEDICINE & TOXICOLOGY | Hand Set Heat Sealer | 1 | Free of Cost |
| 79 | FORENSIC MEDICINE & TOXICOLOGY | Instrument Trolley | 2 | Free of Cost |
| 80 | FORENSIC MEDICINE & TOXICOLOGY | Rectal Thermometer | 1 | Free of Cost |

| 81 | GENERAL MEDICINE | B.P. Apparatus (Mercury containing instruments should be replaced by digital or other suitable alternatives) | 10 | Free of Cost |
| --- | --- | --- | --- | --- |
| 82 | GENERAL MEDICINE | X-ray viewing box | 5 | Free of Cost |
| 83 | GENERAL MEDICINE | Non-invasive B.P. Apparatus | 5 | Free of Cost |
| 84 | GENERAL MEDICINE | Glucometer | 5 | Free of Cost |
| 85 | GENERAL MEDICINE | Weighing scale | 4 | Free of Cost |
| 86 | GENERAL MEDICINE | *Haemoglobinometer | 4 | Free of Cost |
| 87 | GENERAL MEDICINE | *Urinometer | 2 | Free of Cost |
| 88 | GENERAL SURGERY | Digital/Electronic BP Apparatus, Weighing Machine, Stethoscope, Height scale | 2 | Free of Cost |
| 89 | GENERAL SURGERY | X ray viewing box 4 in 1 | 4 | Free of Cost |
| 90 | GENERAL SURGERY | Digital/Electronic BP Apparatus, Stethoscope | 8 | Free of Cost |
| 91 | GENERAL SURGERY | Weighing Machine, Height scale | 8 | Free of Cost |
| 92 | MICROBIOLOGY | Culture Plates/ Petri Dishes | 200 | Free of Cost |
| 93 | MICROBIOLOGY | Glassware including Pasteur Pipetts | 100 | Free of Cost |
| 94 | MICROBIOLOGY | Facility for Heating Slides | 100 | Free of Cost |
| 95 | MICROBIOLOGY | Anaerobic apparatus | 2 | Free of Cost |
| 96 | MICROBIOLOGY | Micrometer eye pieces | 1 | Free of Cost |
| 97 | MICROBIOLOGY | Micrometer stage | 1 | Free of Cost |
| 98 | MICROBIOLOGY | pH determination apparatus | 2 | Free of Cost |
| 99 | MICROBIOLOGY | Serum inspissators | 1 | Free of Cost |
| 100 | MICROBIOLOGY | VDRL shaker | 1 | Free of Cost |
| 101 | MICROBIOLOGY | Vortex Mixer/ blood Mixer | 2 | Free of Cost |
| 102 | MICROBIOLOGY | ATCC strain - Enterococcus fecalis 29213 | 2 | Free of Cost |
| 103 | MICROBIOLOGY | ATCC strain - E.coli 25922 | 2 | Free of Cost |
| 104 | MICROBIOLOGY | ATCC strain - E.coli 35218 | 2 | Free of Cost |
| 105 | MICROBIOLOGY | ATCC strain - Pseudomonas aeruginosa 27853 | 2 | Free of Cost |
| 106 | MICROBIOLOGY | ATCC strain - Staphylococcus aureus 25923 | 2 | Free of Cost |
| 107 | MICROBIOLOGY | ATCC strain - Staphylococcus aureus 29213 | 2 | Free of Cost |
| 108 | MICROBIOLOGY | Bottles for blood culture | 200 | Free of Cost |
| 109 | MICROBIOLOGY | Micropipettes – Multi channel & Single channel | 6 | Free of Cost |

| 110 | MICROBIOLOGY | Digital Thermometers of different temperatures | 2 | Free of Cost |
| --- | --- | --- | --- | --- |
| 111 | MICROBIOLOGY | PPE | 500 | Free of Cost |
| 112 | MICROBIOLOGY | VTM | 2500 | Free of Cost |
| 113 | OBSTETRICS & GYNAECOLOGY | Weighing machine | 3 | Free of Cost |
| 114 | OBSTETRICS & GYNAECOLOGY | Height scale | 2 | Free of Cost |
| 115 | OBSTETRICS & GYNAECOLOGY | View box | 2 | Free of Cost |
| 116 | OBSTETRICS & GYNAECOLOGY | Digital/ Electronic B.P. Apparatus | 6 | Free of Cost |
| 117 | OBSTETRICS & GYNAECOLOGY | Weighing machine | 2 | Free of Cost |
| 118 | OBSTETRICS & GYNAECOLOGY | Height scale | 1 | Free of Cost |
| 119 | OBSTETRICS & GYNAECOLOGY | Glucometer | 1 | Free of Cost |
| 120 | OBSTETRICS & GYNAECOLOGY | X ray View box | 1 | Free of Cost |
| 121 | OPHTHALMOLOGY | Snellen Chart/Snellen drum with or without remote control | 3 | Free of Cost |
| 122 | OPHTHALMOLOGY | Snellen chart/Snellen drum with or without remote control | 1 | Free of Cost |
| 123 | OTORHINOLARYNGOLOGY | Sterilizer | 1 | Free of Cost |
| 124 | OTORHINOLARYNGOLOGY | Digital /Electronic BP Apparatus | 1 | Free of Cost |
| 125 | OTORHINOLARYNGOLOGY | Stethoscope | 2 | Free of Cost |
| 126 | OTORHINOLARYNGOLOGY | Siegel’s speculum | 1 | Free of Cost |
| 127 | OTORHINOLARYNGOLOGY | Jobson Horne probe | 2 | Free of Cost |
| 128 | OTORHINOLARYNGOLOGY | The ward side room/emergency treatment room should have a mobile spotlight. | 1 | Free of Cost |
| 129 | OTORHINOLARYNGOLOGY | Biopod | 1 | Free of Cost |
| 130 | OTORHINOLARYNGOLOGY | Boyle's Apparatus | 1 | Free of Cost |
| 131 | PAEDIATRICS | Oxygen Cylinder | 6 | Free of Cost |
| 132 | PAEDIATRICS | Oxygen regulator | 6 | Free of Cost |
| 133 | PAEDIATRICS | Oxygen Humidifiers | 6 | Free of Cost |
| 134 | PAEDIATRICS | Oxygen head-box (of each size) | 8 | Free of Cost |
| 135 | PAEDIATRICS | Nebulizers | 4 | Free of Cost |
| 136 | PAEDIATRICS | Digital Weighing machine | 4 | Free of Cost |
| 137 | PAEDIATRICS | Infantometer | 3 | Free of Cost |
| 138 | PAEDIATRICS | Stadiometer | 4 | Free of Cost |

| 139 | PAEDIATRICS | Digital Thermometer-Oral *(Mercury based instruments to be replaced with other alternatives), Low reading thermometer | 20 | Free of Cost |
| --- | --- | --- | --- | --- |
| 140 | PAEDIATRICS | BP measuring Instrument with various cuff sizes - Digital (Only Non mercury alternatives to be used) | 6 | Free of Cost |
| 141 | PAEDIATRICS | X-ray view box | 6 | Free of Cost |
| 142 | PAEDIATRICS | Glucometer | 4 | Free of Cost |
| 143 | PATHOLOGY | Stopwatch reading at 1/5 second | 5 | Free of Cost |
| 144 | PATHOLOGY | Haemo-cytometers with redand white pipettes | 5 | Free of Cost |
| 145 | PATHOLOGY | Staining jars for slides | 10 | Free of Cost |
| 146 | PATHOLOGY | Urinometers calibrated (Mercury based instruments to be replaced with other alternatives) | 2 | Free of Cost |
| 147 | PATHOLOGY | Centrifuge tubes graduated round bottem(4", 6", 8" ) (Qty = 200 of each i.e. total 600) | 600 | Free of Cost |
| 148 | PATHOLOGY | Graduated cylinders for various capicities ranging from 100cc to 1000cc | 10 | Free of Cost |
| 149 | PATHOLOGY | Pipettes of various sizes with disposal tips. (10-100, 100-1000, 1-50) (Qty = 4 of each i.e. total 12) | 12 | Free of Cost |
| 150 | PATHOLOGY | Reagent bottles | 50 | Free of Cost |
| 151 | PATHOLOGY | Dropping Bottles | 10 | Free of Cost |
| 152 | PATHOLOGY | Coplin jars | 12 | Free of Cost |
| 153 | PHARMACOLOGY | Special Drug Delivery systems like Metered Dose Inhalers, Spacers, Rotahalers, Nasal sprays, Transdermal patches, Insulin infusion pumps, Insulin pen etc. | 10 | Free of Cost |
| 154 | PHARMACOLOGY | Samples of dosage formulations of various types including rational and irrational FDC, Essential medicines | 10 | Free of Cost |
| 155 | Physical Medicine & Rehabilitation | Kit for Neuro-development assessment | 1 | Free of Cost |

| 156 | Physical Medicine & Rehabilitation | CBR Manual* | 1 | Free of Cost |
| --- | --- | --- | --- | --- |
| 157 | Physical Medicine & Rehabilitation | Multi Gym Exerciser | 2 | Free of Cost |
| 158 | Physical Medicine & Rehabilitation | Vertical & Inclined Sanding | 1 | Free of Cost |
| 159 | Physical Medicine & Rehabilitation | Beads Counting | 1 | Free of Cost |
| 160 | Physical Medicine & Rehabilitation | Wax bath | 1 | Free of Cost |
| 161 | Physical Medicine & Rehabilitation | self Help devices | 2 | Free of Cost |
| 162 | PHYSIOLOGY | Demonstration eye piece | 3 | Free of Cost |
| 163 | PHYSIOLOGY | Double demonstration eye piece | 3 | Free of Cost |
| 164 | PHYSIOLOGY | Stage incubator | 1 | Free of Cost |
| 165 | PHYSIOLOGY | Westergren's pipette for E.S.R. on stand(with space pipette) | 20 | Free of Cost |
| 166 | PHYSIOLOGY | Wintrobe’s pipette for ESR and PCV with stand | 20 | Free of Cost |
| 167 | PHYSIOLOGY | Hemoglobin-meter Sahli's or Hellige (with spaces) | 5 | Free of Cost |
| 168 | PHYSIOLOGY | Hemocytometer | 5 | Free of Cost |
| 169 | PHYSIOLOGY | Tuning fork time marker 100/sec | 2 | Free of Cost |
| 170 | PHYSIOLOGY | Electrodes | 2 | Free of Cost |
| 171 | PHYSIOLOGY | Spirit lamps | 2 | Free of Cost |
| 172 | PHYSIOLOGY | Marey's tambour | 2 | Free of Cost |
| 173 | PHYSIOLOGY | Perimeter Pristely Smith S/LP.984 B & T | 10 | Free of Cost |
| 174 | PHYSIOLOGY | Digital BP Instrument | 10 | Free of Cost |
| 175 | PHYSIOLOGY | Stethoscopes | 10 | Free of Cost |
| 176 | PHYSIOLOGY | Stethoscopes, demonstration with multiple ear pieces | 10 | Free of Cost |
| 177 | PHYSIOLOGY | Venous pressure apparatus | 1 | Free of Cost |
| 178 | PHYSIOLOGY | Van Slyke's apparatus manometric | 1 | Free of Cost |
| 179 | PHYSIOLOGY | Sherrington starling kymograph (Electrically Driven) | 2 | Free of Cost |
| 180 | PHYSIOLOGY | Low voltage unit for tapping 2 and 4 volts for stimulation | 1 | Free of Cost |
| 181 | PHYSIOLOGY | Electromagnetic time marker | 2 | Free of Cost |
| 182 | PHYSIOLOGY | Douglas bag, complete | 1 | Free of Cost |
| 183 | PHYSIOLOGY | Basal metabolism apparatus | 1 | Free of Cost |
| 184 | PHYSIOLOGY | Mosso's Ergograph | 10 | Free of Cost |

| 185 | PHYSIOLOGY | Clinical thermometer | 10 | Free of Cost |
| --- | --- | --- | --- | --- |
| 186 | PHYSIOLOGY | Compass aesthesiometer | 10 | Free of Cost |
| 187 | PHYSIOLOGY | Thermo-aesthesiometer | 10 | Free of Cost |
| 188 | PHYSIOLOGY | Algometer | 10 | Free of Cost |
| 189 | PHYSIOLOGY | Apparatus for passive movement | 1 | Free of Cost |
| 190 | PHYSIOLOGY | Knee hammer | 10 | Free of Cost |
| 191 | PHYSIOLOGY | Stethograph | 10 | Free of Cost |
| 192 | PHYSIOLOGY | Olfactometer | 1 | Free of Cost |
| 193 | PHYSIOLOGY | Schematic eye | 1 | Free of Cost |
| 194 | PHYSIOLOGY | Phakoscope | 1 | Free of Cost |
| 195 | PHYSIOLOGY | Perimeter with charts (Lister’s) | 2 | Free of Cost |
| 196 | PHYSIOLOGY | Color perception lantern Edridge green | 1 | Free of Cost |
| 197 | PHYSIOLOGY | Maddox rod | 1 | Free of Cost |
| 198 | PHYSIOLOGY | Newtons color wheel | 1 | Free of Cost |
| 199 | PHYSIOLOGY | Tuning fork to test hearing 32-10000 cps(sets-100, 256, 512 hz) | 30 | Free of Cost |
| 200 | PHYSIOLOGY | Dynamometer | 1 | Free of Cost |
| 201 | PHYSIOLOGY | Stop watch | 5 | Free of Cost |
| 202 | PHYSIOLOGY | Colorimeter, photoelectric | 1 | Free of Cost |
| 203 | PHYSIOLOGY | Myograph stand | 2 | Free of Cost |
| 204 | PHYSIOLOGY | pH meter electric | 1 | Free of Cost |
| 205 | PHYSIOLOGY | Electronic stimulator | 1 | Free of Cost |
| 206 | PHYSIOLOGY | Thermometers, balances, microslides and glassware | 5 | Free of Cost |
| 207 | PSYCHIATRY | Alcohol breath analyzer | 1 | Free of Cost |
| 208 | PSYCHIATRY | a) Projective tests | 2 | Free of Cost |
| 209 | PSYCHIATRY | b) Intelligence Tests | 2 | Free of Cost |
| 210 | PSYCHIATRY | c) Personality Tests | 2 | Free of Cost |
| 211 | PSYCHIATRY | d) Neuro psychological tests | 2 | Free of Cost |
| 212 | Respiratory Medicine | Peak flow meters | 4 | Free of Cost |
| 213 | Respiratory Medicine | Nebulizer | 4 | Free of Cost |
| 214 | Respiratory Medicine | Digital BP Instrument | 4 | Free of Cost |
| 215 | Respiratory Medicine | Digital Weighing machine ( Adult) | 2 | Free of Cost |
| 216 | Respiratory Medicine | Height Scale | 2 | Free of Cost |
| 217 | SKILL LABS | Trainer Simulators/models/ mannequins for: | 1 | Free of Cost |

.

Section-V (Technical Specification)

Department; ANAESTHESIOLOGY

Anesthesia Machine and Accessories like Laryngoscope, all size

Endotracheal  Tubes,  Nasal  and  Oral  Airways,  Magill’s  For ceps,

Anesthesia machine and accessories like laryngoscope, all size endotracheal tube, nasal and oral airway, Magill’s forceps :

1. Powder coated structure.

2. Active Anti Hypoxic Electronic device - to ensure no nitrous flow with loss of O2 pressure.

3. Ratio Control - 25-28% mixed ratio between 02 & N2O Gas flow

4. Should have Electronic Flow meter for O2, N2O & Air

5. Should have Electronic display of Gas pressure in the user interface display

6. Cylinder Pressure and operating pressure should be indicated electronically in user interface display.

7. Oxygen Driven Nitrous Supply through Electronically

8. Inner connection of tubing’s should be made of alloy metal or PPU tubes

9. Audio Visual alarms should be available for Oxygen Failure warning Device

10. Electronically measured cylinder pressure, quick-release system for docking and release.

11. Electronically/pneumatically controlled technology valve for High precision pressure relief cum non return valve should be available

12. The vaporizer should be electronic/ Manual with an injector. The container of the vaporizer chamber should be minimum 300 ml (10,1 oz.). Refilling of inhalation agent should be possible during running mode. The liquid level in the vaporizer should be displayed both on vaporizer and control display and an alarm should be given when the level becomes too low. The desflurane vaporizer should not require to be warmed up before.

13. Provision to connect two electronic/ manual controlled vapouriser simultaneously.

14. Switching between manual and automated ventilation should be possible. The system shall contain a continuously adjustable pressure relief safety valve settable between 0-70 cm H20. Total system volume should be as low as possible and not exceed 3 liter including absorber.

15. Gas Inlet port for O2 & N2O, pin- index type 1 each & O2,N2O & Air non-interchangable type hose 1 each

16. Direct Central Pipe Line connectivity to Machine.

17. Standard Maggils circuits,all rubber antistatic tested.

18. Alloy metal tray for working / writing table.

19. Castor wheel with Individual breaks. 4 wheels with individual locking brake.

20. Double Oxygen outlet available for circuirs

| 21. Operating gas pressure available on the user interface |  |
| --- | --- |
| 22. 2 - B” type cylinder carrying facility at back side. |  |
| 23. Circle Absorber shall contain a continuously adjustable pressure relief safety valve settable between 0-70 cm H20.Total system volume should be as low as possible and not exceed 3 liter including absorber. |  |
| 24. Ventilator Specifications: |  |
|  | The equipment shall contain functions for volume- and pressure regulated ventilation. |
|  | The ventilator shall contain the functions ”Spontaneous Breathing” and ”Manual Ventilation”. |
|  | The ventilator shall have the following modes: VCV, PCV, PRVC, SIMV |
|  | The ventilator shall have Pressure Support mode with backup functionality |
|  | The ventilator should not contain any moving parts |
|  | The equipment shall handle high- to low flow anaesthesia during both non-rebreathing and partial rebreathing conditions. Switching between different breathing systems shall be a simple operation. |
|  | It shall be possible to regulate the I:E ratio between 4:1 to 1:8 |
|  | The equipment shall include an integrated continuous PEEP-function ranging from 0 to at least 20 cm H2O or more. |
|  | The set PEEP level should be visible on the control display |
|  | The set pressure should be adjustable between 5 - 70 cm H20 |
|  | The set tidal volume should be adjustable between 20-1500 ml. |
|  | The breathing frequency should be adjustable between 3-80 breaths/min or more |
|  | The maximum flow rate should be approximately 200 l/min |
|  | The equipment should provide a fast rise time in Pressure Controlled mode without overshoot. |
|  | The equipment should have an easy accessible timer displaying hours, minutes and seconds. |
|  | The equipment should have tools for performing a lung recruitment maneuver |
|  | The lung recruitment manoeuvre should be automatic |
|  | It should be possible to measure and visualize dynamic compliance breath by breath |
|  | It shall be possible to pause ventilation and fresh gas flow for a defined time period." |
| 25. Should have single/Double absorbent chamber canister. Easily removable for changing. Supplied with all standard accessories including Bains Circuit. |  |
| 26. Closed circuit system ,Sodalime canister, Single/Double Chamber |  |
| 27. Pediatric Circuit (Jackson Rees) |  |
| 28. Auxiliary O2 Supply point |  |
| 29. Mask of Different size (2 each- 1 to 5) |  |

30. Reservoir bag- 3 litre/5 litre/0.5 litre- 2 each

31. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ;ANAESTHESIOLOGY

Electrical Suction Apparatus

Electrical Suction Apparatus :

1. Vaccum /LPM : - 700 MM Hg , 50 Litres/Min

2. Pump Type- Double rotary vane type

3. flutter free vacuum control knob,

4. Wide mouthed 2 x 2 LITRE (light weight, unbreakable and clear) with self-sealing bungs and mechanical over flow safety device.

5. Noise (in dBA)- 50 dB

6. Collection container & its cap, suctions tube tips, a vacuum gauge, two sets of moisture & microbial filters and control knob

7. SiliconeTubing:8 mm ID x 2 mtr (PVC), 2x2 lt jar (one set

8. It should be Mobility, portability.

9. Equipment shall operate on 220V-240V, 50 Hz, single phase electric supply.

10. ISO/ BIS approve model should be offered.

Department ;ANAESTHESIOLOGY

FibreOptic Bronchoscope

FibreOptic Bronchoscope :

1. Flexible Fibre Optic bronchoscope

2. Cold Light Source

3. Camera

4. Monitor

5. Video recording digital, preferably USB based video recording device

6. Trolley

7. The working length of the fibre scope should be 60 cms to allow mounting of Endotracheal tube on the scope before performing Bronchoscopy.

8. The outer diameter should be less than 3.7 mm and allow mounting of tube size less than 4.5 mm to the scope, lesser diameter preferable. Suction channel is a must.

9. Range of bending at the tip should be minimum 180 degree up and 130 degree down approx.

10. Compatible light source. Automatic light adjustments to maintain optimum brightness. High efficiency halogen source should be supplied with compatible fiberoptic cable

11. Camera processor should give at least Two output

12. a) S-video / DVI

13. b) Composite video

14. Compatible single chip camera system to give clear and big view of structures.

15. At least 14” colour monitor display.

16. Digital recording directly from camera processor without computer in between.

17. Four section trolley for organizing fiberscope and facility to hang the fiberscope without bending.

18. Leak tester to be provided.

19. Standard accessories including brush for cleaning the bronchoscope lumen.

20. Standard USB video recorder

21. Should work on 220VAC +/-10%, 50 Hz.

22. Model Should by USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO..

Department ;ANAESTHESIOLOGY

USG MACHINE (DESIRABLE)

Ultrasound Machine :

1. Ultrasound scanner with integrated trolley with probe, soft touch alphanumeric key board with track ball. With panel switches & control's easily operable.

2. Integrated high-resolution Monitor 15" or more.

3. Probes & Gel holder-conveniently place (2each).

4. Following transducers are to be supplied:

-2.0-5.0 MHz Multi frequency Convex Transducer-One.

-5.0-12.0 MHz Multi frequency Linear Transducer-One.

-5.0-8.0 MHz or more Endo Cavitary Probe-One.

(+/- 1 MHz to be allowed for each):

5. All probes should be electronic transducers and multi-frequency preferably three frequencies and should give aperture & depths of scanning.

6. Controls for Depth, gain compensation, body markers with transducers position.

7. Real-time continuous dynamic focus.

8. Auto annotation facility anywhere on image.

9. Image display in B, B/M & M Model (2B & 2D).

10. Zoom facility minimum five times or more.

11. Shades of grey 256 h. inbuilt cine memory.

12. Unite should be capable of measuring BPD, CRL, FL & AC and other GA parameters.

13. Facility for image magnification, inversion, changing, scan, direction, freeze facility.

14. 8 steps STC/GTC should be available.

15. Frame Rate should be 1000 frames / sec or more.

16. Caliper with trackball for the measurement of distances circumferences, area volume etc. should be possible to make different measurement of single image.

17. Alphanumeric key board, p.Panel Switches & Foot Controls.

18. Patient reports for Obs/Gynae including fetal growth trend for Tissue texture & Trend graph for IUGR cases, Urology and orthopaedics.

19. Give the gain adjustable/rang & its steps.

20. Calculations needed, Velocity, Heart rate, Volume addl. modes.

21. Dicom 3.0 compatible.

22. Review of stored images is desirable.

23. Channels: 100000 or more.

24. Depth: 25 to 30 cm.

25. Dynamic range: 140 dB or more.

26. Cine loop preview for minimum 60 secs. or more.

27. Minimum 2 or more active ports should be there.

28. Should work on 220VAC +/-10%, 50 Hz.

29. Model Should by USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

30. Should be supply with suitable capacity online UPS.

Department ;ANAESTHESIOLOGY

PNS (DESIRABLE)

PNS :

1. Portable, small, battery operated and lightweight

2. Should have LCD display, backlit

3. Variable current range

4. Numerical display of selected current / twitch height / TOF ratio / PTC count

5. Graphical display of twitch height / TOF responses / PTC

6. User programmable pulse width and frequency with default setting

7. Audio visual indication for pulse delivery and error

8. Power saving option

9. Low battery indicator

10. Should have Complete selection of stimulus patterns NM Monitoring

a) 1-twitch on demand and repeat at 1 sec and at 10 sec

b) Train of four (TOF) on demand and repeat at 10 sec

c) Tetanus frequency can be selected (100 Hz or 50 Hz).

d) Double burst stimulation (DBS)

e) Post titanic count (PTC)

11. Audible alert during the delivery

12. Recorder facility with the unit (optional)

13. Nerve locator for peripheral nerve block

14. Disposable stimulating needle with port for injection compatible with nerve locator (20 numbers with each equipment).

15. Model Should by USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ;ANAESTHESIOLOGY

Adult Manikin

Adult Manikin :

1. The human patient simulator comprise of a life like mannequin. It should employ multiple models of human physiology including cardiovascular system, pulmonary system, neuromuscular system, and central nervous system. The models should allow the patient to exhibit clinical signs (e.g., spontaneous breathing, eyelid blinking) and monitored parameters (e.g., electrocardiogram, blood pressure) and should automatically respond totherapeutic intervention without any/ minimal input from the instructor.

2. The mannequin should be controlled completely wirelessly and should not be connected to any control system/instructor computer through wires/hoses.

3. The mannequin should have a realistic skeletal structure, providing true- to-life articulated motion.

4. ABP, CVP, Pulmonary artery pressure, Pulmonary artery occlusion (wedge) pressure, cardiac output

5. The patient simulator should have a pulmonary system that simulateetO2, inO2, etCo2, inN2O, etN2O, metabolic gas exchange. (for example, apnea or hypoventilation and should automatically result in hypercarbia, hypoxemia, decreasing oxyhaemoglobin saturation and tachycardia

A. During spontaneous ventilation, the patient mannequin should breathe with a spontaneously controlled respiratory rate and tidal volumeand adequate oxygenation.

B. Positive pressure ventilation or return of spontaneous ventilation should automatically reverse apnea with the response appropriate to the rate and tidal volume or ventilation.

C. The Patient Simulator should automatically responds to the fraction of inspired oxygen present, such as with smoke inhalation or supplemental oxygen.

6. The patient simulator should have a pharmacology system model with automatic drug recognition and calculation of pharmacodynamics for all commonly used intravenous and inhaled medications, yielding appropriate changes in patient clinical signs and monitored parameters. All patient responses to drug administration should be automatic, dose dependent, and follow an appropriate time course.

7. Patient outcome should be solely based on patient physiology and the treatment administered (e.g., ventilation, oxygen therapy, drug therapy) and should not be influenced by subjective assessment of the operator,Thus providing objective evaluation of clinical performance and reducing risk of negative training transfer.

8. The mannequin should have a realistic airway (mouth, oropharynx, larynx, esophagus, trachea, carina) resembling to that of an actual human patient.

A. Depending on head positioning, choice of clinical tools, and other maneuvers, it should be possible to achieve anywhere from a Cormack Class I (e.g., easy intubation) to a Cormack Class IV (e.g., difficult intubation) airway.

B. The mannequin airway should allow use of airway adjuncts (e.g., combitube, laryngeal mask airway) as they are used in real patients, without any special adjustments by the instructor (e.g., activation of posterior swelling to seat the LMA).

C. The success or failure of airway management should be automatically reflected in the resulting ventilation, oxyhemoglobin saturation, and overall cardiopulmonary stability.

9. The patient simulator should have trauma simulation capabilities, such as:

A. Surgical cricothyroidotomy

B. Articulated mandible

C. Neck articulation

D. Simultaneous bleeding at different sites linked to physiology

E. Secretions from eyes, ears, mouth.

F. Bi-lateral pneumothorax needle decompression at the clinically appropriate location

G. Bi-lateral chest tube insertion at the clinically correct location.

10. Each trauma capability should require minimal instructor input and physiological consequences (e.g., improvement in blood pressure, ventilation, and oxyhemoglobin saturation) should be automatic.

11. The patient simulator should have fully independent left and right lungs.

A. A one-sided pneumothorax should result in chest distention on one side, with the other side rising and falling with spontaneous breathing.

B. The simulator should have independent breath sounds linked to ventilation of each lung for both spontaneous and mechanical ventilation.

C. One-lung ventilation should automatically result in appropriate breath sounds, chest excursion, and pulmonary gas exchange.

D. Independent bilateral trauma feature (needle decompression / chest tube)

12. The patient simulator should have independent blinking eyes and reactive pupils. Eye blinking should be automatic and dependent on the underlying patient physiology (i.e., level-of- consciousness, level of neuromuscular blockade). It should be possible to easily set the pupils manually to different settings (i.e., pinpoint, reactive, non reactive, blown).

13. The patient simulator should be capable of physically shaking, giving a visible cue of convulsions, tremors, or other similar conditions.visible cue of convulsions, tremors, or other similar conditions.

14. The patient simulator should have touch activated, bi-lateral palpable pulses in the following locations: Carotid, Brachial, Radial, Femoral, Popliteal, Pedal (dorsalis and tibialis)

15. The patient simulator should have an advanced cardiac life support system in which:

A. Effective chest compressions automatically yield artificial circulation, cardiac output, central and peripheral blood pressures, palpable pulses, and exhaled CO2.

B. Ineffective chest compressions yield inadequate cardiac output and circulation and an absence of exhaled CO2.

C. Defibrillation energy is automatically identified, quantified, and logged physiological response.

D. Pacing current is automatically identified, quantified, and logged, with appropriate physiological response.

16. The patient simulator should include independent simulations of patients (e.g., young healthy male, pregnant female, elderly patient with coronary artery disease) and injury/disease scenarios (e.g., anaphylactic shock, ruptured spleen, subdural hematonia.)

A. It should be possible to combine any patient with any scenario, creating a wide variety of clinical care simulations.

B. It should be possible to run multiple software patients simultaneously to create multi- patient care simulations.

C. It should be possible to run multiple injury/disease scenarios simultaneously on a particular patient to create multi-trauma simulations.

17. The patient simulator should include educationally complete properly documented clinical simulations including information : Clinical background and scene, Pre-hospital and emergency department learning objectives, Student critical actions, Simulation algorithm, Equipment required for the simulation, and Instructor notes etc.

18. Patient simulator should be supplied complete with

a) Disaster/ Casualty kit which allows the patient to automatically physically bleed in various locations simultaneously and excrete body fluids from the eyes, ears, and mouth.

b)   Wireless instructors workstation communicating over RF (radio frequency) allowing it to be located  up to  150  feet away  from the  patient mannequin during simulation /training exercise

c)    Should have the facility to run control software and monitor waveform display on same instructor’s workstations thus allowing to use it for developing simulations at other locations independent of the patient mannequin.

19. Model Should by USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ;ANAESTHESIOLOGY

Pediatric Manikin

Pediatric Manikin :

1. The human patient simulator comprise of a life like mannequin. It should employ multiple models of human physiology including cardiovascular system, pulmonary system, neuromuscular system, and central nervous system. The models should allow the patient to exhibit clinical signs (e.g., spontaneous breathing, eyelid blinking) and monitored parameters (e.g., electrocardiogram, blood pressure) and should automatically respond to therapeutic intervention without any/ minimal input from the instructor.

2. The mannequin should be controlled completely wirelessly and should not be connected to any control system/instructor computer through wires/hoses.

3. The mannequin should have a realistic skeletal structure, providing true-to-life articulated motion.

4. The patient simulator should have a cardiovascular system that automatically calculates dependent variables (e.g., blood pressure, heart rate) in response to changing cardiovascular system status (e.g intravenous fluid administration), including the following:

A. The manikin should have facility to control blood pressure, heart rate, pulse strength automatically to maintain circulation and perfusion

B. Should have a feature of a myocardial oxygen supply (e.g., diastolic blood pressure) and demand (e.g., cardiac output, heart rate) that yields appropriate cardiac response (e.g., cardiac rhythm, cardiac output, Blood Pressure) to myocardial ischemia. Untreated myocardial ischemia should automatically result in cardiovascular decompensation (BP, Capillary refill) with accompanying cardiac rhythms (e.g., ST- segment depression, ventricular tachycardia, ventricular fibrillation, asystole) and ultimately, cardiovascular collapse.

C. Arterial blood gases (e.g., PaO2, PaCO2, and pH) and mixed venous gases (e.g., PvO2, PvCO2) that can be measured by Lab reports.

D. Hematocrit can be automatically calculated to reflect oxyhemoglobin saturation and administration of a variety of intravenous fluids, such as whole blood, packed red cells, colloids, and crystalloidsby using preset lab reports

E. A complete hemodynamic monitoring package that includes the capability to measure and monitor the following:

5. ABP, Left ventricular blood pressure, CVP, Right atrial pressure, Pulmonary artery pressure, Pulmonary artery occlusion (wedge) pressure, cardiac output.

6. The patient simulator should have a pulmonary system thatcalculate etO2, inO2, etCo2, inN2O,etN2O, metabolic gas exchange (For example, apnea or hypoventilation and should automatically result in hypercarbia, hypoxemia, decreasing oxyhemoglobin saturation and tachycardia)

A. During spontaneous ventilation, the patient mannequin should breathe with a spontaneously controlled respiratory rate and tidal volume to maintain etCo2 and adequate.

B. Positive pressure ventilation or return of spontaneous ventilation should automatically reverse apnea with the response appropriate to the rate and tidal volume or ventilation.

C. The Patient Simulator should automatically responds to the simulated fraction of inspired oxygen present, such as with smoke inhalation or supplemental oxygen. ( simulated cases)

7. The patient simulator should have a pharmacology system model with automatic drug recognition andcalculation of pharmacodynamics for all commonly used intravenous and inhaled medications, yielding appropriate changes in patient clinical signs and monitored parameters. All patient responses to drug administration should be automatic, dose dependent, and follow an appropriate time course.

8. Patient outcome should be solely based on patient physiology and the treatment administered (e.g., ventilation, oxygen therapy, drug therapy) and should not be influenced by subjective assessment of the operator,Thus providing objective evaluation of clinical performance and reducing risk of negative training transfer.

9. The mannequin should have a realistic airway (mouth, oropharynx, larynx, esophagus, trachea, carina) resembling to that of an actual human patient.

A. Depending on head positioning, choice of clinical tools, and other maneuvers, it should be possible to achieve anywhere from a Cormack Class I (e.g., easy intubation) to a Cormack Class IV (e.g., difficult intubation) airway.

B. The mannequin airway should allow use of airway adjuncts (e.g., combitube, laryngeal mask airway) as they are used in real patients, without any special adjustments by the instructor (e.g., activation of posterior swelling to seat the LMA).

C. The success or failure of airway management should be automatically reflected in the resulting ventilation, oxyhemoglobin saturation, and overall cardiopulmonary stability.

10. The patient simulator should have trauma simulation capabilities, such as:

A. Airway opening acquired by head tilt, chin lift and jaw thrust, LMA, ET tube , Fiber optic , gastric tube,

B. Articulated mandible

C. Neck articulation

D. Bleeding moulage modules ( makeup kit) , enlarged liver, fontanelle bulge ,limp head

E. Limp, tone , motion , head seizure ,tounge edema, foreign body obstruction , laryngospasm.

F. Pneumothorax - Bilateral and unilateral chest rise and fall , Normal and abnormal breath sounds bi-lateral at clinically correct location ,

G. Simulated Chest tube insertion unilateral at clinically correct location

11. Each trauma capability should require minimal instructor input and physiological consequences (e.g., improvement in blood pressure, ventilation, and oxyhemoglobin saturation) should be automatic.

12. The patient simulator should have fully independent left and right lungs.

A. A one-sided pneumothorax should result in chest distention on one side, with the other side rising and falling with spontaneous breathing.

B. The simulator should have independent breath sounds linked to ventilation of each lung for both spontaneous and mechanical ventilation.

C. One-lung ventilation should automatically result in appropriate breath sounds, chest excursion, and pulmonary gas exchange..

D. Independent trauma feature (Unilateral needle thoracentesis mid-clavicular)

13. The patient simulator should have independent blinking eyes and reactive pupils. Eye blinking should be automatic and dependent on the underlying patient physiology (i.e., level-of- consciousness, level of neuromuscular blockade). It should be possible to easily set the pupils manually to different settings (i.e., pinpoint, reactive, non reactive, blown).

14. The patient simulator should be capable of physically shaking, giving a visible cue of convulsions, tremors, or other similar conditions.

15. The patient simulator should have touch activated, palpable pulses Brachial, and Femoral

16. The patient simulator should have an advanced cardiac life support system in which:

A. Effective chest compressions automatically yield artificial circulation, cardiac output, central and peripheral blood pressures, palpable pulses, and exhaled CO2.

B. Ineffective chest compressions yield inadequate cardiac output and circulation and an absence of exhaled CO2.

C. Defibrillation shock should auto convert the ECG rhythm and to be logged physiological response.

D. Pacing current is automatically identified, quantified, and logged, with appropriate physiological response.

17. The patient simulator should include independent simulations of patients ( Eg Pediatric male, female with coronary disease) and injury / disease scenarios ( Eg anaphylactic shock, trauma )

A. It should be possible to combine any patient with any scenario, creating a wide variety of clinical care simulations.

B. It should be possible to run multiple software patients simultaneously to create multi- patient care simulations.

C. It should be possible to run multiple injury/disease scenarios simultaneously on a particular patient to create multi-trauma simulations.

The patient simulator should include educationally complete properly documemnted clinical simulations including information : Clinical background and scene, Pre-hospital and emergency department learning objectives, Student critical actions, Simulation algorithm, Equipment required for the simulation, and Instructor notes etc.

Patient simulator should be supplied complete with

a) Disaster/Casualty make up kit which allows the patient to show trauma , simulated bleeding in various locations

b) Wireless instructors workstation communicating over RF (radio frequency) allowing it to be located up to 150 feet away from the patient mannequin during simulation /training exercise

c) Should have the facility to run control software and monitor waveform display on same instructor’s workstations thus allowing to use it for developing simulations at other locations independent of the patient mannequin.

18. Model Should by USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ;ANAESTHESIOLOGY

Oxygen Therapy Unit

Oxygen Therapy Unit :

1. Suitable for treatment of Hypoxemic patients with respiratory distress

2. It should be complaint for use on patients in ICU, wards, emergency department and home oxygen therapy

3. It should be single system for treating infants, paediatric and adult patients

4. Inbuilt flow generator capable of delivering wide range of flows:2-25 litres in paediatric mode and 10-60 litres in adult mode

5. Inbuilt Air/O2 blending and Fio2 monitoring, facility to deliver wide range of oxygen concentration (Fio2) from 21 to 100%

6. It should have inbuilt Air source without need for external compressor

7. Integrated heated humidifier

8. Color display to monitor humidity setting, flow, Fio2 and faults

9. Visual and audible alarm indication for:

a. Tubes disconnect leaks, tube blockages and water out and hardware fault with error codes. Audible power failure alarm

10. Disinfection mode with heated disinfection tube for sterilization of the device after patient use

11. Supplied with heated wire patient breathing tube and nasal cannula of different sizes

12. Pediatric nasal cannula should be made of kink proof material and has adhesive wiggle pads to stick on skin to facilitate kangaroo care

13. It should be compatible for use on tracheostomy patients

14. Model Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

15. Complaint with international safety standards and regulations.

Department ;ANAESTHESIOLOGY

Provision for Resuscitation Equipment and CPR Algorithms

Provision for Resuscitation Equipment and CPR Algorithms :

1. Should have Silicone self expanding Manual Resuscitators of all sizes (Infant, Pediatrics. Adult) with mask of all sizes.

2. 100%, autoclavable; reusable

3. Non re-breathing valve with 40 cm of H2O pressure release for infant and paediatric

4. 360° swiveling patient connector with standard 15 mm / 22 mm diameter

5. Oxygen reservoir bag with valves 2 mtr. PVC oxygen tubing

6. Should have Laryngoscope of both straight blades and curved blades of various size Standard blades (Macintosh, Miller), Non-standard blades ( McCoy)

7. Lighted stylet for oral and nasal intubation

8. Oral Airways -Silicone, autoclavable & reusable in sizes 000, 00 & 0, 1.2,3,4

9. adult size nasal airways, one(1) each 20F, 22F, 24F,26F, 28F, 30F, 32F, 34F

10. Four pediatric nasal airways, One(1) each 12F, 14F,16F,18F

11. Magils forceps all standard sizes

12. All size malleable Stylet ( for endotracheal tube for neonate to adult)

13. Gum elastic bougie(size 3/ 4)

14. Laryngeal Mask Airway Classic reusable (sizes-1 ,2, 3,4 )

15. Cricothyrotomy set

16. Automatic External Defib as per following specs

17. Simple operation, dedicated therapy controls, configurable options.

18. Automated External Defibrillator (AED) capability with Shock Advisory System.

19. Data storage, Transmission and retrieval capabilities.

20. Power:

21. Battery Only Configuration-choice of batteries.

22. Dual battery capability.

23. DC Power Adaptor for transportation

24. Batteries should charge while device operates from Power Adaptor.

25. Low Battery Indicator and message.

26. Warm start.

27. Service indicator.

28. Display: LCD, User selectable contrast, minimum 4 secs of ECG

29. Data Management: Report Types- Three Formats types, two full capacity patient records.

30. Communications: PC Card, Internal Modem, External EIA/TIA Modem, Cellular Modem or serial connection.

31. Scope of supply

32. Main Unit – 1 No/Unit

33. Single Electrode pad – 1 No/Unit

34. Carrying case – 1 No/Unit

35. Battery set – 1 No/Unit

36. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ;ANAESTHESIOLOGY

ABG Machine

ABG Machine :

1. Compact system for measuring pH, pCO2, pO2, -HCO3 & four Electrolytes like Na, K, Ca+ and Cl - in blood.

2. All should be measured in a single injection / aspiration of Sample.

3. May have provision of modular platform for future up gradation to include glucose, lactate & hemoglobin the same machine with the inspiration of single sample.

4. Should be able to analyze all parameters using low blood volume directly from syringe or capillaries.

5. Fast and accurate result of test made available in about 60 seconds.

6. Automatic Calibration by liquid calibrators with flexible time mode. Instrument should have Stand-by mode facility and Economy mode.

7. It should not be cartridge based system.

8. Should be provided with Start-up Kit/Regent/ Calibrators/ Consumables/ Accessories are required performing the machine initial 500 tests or 3 Months test which is higher.

9.   Compatible online UPS with battery back up of at least one hour

10. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ;ANAESTHESIOLOGY

TOF Monitor (PNS)

TOF Monitor (PNS) :

1. Portable, small, battery operated and lightweight

2. Should have LCD display, backlit

3. Variable current range

4. Numerical display of selected current/twitch height/TOF ratio/PTC count

5. Graphical display of twitch height/TOF responses/PTC

6. User programmable pulse width and frequency with default setting

7. Audio visual indication for pulse delivery and error

8. Power saving option

9. Low battery indicator

10. Should have Complete selection of stimulus patterns NM Monitoring

a) 1-twitch on demand and repeat at 1 sec and at 10 sec

b) Train of four (TOF) on demand and repeat at 10 sec

c) Tetanus frequency can be selected (100 Hz or 50 Hz).

d) Double burst stimulation (DBS)

e) Post titanic count (PTC)

11. Audible alert during the delivery

12. Recorder facility with the unit (optional)

13. Nerve locator for peripheral nerve block

14. Disposable stimulating needle with port for injection compatible with nerve locator (20 numbers with each equipment).

15. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ;ANAESTHESIOLOGY

Radio Frequency Ablation Device

| Radio Frequency Ablation Device : |  |
| --- | --- |
| - It should maximize energy delivery by continuously monitoring tissue impedance & adjusting output accordingly. |  |
| - It should have electrode cooling lowers impedance, allowing more current to pass through the tissue to create a larger ablation zone. |  |
| - It should Minimize treatment time – maximum effect in approx. 12 minutes. |  |
| - It should Monitor all critical performance variable. |  |
| - It should be a User-friendly system. |  |
| - It should enhance Physician control through simple, intuitive needle design & reduce potential injury to vital strictures. |  |
| - It should have Unique electrode design circulates water internally reducing increment in tissue impedance & improves energy delivery. |  |
| - It should have Single & cluster electrodes of assorted sizes are available. |  |
| - It should have 17-gauge electrodes which are easy to insert, minimizing patient trauma. |  |
| - It should have Four different lengths offer flexibility in laparoscopicand percutaneous procedures. |  |
| - It should have Temperature probe which measures ablation zone. |  |
| Technical Specification |  |
| - Volt: 220 V |  |
| - Voltage Input range: 200 to 240 V |  |
| - Maximum Input Voltage: 260 Vac |  |
| - Maximum Voltage on any Output Connector: 260 Vac |  |
| - Max Input Power: 420 VA |  |
| - System has an automated 12- minute tissue ablation. |  |
| - Radio Frequency energy is used to heat & coagulate target tissue |  |
| - The System has 200 Watt power pulses energy for generating larger coagulation volumes. |  |
| - System adjusts output energy based on tissue impedance using feedback algorithm |  |
| - System is compatible with Switch Box for multiple ablations at a time. |  |
| - Impedance: Range – 0 to 1000 Ohms |  |
| - Current |  |
| - Power Range – 0 to 200 Watt |  |
| - Time: Range 1 to 30 minutes |  |
| - Temperature: Range 10 Degree to 99 Degree |  |
| - Frequency: 480 k Hz + 10 % |  |
| - Safety Parameter: Impedance Cutoff < 25 ohms ;> 1000 ohms |  |
| - Temperature < 10 degree; > 99 degree |  |
| 19. External Peristaltic Pump for cooling the tip of the electrode |  |
| 20. System is compatible with Both Single & Cluster Electrode. |  |
| 21. Electrode has a separate lumen to allow water to pass through for cooling the tip. |  |
| 22. System is compatible with Single and Cluster Electrode. The tip of the electrodes is cooled by water (pumped through external Peristaltic Pump). |  |
| 23. Flow Rate for Peristaltic Pump: Max Output Flow Rate: 140 ml/min |  |
| 24. Minimum Output Flow Rate: 80 ml/min through 1.6 mm (ID) tubing. |  |
| 25. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO. |  |

Department ;ANAESTHESIOLOGY

Patients Controlled Analgesia System (Portable)

Patients Controlled Analgesia System (Portable) :

1. Must accommodate any syringe size of: a) 10 ml b) 20 ml c) 50 ml.

2. Flow must be adjustable between 0.1 ml to 999 ml/hr.

a) Accuracy +/- 2 %.

3. Set flow rate and volume infused should be digitally displayed.

4. Delivery rate should be preset on Delivery rate and on volume and time pre selection.

a) Volume infused should be displayed.

5. Must have a fast mode of infusion independent of set flow rate.

6. Should not allow change of flow rate or fast delivery without stopping the pump.

7. Should have alarms for :

a. Occlusion

b. Syringe almost empty

c. Low Battery

8. Occlusion sensitivity should be adjustable by the operator.

9. Should have a built in rechargeable battery and a battery charger.

10. Should provide clamp for fixing on IV pole and should be stackable.

11. It should have system to give bolus volumes of 5ml or more than 5ml during infusion.

12. It should be upgradeable.

13. It should be provide security against tampering with ability to record and retrieve drug pump/Microprocessor malfunction. The dosing modes: PCA,CBI, PCA + CBI and loading.

14. It should provide printer output port.

15. Dosage adjustments should be possible in mg/hr and ml/hr

16. To be supplied with standard 50ml syringe/tubings:200 Nos each pump

17. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department; ANAESTHESIOLOGY

OT Table Fluoroscopy Compatible

OT Table Fluoroscopy Compatible:

1. Electro- Hydraulic operation table

2. heavy weight of the patient. Complete with all accessories.

3. Electro-hydraulically powered for movements of height, tilting and trendelenburg.

4. The base of stable construction with large castors for easy travel and manoeuvring

5. Table top 500 mm x 1940 mm, Height range 600- 750 mm x 950- 1050 mm

6. Trend/Antitrend-20-25, latel tilt- 15-20 degree

7. The Operation table provided with corded hand control and override – control panel for additional safety.

8. The tabletop divided in 5 parts i.e., head plate (removable), back plate, seat plate and leg plates (removable).

9. Guided rails for X-Ray cassettes inserted from the head end.

10. Back plate is motorised for easy movements

11. The back plates are gas spring adjusted offering easy and smooth control.

12. The Operation table should have safe positioning and adjustment due to optimal covering of moving joints.

13. The Operation table should provide optimal patient positioning in a supine position. The longitudinal shift feature of 290mm should give unobstructed imaging of vertebrae.

14. Mobile Operating Table, Electrohydraulic, SFC padding

15. Standard Accessories:

16. Head Rest, Leg plates, pair, hinged and abductable, SFC padding

17. Arm board, Radial Setting Clamp, Geopel knee crutch, Body Strap

18. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ;ANAESTHESIOLOGY

Nerve Locator

Nerve Locator :

1. Portable, small, battery operated and lightweight

2. Should have LCD display,backlit

3. Variable current range

4. Numerical display of selected current/twitch height/TOF ratio/PTC count

5. Graphical display of twitch height/TOF responses/PTC

6. User programmable pulse width and frequency with default setting

7. Audio visual indication for pulse delivery and error

8. Power saving option

9. Low battery indicator

10. Should have Complete selection of stimulus patterns NM Monitoring

a. 1-twitch on demand and repeat at 1 sec and at 10 sec

b. Train of four (TOF) on demand and repeat at 10 sec

c. Tetanus frequency can be selected (100 Hz or 50 Hz).

d. Double burst stimulation (DBS)

e. Post titanic count ( PTC)

11. Audible alert prior to actual delivery

12. Recorder facility with the unit (optional)

13. Nerve locator for peripheral nerve block

14. Disposable stimulating needle with port for injection compatible with nerve locator (20 numbers with each equipment )

15. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; ANAESTHESIOLOGY

Transcutaneous Electric Nerve Stimulating Machine

Transcutaneous Electric Nerve Stimulating Machine :

1. Portable, small, battery operated and lightweight

2. Should have LCD display, backlit

3. Variable current range

4. Numerical display of selected current/twitch height/TOF ratio/PTC count

5. Graphical display of twitch height/TOF responses/PTC

6. User programmable pulse width and frequency with default setting

7. Audio visual indication for pulse delivery and error

8. Power saving option

9. Low battery indicator

10. Should have Complete selection of stimulus patterns NM Monitoring

a) 1-twitch on demand and repeat at 1 sec and at 10 sec

b) Train of four (TOF) on demand and repeat at 10 sec

c) Tetanus frequency can be selected (100 Hz or 50 Hz).

d) Double burst stimulation (DBS)

e) Post titanic count (PTC)

11. Audible alert prior to actual deliver

12. Recorder facility with the unit (optional)

13. Nerve locator for peripheral nerve block

14. Disposable stimulating needle with port for injection compatible with nerve locator (20 numbers with each equipment)

15. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; ANAESTHESIOLOGY

Anaesthesia Machine with Resuscitation Equipment

Anaesthesia Machine with Resuscitation Equipment :

1. Powder coated structure.

2. Active Anti Hypoxic Electronic device - to ensure no nitrous flow with loss of O2 pressure.

3. Ratio Control - 25-28% mixed ration between 02 & N2O Gas flow

4. Should have Electronic Flow meter for O2, N2O & Air

5. Should have Electronic display of Gas pressure in the user interface display

6. Cylinder Pressure and operating pressure should be indicated electronically in user interface display.

7. Oxygen Driven Nitrous Supply through Electronically

8. Inner connection of tubing’s should be made of alloy metal or PPU tubes

9. Audio Visual alarms should be available for Oxygen Failure warning Device

10. Electronically measured cylinder pressure, quick-release system for docking and release.

11. Electronically/pneumatically controlled technology valve for High precision pressure relief cum non return valve should be available

12. The vaporizer should be electronic/ Manual with an injector. The container of the vaporizer chamber should be minimum 300 ml (10,1 oz.). Refilling of inhalation agent should be possible during running mode. The liquid level in the vaporizer should be displayed both on vaporizer and control display and an alarm should be given when the level becomes too low. The desflurane vaporizer should not require to be warmed up before.

| 13. Provision to connect two electronic/ manual controlled vaporiser simultaneously. |  |
| --- | --- |
| 14. Switching between manual and automated ventilation should be possible. The system shall contain a continuously adjustable pressure relief safety valve settable between 0-70 cm H20. Total system volume should be as low as possible and not exceed 3 liter including absorber. |  |
| 15. Gas Inlet port for O2 & N2O, pin- index type 1 each & O2,N2O & Air non-interchangable type hose 1 each |  |
| 16. Direct Central Pipe Line connectivity to Machine. |  |
| 17. Standard Maggils circuits,all rubber antistatic tested. |  |
| 18. Alloy metal tray for working / writing table |  |
| 19. Castor wheel with Individual breaks. 4 wheels with individual locking brake. |  |
| 20. Double Oxygen outlet available for circuirs |  |
| 21. Operating gas pressure available on the user interface |  |
| 22. 2 - B” type cylinder carrying facility at back side. |  |
| 23. Circle Absorber shall contain a continuously adjustable pressure relief safety valve settable between 0 -70 cm H20. Total system volume should be as low as possible and not exceed 3 liter including absorber. |  |
| 24. Ventilator Specifications: |  |
|  | The equipment shall contain functions for volume- and pressure regulated ventilation. |
|  | The ventilator shall contain the functions ”Spontaneous Breathing” and ”Manual Ventilation”. |
|  | The ventilator shall have the following modes: VCV, PCV, PRVC, SIMV |
|  | The ventilator shall have Pressure Support mode with backup functionality |
|  | The ventilator should not contain any moving parts |
|  | The equipment shall handle high - to low flow anaesthesia during both non-rebreathing and partial rebreathing conditions. Switching between different breathing systems shall be a simple operation. |
|  | It shall be possible to regulate the I:E ratio between 4:1 to 1:8 |
|  | The equipment shall include an integrated continuous PEEP-function ranging from 0 to at least 20 cm H2O or more. |
|  | The set PEEP level should be visible on the control display |
|  | The set pressure should be adjustable between 5 - 70 cm H20 |
|  | The set tidal volume should be adjustable between 20-1500 ml. |
|  | The breathing frequency should be adjustable between 3-80 breaths/min or more |
|  | The maximum flow rate should be approximately 200 l/min |

|  | The equipment should provide a fast rise time in Pressure Controlled mode without overshoot. |
| --- | --- |
|  | The equipment should have an easy accessible timer displaying hours, minutes and seconds. |
|  | The equipment should have tools for performing a lung recruitment maneuver |
|  | The lung recruitment manoeuvre should be automatic |
|  | It should be possible to measure and visualize dynamic compliance breath by breath |
|  | It shall be possible to pause ventilation and fresh gas flow for a defined time period." |
| 25. Should have single / Double absorbent chamber canister. Easily removable for changing. Supplied with all standard accessories including Bains Circuit. |  |
| 26. Closed circuit system, Sodalime canister, Single / Double Chamber |  |
| 27. Pediatric Circuit (Jackson Rees) |  |
| 28. Auxiliary O2 Supply point |  |
| 29. Mask of Different size (2 each- 1 to 5) |  |
| 30. Reservoir bag- 3 litre/5 litre/0.5 litre- 2 each |  |
| 31. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO. |  |

Department ; ANAESTHESIOLOGY

Nerve Stimulator

Nerve Stimulator :

1. Portable, small, battery operated and lightweight

2. Should have LCD display ,backlit

3. Variable current range

4. Numerical display of selected current/twitch height/TOF ratio/PTC count

5. Graphical display of twitch height/TOF responses/PTC

6. User programmable pulse width and frequency with default setting

7. Audio visual indication for pulse delivery and error

8. Power saving option

9. Low battery indicator

10. Should have Complete selection of stimulus patterns NM Monitoring

a) 1-twitch on demand and repeat at 1 sec and at 10 sec

b) Train of four (TOF) on demand and repeat at 10 sec

c) Tetanus frequency can be selected (100 Hz or 50 Hz).

d) Double burst stimulation (DBS)

e) Post titanic count (PTC)

11. Audible alert prior to actual delivery

12. Recorder facility with the unit (optional)

13. Nerve locator for peripheral nerve block

14. Disposable stimulating needle with port for injection compatible with nerve locator (20 numbers with each equipment )

15. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; ANAESTHESIOLOGY

Patient Examination Table,

Patient examination table :

1. Overall Approx. size: 187L x 51W x 81H.

2. Adjustable backrest by gas spring system.

3. Two section Foam padded, upholstered top.

4. Provided with three Drawers and three Cabinets.

5. Tray for B.P. Apparatus provided near head rest.

6. Finish: Pretreated & Epoxy Powder Coated.

7. Carrying capacity: Approx. 150 - 180 kg

8. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; GENERAL MEDICINE

Portable Suction Machine

Portable Suction Machine :

1. Vaccum /LPM : - 700 MM Hg , 50 Litres/Min

2. Pump Type- Double rotary vane type

3. flutter free vacuum control knob,

4. Wide mouthed 2 x 2 LITRE (light weight, unbreakable and clear) with self-sealing bungs and mechanical over flow safety device.

5. Noise (in dBA)- 50 dB

6. Collection container & its cap, suctions tube tips, a vacuum gauge, two sets of moisture & microbial filters and control knob

7. SiliconeTubing:8 mm ID x 2 mtr (PVC), 2x2 lt jar (one set

8. It should be Mobility, portability.

9. Equipment shall operate on 220V-240V, 50 Hz, single phase electric supply.

10. ISO/ BIS approve model should be offered.

Department : ANATOMY

Drill machine

1. The table top drill mounted on stand should have spring loaded lever of high quality.

2. Should have variable speed operation, high speed drill.

3. The single phase ac motor should have approx. 690 W.

4. It should be compact with overall length (320mm Approx.) and weight (Approx. 1.8Kg)

5. Aluminum die-casting inner cover and gear cover should be used for increased durability.

6. Should be easy to operate with a finger sized trigger switch and a speed control dial

7. It should have a reliable and convenient push button type forward/reverse change over switch.

8. Should have 2 speed transmission.

9. Shortest chuck offset (23.5 mm (15/16"))

10. It should have soft grip handle for comfortable operation

11. It should have low operation noise 79dB

12. To be provided with high speed tungsten carbide bits of 1/2", 3/8", 1/4", 3/16", 1/8" or equivalent size

13. Should be ISO / BIS / ISO approved model should be offered

Department : ANATOMY

Hand saw (preferably metal),

1. Should be high carbon steel blades , hard and durable, anti rust and smooth, which lead sto less friction.

2. Blades should be chrome plated to prevent body fluid / chemical corrosion.

3. For HRC, Blade materials should reach 52 degrees and teeth reaches 60 degrees(+/ - 3)

4. Chrome treatment, antirust smoothening should be done to reduce the friction.

5. It should have strong handles.

6. It should be ergonomic and rust proof.

7. Special teeth design which should be sharp at edges.

8. Should be supplied with essential accessories.

9. Power Supply 200VAC +/- 10 %, 50Hz fitted with Indian plug.

10. Should be ISO/BIS / ISO approved approved model should be offered

Department : ANATOMY

Band saw

1. Body & meat cutting band saw

2. It should attach with a large movable stainless steel work table (electrically powered)

3. It should be fitted with 1 H.P. motor , 50 Hz, 230V

4. Should be supplied with stand, protection and spare blades (2 nos.)

5. Power Supply 200VAC +/- 10 %, 50Hz fitted with Indian plug.

6. Should be ISO / BIS / ISO approved model should be offered

Department : ANATOMY

Storage tank to hold 10 cadavers. Static /movable ,Durable tank with

input and output facility with lid

1.   Storage to hold 10 cadavers, static/movable, durable tank with input and output facility with

lid

2.   Should be ISO / BIS / ISO approved model should be offered

Department : ANATOMY

Dissecting instruments for cadaveric dissection

| 1. | Technical Specifications: |  |
| --- | --- | --- |
| A. Student Anatomy Dissecting Kit B. Should contain most widely used instruments of high quality. |  |  |
| 2. | Dissection Kit contents: |  |
| A. Cartilage knife 2" blade B. Scalpel with screw lock blade |  |  |
|  | C. | Narrow blade scalpel 1.5" blade |
| D. Forceps 4.5" with guide pin, medium points E. Forceps 4.5" with guide pin curved, fine points |  |  |
|  | F. | Dissecting scissors, Iris 4.5" |
| G. Probe and hook chrome H. Dissecting scissors with one point sharp & one point blunt 5.5" |  |  |
| I. Teasing needle straight |  |  |
|  | J. | Teasing needle bent |
|  | K. | Ruler 6" plastic |
|  | L. | Dissecting chain and hook chrome |
| M. Dissecting blow pipe 6" N. Double fold vinyl case |  |  |
| O. B.P. Handle (No.-4) & B.P. Blade (No.-22 & 24) |  |  |
| P. Tooth Forceps |  |  |
| Q. Plain Forceps R. Scissor Straight |  |  |
|  | S. | Magnifying glass |
|  | T. | Curved Scissor |
| 3. | Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO. |  |

Department : ANATOMY

Embalming Machine

1. Fluid delivery rate should be 10ltrs/hr.

2. Inner tank to store embalming fluid with capacity 5-7 ltrs. should be of stainless steel.

3. Pump: pump should be of ELECTROMAGNETIC DOSING PUMP with capacity 0-5 ltrs. per hour and pressure 3 kg/cm square.

4. The equipment should be mounted on castors for easy movement and the hand grip should be provided for lifting.

5. I.V. stand fixed for mounting cannula tubing and mains cable.

6. Indicator for mains on & in use should be present.

7. The outer body should be of complete stainless steel.

8. Should be ISO / BIS / ISO approved model should be offered.

Department : ANATOMY

Dissection microscope

1.   Eye piece: Straight binocular type wide field (10 x)

2.   Optic carrier with five steps magnification.

3. Fine focusing manual

4. Objective 250 mm f & 400 mm f.

5. Light co-axial illumination additional 10 spare bulbs.

6. Solid metallic body with sturdy stand riding on heavy castor wheels with locking breaks.

7. Halogen illumination 6V/15W with power supply

8. Should have 3 spare lamps with each unit.

9. Power Supply-Power input to be 220-240VAC, 50Hz.

10. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department : ANATOMY

Microtomes Rotary

1. Fully automated motorized rotary microtome along with manual operation having microprocessor controlled panel with provision for motorized cutting via operating panel or foot pedal control

2. Precise micrometer feed system via stepper motor permits precision sectioning selectable at least from 2.0-40/60 micron in 0.5 micron increments.

3. Trimming section selectable from 2 micron onwards.

4. The vertical specimen stroke length of 70mm, larger specimen can be sectioned. The specimen holder should be clamp type and hold 60mm size block.

5. Suitable knife holder for high profile and low profile should be provided

6. The specimen retraction should occur on return stroke.

7. Knife angle position locking facility should be provided.

8. Cold light source

9. Precise specimen orientation with zero point indication, with an orientation 6º X-Y axis helps in marking perfect orientation of the sample for sectioning.

10. Motorised coarse feed in two speeds 30 micron/sec and 90 micron/sec. Variable sectioning speed adjustable from 0.5 to 420 mm/sec.

11. Disposable blade holder with lateral displacement feature that can hold both high and low profile blades and knife holder which can accommodate 16 cm C&D type knifes.

12. Knife holder should be vibration free.

13. Integrated section waste tray.

14. Accessories

15. Microtome disposable blades (high profile coated) - 20 packets (50 blades/pack). Total 1000 Nos.

16. Microtome disposable blades (low profile coated) - 2 packets (50 blades/pack).

17. C type knife 16 cm - 3 Nos

18. Cold plate (Dry type)

19. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department : ANATOMY

Paraffin Bath Embedding

1. Embedding station should have 2 separate components cold plate heated paraffin dispensing module.

2. Module should allow flexibility to arrange embedding work flow in the either direction.

3. Capacity of Paraffin reservoir should be 3 lit. capacity with vacuum lid, easy to clean, scratch resistant, heated wailing surface with paraffin drain system.

4. Adjustable paraffin flow rate.

5. Two heated removable paraffin waste tray.

6. Paraffin outlet:manual or foot pedal operated paraffin outlet.

7. Removable and interchangeable cassette and mould warmer.

8. Large peltier cooling spot, even for super mega cassettes.

9. Adjustable temperature for all modules between 55 to 70 Deg C

10. Heated forceps holder which is easily removed for cleaning.

11. Capacity of approximate ranging 100 to 150 cassettes.

12. Programmable timing and temperature.

13. Magnifier (essential)

14. Illuminator (essential)

15. Main: 230V, 50 Hz

16. Height - 400 mm

17. Width - 500 mm

18. Depth - 700 mm

19. Weight - 30 Kg

20. Size of working surface - 2 x 175 mm

21. Size of cold spot - 70 mm

22. Accessories

23. Removable holder.

24. Paraffin scrapper

25. Freezer

26. Replacement bulb

27. Set of fuse

28. Set of power cords

29. Cassette - 50,000 labels

30. Stainless steel capsules for matis mould.

31. ONLY Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; BIOCHEMISTRY

Analytical Balance : upto 200g/1gm increment

Analytical Balance : upto 200g/1gm increment :

1. Microprocessor based single pan Analytical Balance with High accuracy & precision is required. Reading of the weight by digital display.

2. Weigh accurately up to 3rd decimal place

3. Auto self-calibration facility

4. Auto zero Setting, TARE setting

5. One touch calibration

6. Weighing capacity up to 200g

7. Readability and repeatability 0.001g

8. Stabilization time < 5 second

9. Liquid Crystal Display (LCD) for display

10. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; BIOCHEMISTRY

Semi Autoanalyzer

Semi Autoanalyzer :

1.   Should be microprocessor controlled general purpose bi-chromatic photometer system with at least 6 filters/grating system ranging from 340 to 630 nm.

2. Temperature 37 self monitoring built-in incubation systems for temperature controlled absorbance reading.

3. Light source: Tungsten/ halogen or higher grade with 2 (two) additional bulb.

4. Should have end point, kinetic and two point kinetic measurement modes.

5. Should have flow cell measuring device having capability of measuring with 500 μl or less volume of reaction mixture.

6. Should have inbuilt printer.

7. Should have a measurement range from 0.001 to 2.500Abs or more

8. Should have facility for reading results on LCD display.

9. The system should have good repeatability and precision

10. Should have quality control – two control/test QC survey of at least 30 points, Levy Jenny plot.

11. Should have a filter half bandwidth of 10nm or lesser.

12. Should have a test programme memory of 50 or more.

13. Should be provided with sample carry over prevention facility.

14. Should be supplied with 1 variable (10-100μl) and one fixed volume 500μl pipettes.

15. Aspiration should be based on Bellow/Peristaltic Pump/ Vacuum pump.

16. Should provide 500 ml of reagents with calibrators for urea, Serum creatinine, Serum bilirubin, Plasma Glucose, Cholesterol, and Quality control 5ml one each for normal and abnormal for installation and demonstration. Should be supplied as and when required by users.

17. Should be provided with Start-up Kit/Regent/ Calibrators/ Consumables/ Accessories are required performing the machine initial 500 tests or 3 Months test which is higher.

18. Should be supplied with on line pure sine wave UPS of sufficient capacity for a minimum back of 30 minutes.

19. Should be provided with calibration certificate issued by the manufacturer at the time of installation and calibration certificate should be issued for the machine by the supplier during preventive maintenance visit in the warranty / AMC period if demanded by the end user.

20. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; BIOCHEMISTRY

Constant Temperature Water Bath Tank Capacity:

(Temperature Range 5 to 80 degree Celsius)

Constant Temperature Water Bath Tank Capacity: (Temperature Range 5 to 80 Degree Celsius) :

1. Double Wall, both outer and Inner chamber are Stainless steel with insulation between outer and inner chambers.

2. Special type of transparent PP/PVC/acrylic cover in the shape of inverted V to avoid evaporation of water and draining of condensed water back into water bath.

3. Shaking speed (Approximately): 25 to 200 strokes/min.

4. Temperature Range: from 5 to 80 o C with accuracy of ± 1 o C

5. Thermostatically controlled with Regulator.

6. Mercury thermometer must be provided

7. Digital temperature indicator cum Regulator

8. Removable stainless steel tray with the following capacities (Approximately) must be provided.

a. 100 x 10-12 flask

b. 250 x 4 - 6 flask

c. 500 ml x 2 - 4 flask

d. Test tube racks.

9. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

10. Should be provided with Start-up Kit/Regent/ Calibrators/ Consumables/ Accessories are required performing the machine initial 500 tests or 3 Months test which is higher.

Department ; BIOCHEMISTRY

Complete Chromatographic Unit for Paper & TLC

Complete Chromatographic Unit for Paper & TLC :

1. TLC Basic Kit, consisting of (Nanomat 4)

2. smartCut plate cutter to cut TLC/HPTLC glass plates up to 20 × 20 cm

3. Twin Trough Chamber for 10 × 10 cm plates, with stainless steel lid

4. Twin Trough Chamber for 20 × 20 cm plates, with stainless steel lid

5. smartAlert solvent front monitor (only suitable for glass plates)

6. Viewing Box 4 for CAMAG UV lamps of the 022.91XX series

7. UV lamp 4 dual wavelength 254/366 nm, 2 × 8 W

8. Glass Reagent Sprayer, all glass, with 100 mL erlenmeyer flask

9. Saturation pads, pack of 100 (20 × 20 cm)

10. Capillary dispenser consisting of universal capillary holder (022.7786), one dispenser

11. magazine for 1 μL capillaries (022.7661and one package of 5 × 100 disposable capillary pipettes 1 μL

12. Dispenser magazine for 2 μL capillaries, without capillaries

13. Dispenser magazine for 5 μL capillaries, without capillaries

14. Disposable capillary pipettes 2 μL, pack of 5 × 100

15. Disposable capillary pipettes 5 μL, pack of 5 × 100

16. MERCK TLC plates silica gel 60 F 254, 20 × 20 cm, pack of 25

17. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

18. Should be provided with Start-up Kit/Regent/ Calibrators/ Consumables/ Accessories are required performing the machine initial 500 tests or 3 Months test which is higher.

Department ; BIOCHEMISTRY

Centrifuge > 16 Tubes

Centrifuge Machine (16 Tube) :

1. Should have a maximum speed of 5000 RPM with steeples regulator

2. Should be supplied with safety lid and lock. (Brushless Motor)

3. Should have digital speedometer and timer.

4. Should have imbalance detector and automatic cut off.

5. Standard mark on quality and safety is Compulsory.

6. Should work on a220VAC +/- 10 %, 50Hz AC Supply.

7. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; BIOCHEMISTRY

Complete Electrophoresis Apparatus with Power Supply

(Paper, PAGE, Agarose)

| Complete Electrophoresis Apparatus with Power Supply (Paper, PAGE, Agarose) : |  |
| --- | --- |
| 1. Electrophoresis apparatus with power pack capable of running both paper and gel electrophoresis |  |
| 2. Dual Chambers, each with buffer capacity of 300ml |  |
| 3. Power pack with both constant voltage and current indicators with user manual. |  |
| 4. Vertical Gel Appartus |  |
| 5. Should be able to run two slab gels appr 10x8cm simultaneously |  |
| 6. Complete Gel casting system for casting multiple gels |  |
| 7. Power connecor integral with safety lid |  |
| 8. Silicon rubber gaskets |  |
| 9. Supply atleast four sets of 1 mm spacer plate and short glass plates |  |
| 10. Supply atleast four Nos of 1 mm thick comb of 8well, 10 well and 15 well. |  |
| 11. Horizondal gel Appartus |  |
| 12. Gel tank size 8x11 inches (midi gel appartus) and 3x6 inches (mini gel appartus) with platinum electrodes and dams |  |
| 13. Complete Gel casting system for casting multiple gels |  |
| 14. power connecor integral with safety lid |  |
| 15. Supply atleast four sets of gel casting trays |  |
| 16. Supply at least six numbers of 1 mm thick comb for 8-20 samples |  |
| 17. Compatible DC Power supply |  |
|  | Compatible microprocessor-based power supply to run at least 2 units at constant voltage or current with automatic cross over |
|  | Output range programmable,10-500V, 4-500 mA in 1 mA step, 100 W maximum |
|  | Single-unit increments in settings and read-outs for precision and reproducibility |
|  | Easy to read digital display |
|  | Ensure safety features for overload, sudden load change, short circuit protection etc |
| 18. Electro Blotter for transfer of proteins from acrylamide gel to nitrocellulose |  |
| 19. Vertical Tank with safety lid and connecting leads |  |
| 20. Blotting pads with holders -2 Nos. compatible with the vertical gel apparatus |  |
| 21. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO. |  |
| 22. Should be provided with Start-up Kit/Regent/ Calibrators/ Consumables/ Accessories are required performing the machine initial 500 tests or 3 Months test which is higher. |  |

Department ; BIOCHEMISTRY

Densitometer with Computer

Densitometer with Computer :

1. Densitometer (or) Gel scanner with the necessary accessories and software either of these with the following features to be procured along with electrophoresis system

2. Scanning & processing all gels including those specified above

3. Facility to store the scanned image of the gel

4. Facility for curve editing and entry of patient demographics

5. Availability of quantification and quality control features

6. Storage of patient data and results – up to a minimum of 10000 Samples

7. Facility to generate a comprehensive report containing patient demographics, scanned image of the gel, curve and quantification data.

8. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

9. System, Configuration and Accessories

All necessary standard accessories like those required for sample application to be provided along with the instrument. Others PC with software should be provided.

Department ; BIOCHEMISTRY

Fume Cupboard

Fume Cupboard :

1-Operational Requirements

The cupboards shall be fume cupboards with exhaust systems comprising of PVC centrifugal fan, Duct work, discharge cowl and supports.

2-Technical Requirements

Service

GPOs are to be located on the outside face of the fume cupboard and are to have RCD protection. The fume cupboard Microprocessor control panel shall be located on the outside face of the fume cupboard and shall include the exhaust system start, gas and power start, light on/off, emergency shut down and a visual light to indicate power is live to the fume cupboard.

Other (piped) Services:

The outlets to other services (e.g. gas, water) shall be finished in polycoat powder lacquer, and shall be located inside the fume cupboard and be colour coded to match control valves. All pipe work between control valves and outlets shall be preplumbed and tested by the fume cupboard supplier.

Controls:

All services, other than electrical, shall be individually controlled and the controls shall be located on the outer face of the fume cupboard and not protrude beyond the line of the face of the fume cupboard. The unit shall include battery back-up that is automatically charged when the fume cupboard is running.

Fixed Volume fume Cupboard

The quantity of air extracted from the conditioned air space through each fume cupboard shall remain constant. The quantity of air through the open sash of the fume cupboard shall be modulated to maintain an average air velocity of 0.5m/s across the open sash in any sash position. As the sash is lowered it will open up a bypass baffle at the top of the fume cupboard, allowing air to flow into the top of the fume cupboard thus balancing the fume cupboard.

Fixed

The sash shall be easily moveable vertically rising panels, glazed with armour plate glass. The sash to be counterbalanced by a one piece balanced weight, supported by cables running through a stainless steel ball bearing pulley. The bottom of the sash will have a full width pull handle of a type that will minimize the airflow disturbance across the sash opening.

Maximum Working Sash Opening:

To reduce the volume of air exhausted from the laboratory the fume cupboards maximum working sash opening shall be limited to 500mm by a physical stop preventing the sash being opened further.

Minimum Sash Opening:

Stops shall be fitted to provide a minimum opening of at least 50mm below the sash.

Design of Sash Opening:

The sash opening shall incorporate aerodynamic features to promote non-turbulent entry of room air into the fume cupboard working chamber.

Baffle

Internal removable air distribution baffles shall be fitted to the rear and top of the fume cupboard chamber. An anti-roll baffle shall be fitted to the top front of the fume cupboard chamber

To ensure scavenging of fumes from behind the sash. The baffle fastening system shall also serve as fixing points for Lab zone laboratory equipment scaffold.

Lightining

Light capable of providing illumination at the work surface shall not be less than 600 lux and shall be located outside the fume cupboard working chamber behind a 5mm non-opening transparent or translucent panel, sealed from the interior chamber. For cleaning or replacement a three-pin plug will be provided.

Sink

The work surface is to be provided with sinks as specified in the Fume Cupboard Schedule.

CONSTRUCTION

The fume cupboard shall not have wall cavities but be of single wall construction.

The chamber baffles and floor should be constructed from nonporous white rigid PVC.

The sash whether acrylic or glass must be suitable for the application.

Exhausts

Exhausts from different fume cupboards shall not be combined.

Fan

The fan shall be the made from corrosion resistant materials. The impeller shall be a one piece 36 blade polypropylene construction, statically and dynamically balanced. The casing shall be either PVC or Polypropylene. The pedestal shall be galvanized steel. A drain point shall be provided in the fan cowling.

Fan Stack Discharge

A velocity cowl shall be fitted to the fan stack discharge to maintain a maximum fume discharge velocity of 10m/s.

UPVC Duct Work

Fume cupboard exhaust ductwork and associated vertical discharge cowls shall be made from light weight corrosion resistant unplasticated polyvinyl chloride (UPVC).

Circular ductwork with a minimum wall thickness of 3mm should be used where possible. Blends should be of moulded radiused construction with a minimum wall thickness of 2.5mm. Duct is to be joined via spigots or sockets with an approved solvent or hot air welding. All fabricated ductwork whether rectangular or circular, shall be hot air welded. All joints are to be air tight.

Schedule Of Fume Cupboard

The attached schedule details the fit-out requirements of each fume cupboard.

Fume Cupboard monitoring

All fume cupboards shall be mounted on white powder coated steel under bench with under bench cupboard complete with doors, handles and shelves. Infill panels above the fume

Cupboard to ceiling shall be provided by the builder and adequate access panels shall be provided for servicing of fume cupboard light, controls and sash movement systems.

Variable Flow System

The fan inverter shall indirectly work off sash position and incorporate a pressure transducer that monitors the duct pressure and compares the sash position against the expected duct pressure and if the pressure is below what is expected (in any sash position) the fume cupboard shall initiate the emergency shutdown procedure. Provision may be made to boost the exhaust flow rate above the face velocity requirements

Fume Eliminator

The fume eliminator shall be for horizontal air flow and shall be the deflector blade design incorporating packs of hooked panel blades. Wide angle, full coverage, corrosion resistant water spray jets (with solenoid control valve) shall be positioned in the inlet duct of the eliminator.

Inlet and outlet transitions shall be designed to achieve good flow distribution across the whole blade area. Materials of construction should be either PVC or Polypropylene. Water to the jets shall be arranged to turn on and off with the exhaust fan.

Fume Wet Scrubber

The fume scrubber should be designed to either fit directly onto the top of the fume cupboard (to eliminate ducting) with the exhaust gases passing vertically up through the scrubber, or where site space restrictions do not permit the installation of a vertical scrubber on top of the fume cupboard, a horizontal roof mounted scrubber can be used. In either case the scrubber shall be of the packed design.

Vertical Scrubber:

The scrubber shall be designed to fit directly onto the fume cupboard exhaust outlet. The exhaust air shall pass up through a packed bed of polypropylene pall rings. Inlet and outlet transitions shall be designed to achieve good flow distribution across the whole section of the packed bed. Immediately above the packed bed will be wide angle, full coverage, corrosion resistant water

spray jets with solenoid valve. A transparent window shall be incorporated into the scrubber body, providing a view of the working water spray jets. Above the sprays there shall be a mist eliminator bank effective to 99%‡15 of the saturated exhaust air before it enters the duct. Water to the jets shall be arranged to turn off with the exhaust fan.

Spray Water Recirculating Tank:

A water recirculating tank shall be positioned on the underside of the fume eliminator or horizontal wet scrubber. The recirculating tank shall incorporate a level float connected to the water supply and a 240 Volt pump providing water at pressure to the sprays. The tank shall include an inspection window in its side.

Connection between the eliminator or scrubber and the water recirculating tank shall be by hand toggle clamps.

Baffle Wash down Facilities:

Water sprays shall be positioned behind the cupboard baffles and be arranged to activate on fan shut down. The cupboard work surface should have an integral trough/slump to collect wash down water.

Heat Shields

Heat shields or heat resistant baffles, where fitted, shall be designed to provide adequate protection

for the inner surfaces of the fume cupboard. They shall be easily removable for cleaning .Heat shields shall not be positioned so that they compromise the safe and efficient operation of the fume cupboard.

Thermal Detector

1-Operational Requirements

The cupboards shall be fume cupboards with exhaust systems comprising of PVC centrifugal fan, ductwork, discharge cowl and supports.

Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; BIOCHEMISTRY

Digital Analytical Balance

Digital Analytical Balance :

1. Microprocessor based single pan Analytical Balance with High accuracy & precision is required. Reading of the weight by digital display.

2. Weigh accurately up to 3rd decimal place

3. Auto self-calibration facility

4. Auto zero Setting, TARE setting

5. One touch calibration

6. Weighing capacity up to 120g

7. Readability and repeatability 0.001g

8. Stabilization time < 5 second

9. Liquid Crystal Display (LCD) for display

10. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; BIOCHEMISTRY

Balance Micro

Balance Micro :

1. Readability 1 μg (0.001 mg)

2. Weighing range 6 g

3. Reproducibility 0.0015 mg (0 to 0.2 g) 0.0025 mg (0.2 to 6 g)

4. Linearity +/- 0.007 mg

5. Stabilization time 8 Sec

6. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO..

Department ; BIOCHEMISTRY

Spectrophotometer

Spectrophotometer :

1. Basic mode: Absorbance, %T, concentration

3. Wavelength scanning mode

4. Kinetics mode – including time/drive

5. DNA/Protein analysés mode

6. Multi – wavelength

7. Performance validation mode

8. Display: LCD

9. Light source: Tungsten – Halogen and deuterium

10. Monochromator: Littrow type with 1200lines/mm grating

11. Detector: Two silicon photodiodes (and 2 spares)

12. Wavelength range (approximately): 190 – 1100nm

13. Wavelength accuracy: ± 0.3nm

14. Wavelength resolution: 0.1nm

15. Wavelength repeatability: ±0.05nm

16. High resolution Varaible spectral band pass ( 0.5, 1 ,2 4 nm)

17. Photometric range: -3.5 to +3.5 A, 0 –300%T and 0 – 9999conc

18. Photometric accuracy ,±0.002A(0to0.5A),±0.004A(0.5to1A

19. Scan speed: up to 2500nm /min (return 3000nm/min)

20. Noise drift: Less than 0.001A/hr after warm up

21. Noise: Less than 0.0001A @500nm OA

22. Baseline flatness: ± 0.002A

23. Stray light: Less than 0.05%T @220nm and 340nm

24. Cuvette volume: Maximum 1ml

25. No of cuvettes minimum 5 quartz Cuvette

Accessories

26. PC and laser printer should be provided User manual and service manual to be provided

27. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

28. Should be provided with Start-up Kit/Regent/ Calibrators/ Consumables/ Accessories are required performing the machine initial 500 tests or 3 Months test which is higher.

Department ; BIOCHEMISTRY

ELISA (Demonstration)

ELISA (Demonstration) :

1. Predilution facility should be present

2. Random access, provision for multiple assays.

3. Parallel sample pipetting should be present

4. 6 filters with a minimum range of 400-700 nm are required.

5. Minimum 10 measurements.

6. Volume of wash liquid dispensed variable & adjustable.

7. User friendly software with option for manual intervention

8. Residual volume/ well should be <6μl

9. User friendly software with option for manual intervention

10. Temp. range-room temp to 37 deg C +/- 1 C.37 deg

11. There should be no carryover of sample.

12. Password protection to prevent unauthorized person’s access to software.

13. Printer to provide printed reports of tests. Patient name, ID keyboard entry & individual report printouts in preset format.

14. Facility to store at least 50 assay protocols

15. Facility to program samples, standards and controls in replicates

16. Display of assay scheduling, start and finish times

17. Automatic Quality Control Equations like Westgard rules, Levy Jennings charts.

18. Curve fits like Cubic Spline, Sigmoid, Polygon, LogLog, 4-parameter, point to point, Linear, Quadratic, spline, lin/log.

Reader Specification

1. Dynamic Range 0.0-3.0 OD

2. Precision <2% CV (2-3.0 OD)

3. Linearity +/- 1% (<2.00 OD)

4. Accuracy ± (1% of the reading +0.005A from 0 to 1.5A) ± (2% of the reading +0.005A from 1.5 to 3.0A)

Washer Specifications

5. Manifold configuration 8-Way

6. Wash cycles possible – not less than five

7. Incubator Specifications –

8. Facility for shaking

9. 4 System Configuration Accessories, spares and consumables

System as specified-

10. All consumables required for installation and standardization of system to be given free of cost.

11. 20 nos. uncoated micro well plates to be supplied

Environmental factors

12. The unit shall be capable of being stored continuously in ambient temperature of 0 -50deg C and relative humidity of 15-90%

13. The unit shall be capable of operating continuously in ambient temperature of 10 -40deg C and relative humidity of 15-90%

Power Supply

14. Power input to be 220-240VAC, 50Hz fitted with Indian plug

15. Resettable overcurrent breaker shall be fitted for protection

16. Suitable voltage corrector/stabilizer

17. Suitable UPS with maintenance free batteries for minimum one-hour back-up should be supplied with the system.

18. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

19. Should be provided with Start-up Kit/Regent/ Calibrators/ Consumables/ Accessories are required performing the machine initial 500 tests or 3 Months test which is higher.

Department ; BIOCHEMISTRY

PCR Machine

PCR Machine :

1. It should be Tabletop model.

2. Complete system including basic system, essential accessories, the state-of- art computer workstation, acquisition and analysis software, startup kit inclusive of calibration standards etc.

3. Open system to accommodate Taqman, SYBR green and all other fluorescent dye-based chemistries.

4. Peltier based 96 well block

5. Standard optical 96 well plates, 0.2 ml strips, 0.2ml tubes compatibility

6. Minimum sample volume requirement -5μl

7. CCD camera with halogen/LED and at least five excitation and five emission filters

8. Multiplex inability upto five dyes in a single run,

9. Calibrated dyes at installation: FAM/SYBR Green, VIC/JOE, NED/TAMRA/Cy3, ROX/Texas Red®, and Cy5, Should offer flexibility in dye selection.

10. Facility to calibrate new dye within the wavelength range without addition of new filters

11. Passive reference dye ROX or any other calibrated dye and should be optional

12. Option for melt curve analysis

13. Temperature range 40C to 1000C

14. Sensitivity: Detection of 1 copy of template

15. Software applications: Comparative Ct, Standard Curve, Relative Standard Curve, Allelic Discrimination / SNP Genotyping, Plus/Minus, dissociation / melt curve

16. Voltage: 220 V /50Hz. All accessories should be provided with latest version of software and upgraded yearly with free of cost

17. The IQ, OQ and PQ of the instrument should be performed at the time of installation.

18. CE-IVD compliant along with the tools like security access, auditing and esignatures.

19. Each unit supply with suitable Online UPS with 30 minutes back up.

20. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

21. Should be provided with Start-up Kit/Regent/ Calibrators/ Consumables/ Accessories are required performing the machine initial 500 tests or 3 Months test which is higher.

Department ; BIOCHEMISTRY

Auto Analyzer

Auto Analyzer :  The system should be fully automated, random access, for routine chemistries

 The system should provide endpoint and Kinetic assay.

 The through put should be 300 tests or more per hour (without ISE).

 The minimum number of test parameters on board at a time should be 30.

 The minimum sample capacity should be 50 at a time.

 Minimum number of STAT samples at a time should be 15.

 The system should have preferential treatment for STAT samples.

 The sample loading: Continuous/sample tray.

 The reaction cuvette must be reusable and durable. the reaction cuvettes should be replaced at free of cost when ever required during the warranty period and CAMC period.

 Sample cups should be reusable.

 Reagent requirement should be less than 300μl.

 Required wavelengths, ranging from 340 to 750 or more. Wavelength measured by diffraction gratings.

 The system should have on board cooling system (refrigerated / Peltier).

 Should have on board quality control programme including Levy Jennings graph.

 Should have reagent storing capacity of 50 or more bottles in reagent 1 and reagent 2 together.

 Data Management- System should store calibration and quality Control data, and patient data for 10000 tests. Report should be generated in free format.

 Water plant of required capacity should be supplied along with the instrument and warranty includes the consumables.

 Free maintenance of the water plant including the resin replacement should be covered under the warranty period.

 Should be supplied with 10,000 sample cups.

 Should be floor model type or bench top with suitable stand.

|  | Should have facility for reading results on monitor and print out facility. 80 column dot matrix/ Laser printer should be supplied along with the unit. |
| --- | --- |
|  | Should provide 1000ml of reagent for urea, uric acid, creatinine, sugar, Cholesterol, TP, albumin, SGOT, SGPT, alkaline phosphatase, Bilirubin (Total & Direct) calibrator 3mlx5 and quality control 5mlx1 -5 Nos each for normal and abnormal. |
|  | The pack size and rate of one pack of reagent for Year 1, 2, and 3 (Period will be starting from the date of price bid opening) for - urea, Uric acid, Creatine, Sugar, Cholesterol, TP, Albumin, SGOT, SGPT and Alkaline phosphatase, Bilirubin-Total and Bilirubin Direct shall be quoted separately which will be taken for evaluation as per the amended sample price calculation. |
|  | Should be supplied with variable micropipette 10 to100μl and fixed 500μl micropipette– 5 Nos each. |
|  | Refrigerator (300L) with temperature display and stabilizer should be provided by the supplier. |
|  | Should work on 200 - 240Vac 50Hz power supply. |
|  | Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO. |
|  | Should be provided with Start-up Kit/Regent/ Calibrators/ Consumables/ Accessories are required performing the machine initial 500 tests or 3 Months test which is higher. |

Department ; BIOCHEMISTRY

Hot Air Oven (More than 200 litres)

Hot Air Oven (More than 200 litres) :

1. Double Walled, inner wall made of aluminium and outer wall Mild Steel sheet.

2. Shelves: 3 to 5 stainless steel, perforated shelves, could be adjusted at any levels.

3. Heating Elements: placed at the bottom or the sides.

4. Temperature control: Thermostatic control from 50 to 250 o C with Regulator and digital temperature-indicator.

5. Automatic timer 0 - 60 minutes.

6. Air circulating fan.

7. Volume : 200 L

8. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; BIOCHEMISTRY

Water Baths

Water Bath :

1. Should have a double walled construction.

2. The inner chamber and top lid should be made of stainless steel.

3. The space between the two walls should be packed with thick glass wool.

4. Should provide with a thermostat control.

5. Working temperature should be from ambient+5 ºC to 80ºC having an accuracy of +1 ºC

6. Should have an approximate inner chamber dimension of 350mm x 250mm x 125mm.

7. Should be operated on 230V, 50 Hz single phase AC supply

8. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; BIOCHEMISTRY

Incubator 37oC

Incubator 37oC :

1. Capacity: 100 L or more

2. Interior chamber: SS 316 for easy cleaning and decontamination

3. Timer: 1 min. to 100 hours and hold position

4. Minimum turbulence and no cross contamination

5. Adjustable safety thermostat for temp setting at 1 deg C increment

6. Temp Accuracy : +/-1%

7. Internal glass door for the observation

8. With minimum two adjustable shelfs

9. Audiovisual Alarm to Indicate when temperature deviates more than 1°C from setpoint, and when program or time has finished. Equipment should have provision to mute the alarm.

10. Peltier heating with continuous air circulation and Heating by natural/forced convection for homogenous temperature distribution

11. Temperature range: +5° C to 60°C

12. There should be a Membrane Keypad with LCD/LED Display to set and display operating parameters, current status, running time and alarm conditions for time and temperature.

13. The Equipment should have Interior lighting facility, and insulated door fitted with heavy hinges

14. Power Supply:230V,50 Hz

System Configuration Accessories, spares and consumables

a. Incubator as per specification

b. All consumables required for installation and standardization of system should be provided free of cost

c. Should be supplied with all the accessories required for the functioning of the equipment.

d. Power Cord

Standard, Safety

 Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; COMMUNITY MEDICINE

Centrifuge Clinical.

Centrifuge Machine (16 tube) :

1. Should have a maximum speed of 5000 RPM with steeples regulator

2. Should be supplied with safety lid and lock. (Brushless Motor)

3. Should have digital speedometer and timer.

4. Should have imbalance detector and automatic cut off.

5. Standard mark on quality and safety is Compulsory.

6. Should work on a220VAC +/- 10 %, 50Hz AC Supply.

7. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; COMMUNITY MEDICINE

Baby Weighing Machine.

Digital Weighing Machine for Baby :

1. Tabletop, light and portable.

2. Built in rechargeable battery.

3. Easy to clean baby tray (acrylic).

4. Zero weight adjustment facility.

5. Quick, clear digital read outs.

6. Measurement does not change with position of baby on the pan.

7. Provision to measure the height of the baby in its laying position.

8. Accuracy: 5g, resolution: 1g, limit: 10gm to 15kg.

9. Auto setting to 0.00 once a the machine is switched on or when no external weight has been put on.

10. 4XAA battery(rechargeable) or equivalent; one hour backup.

11. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

12. Should have model approval from Legal Metrology Dept., Govt. of India.

Department ; COMMUNITY MEDICINE

Otoscope.

Otoscope :

1. 2.5 V(White Led Light) Pocket Scope

2. Auriscope With Aa Battery Handle

3. Wide-Angle Viewing Lens Allows For Instrumentation Under Magnification.

4. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; COMMUNITY MEDICINE

Ophthalmoscope.

Ophthalmoscope (Direct) :

1.   Available with LED/Halogen light

2.   Magnification up to x15 from direct vision to maximum magnification.

3. Red-free, blue and polarization filters and Anti-reflection lens.

4. Should have small and large spot sizes, fixation targets, slit aperture, hemispot and cobalt blue filter.

5. Should be rechargeable battery with Charger / battery/ mains operated.

6. At least 3 apertures and fixation star.

7. Range of lenses not smaller than - 30D to +20D with steps not greater than 1D.

8. Dust free sealed optics and aspherical optical system.

9. Internal batteries, rechargeable preferred compatible with both 2.5 V and 3.5 V batteries or handles provided; Led display indicating the charging status.

Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; COMMUNITY MEDICINE

Incubator, Electric.

Incubator, Electric :

1. Capacity: 120 L

2. Interior chamber: SS 316 for easy cleaning and decontamination

3. Timer: 1 min. to 100 hours and hold position

4. Minimum turbulence and no cross contamination

5. Adjustable safety thermostat for temp setting at 1 deg C increment

6. Temp Accuracy : +/-1%

7. Internal glass door for the observation

8. With minimum two adjustable shelfs

9. Audiovisual Alarm to Indicate when temperature deviates more than 1°C from setpoint, and when program or time has finished. Equipment should have provision to mute the alarm.

10. Peltier heating with continuous air circulation and Heating by natural/forced convection for homogenous temperature distribution

11. Temperature range: +5° C to 60°C

12. There should be a Membrane Keypad with LCD/LED Display to set and display operating parameters, current status, running time and alarm conditions for time and temperature.

13. The Equipment should have Interior lighting facility, and insulated door fitted with heavy hinges

14. Power Supply:230V,50 Hz

15. System Configuration Accessories, spares and consumables

16. Incubator as per specification

17. All consumables required for installation and standardization of system should be provided free of cost

18. Should be supplied with all the accessories required for the functioning of the equipment.

19. Standard, Safety

20. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; DENTISTRY

Dental X-Ray-Intra Oral with RVG

Dental X-Ray-Intra Oral with RVG :

1. RVG is used for digital dental X-rays which can be instantly viewed and evaluated with minimal radiation exposure

2. High resolution RVG based on CCD/CMOS technology with optical fiber.

3. Maximum reduction in patient radiation as compared to X-ray film.

4. Should have easy positioning of sensor inside the mouth

5. Should supply sensor with minimum active area 600mm².

6. Should have a Pixel size: 20 μm x 20 μm

7. Size of the sensor in mm x mm with tolerance +/-2mm. Adult 43mm x 31mm & pediatric 40mm x 30mm.

8. Sensor cable length (meters) 3.5m - 4m.

9. Thickness of the sensor should be less than 5 mm.

10. Spatial resolution approx. 20-25-line pairs/mm.

11. Computer with LCD/LED color monitor 20-inch screen, latest processor, HDD 1TB, 8 GB RAM & Laser Printer & 60 minutes backup Online UPS.

12. Intra oral X-ray unit should be DC based. (8mA/65-70 kv) with option for wall mount/ mobile stand.

13. Should have positioning devices a. Bitewing b. Periapical c. Endodontic.

14. X-ray unit should be supplied with lead apron, thyroid collar.

15. Power input to be 220-240VAC, 50Hz fitted with Indian plug

16. Servo Voltage stabilizer of appropriate ratings

17. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; DENTISTRY

Dental Instruments

| Dental Instruments List : |  |
| --- | --- |
| Extraction Forceps (Set of 7 ) | Two sets |
| Elevators (Cryers, Straight, Coupland's, Apexo, Cross bars, Lyndo levien) | Two sets each |
| Periosteal elevators | 5(five) pieces |
| Mirror handle | 20(twenty) pieces |
| Miror top – 4 No. & 5 No. | 50 (fifty) |
| Probe – Curved & Straaight | 20(Twenty) pieces |
| Tweezer | 20(Twenty) pieces |
| Scissors- Small & Medium | 5 (five) pieces |
| Needle holder – Medium | 5 (five) pieces |
| Bone rongeur | 2 (two) pieces |
| Bone file | 2 (two) pieces |
| Kidney tray – Medium | 10 (ten) pieces |
| Artery forceps – Medium | 5 (five) pieces |
| Wire cutter | 2 (two) pieces |
| Austin's retractor | 5 (five) pieces |
| Howarth's elevator | 5 (five) pieces |
| Ehrich arch bar | 100(hundred) pieces |
| Stainless steel wire – 26 SWG | 100(hundred) spools |
| Stainless steel wire – 24, 26 | 20 each |
| Surgical burs -702 & 703 | 50 each |
| Autoclave apparatus | 1 unit |
| Instrument lifters | 4 pieces |
| Probe EXD57 | 5 pieces |
| Enamel Hatched | 2 pieces |
| GMT | 2 pieces |
| Burs (Round,tapered, straight, inverted cone | 10 pieces each |
| Flat end tapered bur | 10 pieces |
| Cement carrier | 2 pieces |
| Excavators IN31011 |  |
| Endodontic excavators IN31051 | 2 pieces |
| Crown cutting burs | 20 sets |
| Micro motor burs Round | 10 sets |
| Composite filling instruments | 2 sets |
| Composite finishing kit | 2 sets |
| Endodontic access bur | 50 pieces |
| Paseo reamers | 10 pieces |
| GG drills 1-6, 31mm | 2 sets |
| Endo block | 5 sets |
| Broaches | 10 boxes each |
| K files 21mm 15-40,25mm 40-80 | 10 boxes each |
| Finger spreader | 2 pieces |
| Cement spreader | 2 pieces |
| Bur stand | 2 units |
| Locking tweezer | 2 pieces |
| Haemostat straight | 2 pieces |
| Endomotor | 1 unit |
| Apex locator | 1 unit |
| Fixed Orthodontic Instruments | 1 set |
| Edentulous Trays | Two (2) |
| Dentulous Trays | Two (2) |

Department ; DERMATOLOGY, VENEROLOGY & LEPROSY

Hyfrecator Electrosurgical instrument

Hyfrecator/ Electrosurgical instrument :

1. Remote Control Handpiece: change power settings to maintain the sterile field while at the same time decreasing procedure time, power up/down hand switching pencil with 10 foot cord

2. Easily switch back and forth between settings depending on the procedure being performed Store up to 3 power settings (one per mode)

3. Continuously monitors the system for reliability and patient safety Self Diagnostic

4. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; DERMATOLOGY, VENEROLOGY & LEPROSY

Cryotherapy with liquid Nitrogen

Cryotherapy with liquid Nitrogen :

1. Liquid Nitroge cryo system

2. High Quality Dewar Flask.

3. Elegant and stable Base

4. Light weigh easy to use with special ACNE spray

5. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; DERMATOLOGY, VENEROLOGY & LEPROSY

Iontophoresis Machine

Iontophoresis Machine :

1. Safe medical procedure that involves sending mild electrical currents through water and into the skin

2. To treat sports injuries by delivering anti-inflammatory medications directly into the skin

3. Adjustable current strength

4. No electric shocks (electric fence effect has been eliminated

5. Compact and lightweight (only 2.5 kg) ->Smart carrying case

6. Advanced direct delivery technology. Carrying case and Aluminium electrodes are not required

7. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; EMERGENCY MEDICINE

ICU Ventilator)

| 1. | Should have facility for Invasive and Non-Invasive ventilation. |  |  |
| --- | --- | --- | --- |
| 2. | Microprocessor Control suitable for adult ventilation. |  |  |
| 3. | Electromagnetic Compatible Hinged arm holder for holding the circuit. |  |  |
| 4. | Should be able to record and analyse various parameters. Breath to breath / pulmonary functions, loops to be stored in the memory with feasibility of trend analysis on a TFT touch screen. Screen size should be more than 10 to 15” so that be easily seen from distance. Compiled trend analysis at Minimum 48 hours for all measured parameters. Monitoring during mechanical ventilation includes measurement of peak and plateau pressures, intrinsic positive end-expiratory pressure, and work of breathing. |  |  |
| 5. | Machine should be Compressed air / oxygen driven / turbine driven. |  |  |
| 6. | Should have Nebulization assembly facility. |  |  |
| 7. | Ventilator, Compressor & Humidifier should be Same Trolley / cart mounting for easy transpo |  | rtation. |
| 8. | Should have Internal rechargeable battery at least 05 Hrs. backup. |  |  |
| 9. | Automatic Self - test compliance and leakage compensation for circuit and ET Tube. |  |  |
| 10. Should have the following modes: |  |  |  |
| a. Volume Controlled, CMV, PCV or IRV, SIMV (volume cycled & pressure limited), |  |  | SIMV+PS |
|  | (volume cycled & pressure limited), CPAP, Bilevel, APRV. IMPRV, PRVC, AUTOMODE, |  | Minimum |
|  | minute ventilation, Mandatory rate ventilation, Proportional assist, Mandatory ventilation. |  | minute |
| 11. Should have the following setting. |  |  |  |
| a. |  | Tidal Volume: minimum 200ml maximum of 2000 ml in Volume control is essential. |  |
| b. |  | PEEP 0 to 30 cmH20 or more |  |
| c. |  | Oxygen Concentration 21 –100 % |  |
| d. |  | Inspiratory Pressure 1-80 cmH20). |  |
| e. |  | Respiratory rate 1 to 80 bpm. |  |
| 12. Alarm |  |  |  |
| a. |  | Adjustable Alarm. - Low/high minute volume, low/high pressure, low/high tidal volume, low/high rate, apnea time, low/high oxygen, low/high SpO2 |  |
| b. |  | Special alarm - O2 cell Failure , flow sensor, battery, power supply, gas supply, oxygen concentration, |  |
| 13. Compressor |  |  |  |
| a. |  | Should be supplied with External Compressor. |  |
| b. |  | The compressor has been designed to supply the ventilator with dry, filtered compressed air. |  |
| c. |  | Compressor should be oil-free. |  |
| d. |  | Portable & fitted with ventilator cart. |  |
| e. |  | Air filtration 5 microns. |  |
| f. |  | Noise level dB 40 – 50. |  |
| g. |  | Peak flow of 200 lpm. |  |
| 14. Humidifier |  |  |  |
| a. |  | Servo controlled heated Respiratory Humidifier. |  |
| b. |  | Display Should be of LED /LCD. |  |
| c. |  | Temperature control settings & Temperature range: 28-40 deg. |  |
| d. |  | Temperature should be adjustable. |  |
| e. |  | Jar should be autoclavable |  |
| 15. Standard Accessories / spare & Consumable. |  |  |  |
| a. |  | Silicon breathing circuit (Adult reusable) - 5 complete set. |  |
| b. |  | Nebulization assembly compatible circuit 5 complete set. |  |
| c. |  | Humidifier - 1 No. |  |
| d. |  | O2 Pressure Regulator with hose - 1 No. |  |
| e. |  | Hose for O2 connection with connector - 5 mts. |  |
| f. |  | Hose for compressed air with connector - 5 mts. |  |
| g. |  | Test lung - 1 No. |  |
| h. |  | Non-invasive ventilator mask reusable for adult (3sizes) - each size 5 No. |  |
| i. |  | ET tube cuff pressure monitor and HME filter – 10 no |  |
| j. |  | Inbuilt / integrated nebulizer-1 No. |  |
| k. |  | All sensors and other non-consumable items (other than reusable silicon ventilator circuits) should be free of cost during warranty and CMC. |  |
| 16. Ventilator, Humidifier & Compressor Power Supply input to be 200-240VAC, 50 Hz fitted with Indian conditions plug. |  |  |  |
| 17. Suitable online UPS with commensurate capacity for all ventilators including compressor & Humidifier with maintenance free batteries for minimum one-hour back-up should be supplied. |  |  |  |
| 18. Ventilator, Humidifier & Compressor Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO. |  |  |  |
| NOTE: |  |  |  |
| 1) Reusable consumables (other than reusable silicon ventilator circuits) should last during the warranty period. |  |  |  |
| 2) Ventilator & Humidifier any additional reusable consumables are required during the warranty period those will be supplied free of charge by the supplier. |  |  |  |
| 3) The life expectancy of the reusable consumable is expected to be of at least one year from the date of installation of the same. The reusable consumables will be procured at the prices accepted as per the contract. |  |  |  |
| 4) The bidders should submit all reusable consumable items price & their authorised local office/ distributor name in the financial bid. |  |  |  |

Department ; EMERGENCY MEDICINE

Point -of care laboratory for quantitative estimation of cardiac enzymes, ABG and electrolytes

Point -of care laboratory for quantitative estimation of cardiac enzymes, ABG and electrolytes :

1. For this Purpose each of the following equipment is required (one in number) the specifications of which are as following:-

2. cardiac marker machine:- Should be able to measure the high sensitive quantitative measurement of Troponin-T / Troponin I.

3. Should be at least a third[1]generation assay. It should be a portable unit and have an inbuilt battery backup.

4. It should have a parameterwise QC checking facility. Should have a memory capacity of a minimum of 500 test results with date time and comments.

5. Should have memory storage for results of minimum 200 cartridges.

6. Should have a minimum sample volume of approximately 300 microlitres.

7. Should have an operating temperature ranging from 18- 30 degrees Celsius.

8. Should work on 200-240Vac 50Hz power supply.

9. It should have an internal printing facility or supply with an external thermal printer.

10. Should be supplied with 50 strips of Troponin T/I Parameter.

11. Should be provided with a calibration certificate issued by the manufacturer at the time of installation.

12. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

13. Test parameters must be incorporated bellows specification

14. Test parameters: CKMB, Myoglobin, Troponin T and NT-pro BNP Infection Marker: CRP, Coagulation Marker: D-dimer , Pregnancy marker: beta Hcg, Sepsis Marker: Procalcitonin( PCT) & customizable test selection

2. ABG machine :-

1. Compact system for measuring pH, pCO2, pO2, - HCO3 & four Electrolytes like Na, K, Ca+ and Cl - in blood.

2. All should be measured in a single injection / aspiration of Sample.

3. May have provision of modular platform for future up gradation to include glucose, lactate & hemoglobin the same machine with the inspiration of single sample.

4. Should be able to analyze all parameters using low blood volume directly from syringe or capillaries.

5. Fast and accurate result of test made available in about 60 seconds.

6. Automatic Calibration by liquid calibrators with flexible time mode. Instrument should have Stand- by mode facility and Economy mode.

7. It should not be cartridge based system.

8. Startup Kit, Calibrators, Consumables, Accessories and spares required performing initial 500 tests.

9. All the consumables and spares should be quoted separately unit wise.

10. Compatible online UPS with battery back up of at least one hour

11. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

12. Should be provided with Start-up Kit/Regent/ Calibrators/ Consumables/ Accessories are required performing the machine initial 500 tests or 3 Months test which is higher.

Department ; EMERGENCY MEDICINE

ACLS, BLS and Airways Mannequins (Child and Adult)

ACLS, BLS and Airways Mannequins (Child and Adult) :

For adult it should consist of a full-body adult mannequin with CPR and advanced airway capability and for child it should consist of pediatric manikin with CPR and advanced airway capability.

The mannequin should have the following features and items :-

1. All parameters of CPR should be electronically measured and recorded with a wireless/wired device with details about the rate, depth, the effectiveness of compression and ventilation.

2. Should be capable of interfacing with a clinical manual defibrillator and AED for rhythm analysis and electrical therapy including pacing, synchronized shock and defibrillation.

3. Should be supplied with a monitor to display ECG, SPO2, BP, ETCO2, and other waveforms with supportive software.

4. Should have the capability to display oesophageal intubation

5. Should have an intubation facility and display effective ventilation.

6. Should have adjustable chest compliance

7. Iv arm for insertion of a cannula, the ability for detection of pulse, the ability for IO line placement

8. Should have software and hardware support to mimic and demonstrate various ECG patterns, Breathing sounds,and ability to program scenarios.

9. Should have an effective feedback system that can be demonstrable, results storable and retrievable for debriefing

10. The software support and audiovisual support for the above purpose should be supplied allowing upgradability to the universal guidelines from time to time and with no additional cost.

11. Should run on KSEB power source

12. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; EMERGENCY MEDICINE

Rapid autoclave machine

1. Portable High Speed microprocessor-controlled rapid autoclave suitable for sterilization of unwrapped instruments, Wrapped instruments, Packs and special cycle for liquids.

2. Have 2 or more programmable cycles to allow custom creation of different cycle parameters for special sterilization needs. The programmable cycle should allow change of time, temperature, Dry time and Vent.

3. Pre-programmed cycles should have selection of 2 temperatures (132 0C and 1210C) and pressures (186 kPa and 104 kPa).

4. Have a microprocessor-based fault detection circuit for monitoring all functions of sterilizer during a cycle, giving necessary signals to alert operator.

5. Should be fully automatic and have a drying cycle.

6. Should meet the requirements of ASME Boiler and Pressure Vessel Code.

7. Chamber size should be 27 cm dia or more and chamber depth should be 45 cm or more.

8. Equipment should be compact, not larger than 58 cm in length, 46 cm width, and 46 cm height

9. Printer should be provided for recording cycle, cycle time, temperatures and pressure.

10. Should have 3 Air removal purges, prior to sterilization.

11. Should have Front filling and Drain facility behind front door, for safety.

12. Should incorporate all safety features possible, including fault detection circuit, continuous monitoring of chamber temperature to prevent overheat condition.

13. Display should indicate cycle selected, cycle temp and exposure time for selected cycle. During cycle, display should reflect messages indicating status of cycle and errors.

14. During sterilization mode, LCD display should show --remaining cycle time, temperature and pressure.

15. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ;    EMERGENCY MEDICINE

Portable Suction Machine

Portable Suction Machine :

1. Vaccum /LPM : - 700 MM Hg , 50 Litres/Min

2. Pump Type- Double rotary vane type

3. flutter free vacuum control knob,

4. Wide mouthed 2 x 2 LITRE (light weight, unbreakable and clear) with self-sealing bungs and mechanical over flow safety device.

5. Noise (in dBA)- 50 dB

6. Collection container & its cap, suctions tube tips, a vacuum gauge, two sets of moisture & microbial filters and control knob

7. SiliconeTubing:8 mm ID x 2 mtr (PVC), 2x2 lt jar (one set

8. It should be Mobility, portability.

9. Equipment shall operate on 220V-240V, 50 Hz, single phase electric supply.

10. ISO/ BIS approve model should be offered.

Department ;    EMERGENCY MEDICINE

Oxygen Concentrator

| Specification | Specification Name |  | Bid Requirement (Allowed Values) |
| --- | --- | --- | --- |
| Flow Rate | Flow Rate (Minimum) at 93% + - 3% oxygen purity (Litres per minute) |  | 1 - 10 |
| Oxygen outlet and pressure and tank capacity | Oxygen outlet(s) |  | Single, Single or Dual, Others, Dual |
|  | Oxygen tank capacity (Ltrs) |  | 5.0 - 25.0 Or higher |
| Power Source and Battery | Power Source |  | Electrical cum Battery Operated |
|  | Battery duration in hours (If Battery operated or Electrical cum Battery operated) |  | 3, 6, 9, 12 Or higher |
| operating temperature humidity and altitude | Operating altitude Range up to (in meters) above sea level |  | 0 - 10000 |
| Type of Operation | Type of operation |  | Pulse Dose Cum Continuous Flow Type |
| Air Delivery | Adjustable Air Delivery mechanism |  | Yes |
|  | Continuous flow rate (Litre per Minute) up to |  | 25 |
| Noise Level | Noise Level (dB) |  | less than 10 dB silent type, Less than 40 db, Less than 60 db |
| Alarms and Alerts | Digital Display |  | LCD |
| Weight and Dimensions | Weight (Kg) |  | Up to 35 |
| Flame Retardancy | Flame Retardant |  | Yes, No |
| Warranty | Manufacturer warranty Period (Years) |  | 3 Or higher |
| specification | Continious flow setting |  | 0.5 LPM, 1 LPM, 1.5 LPM, 2 LPM |
| Specification Parameter Name |  | Bid Requirement (Allowed Values) |  |
| Quality Certificate |  | Manufacturer should be ISO 13485 certified and product should be USFDA/EU CE (From Notified Body/CE/BIS / ISO |  |

Department ; FORENSIC MEDICINE & TOXICOLOGY

Digital Spectrophotometer

Digital Spectrophotometer :

1. Basic mode: Absorbance, %T, concentration

2. Quantitative mode: Concentration calibration curves

3. Wavelength scanning mode

4. Kinetics mode – including time/drive

5. DNA/Protein analysés mode

6. Multi – wavelength

7. Performance validation mode

8. Display: LCD

9. Light source: Tungsten – Halogen and deuterium

10. Monochromator: Littrow type with 1200lines/mm grating

11. Detector: Two silicon photodiodes (and 2 spares)

12. Wavelength range (approximately): 190 – 1100nm

13. Wavelength accuracy: ± 0.3nm

14. Wavelength resolution: 0.1nm

15. Wavelength repeatability: ±0.05nm

16. Bandpass: 1.8nm

17. Photometric range: - 0.3 to 3.0 A, 0 – 200%T and 0 – 9999conc

18. Photometric accuracy: Better than ±0.005A @1A

19. Scan speed: up to 2500nm /min (return 3000nm/min)

Digital Spectrophotometer :

20. Noise drift: Less than 0.001A/hr after warm up

21. Noise: Less than 0.0001A @500nm OA

22. Baseline flatness: ± 0.002A

23. Stray light: Less than 0.05%T @220nm and 340nm

24. Cuvette volume: Maximum 1ml

25. No of cuvettes minimum 5 quartz Cuvette

26. Accessories

 PC and laser printer should be provided User manual and service manual to be provided 27. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

 Should be provided with Start-up Kit/Regent/ Calibrators/ Consumables/ Accessories are required performing the machine initial 500 tests or 3 Months test which is higher.

Department ; FORENSIC MEDICINE & TOXICOLOGY

Chemical Balance

Chemical Balance:

1. Microprocessor based single pan Analytical Balance with High accuracy & precision is required. Reading of the weight by digital display.

2. Weigh accurately up to 3rd decimal place

3. Auto self-calibration facility

4. Auto zero Setting, TARE setting

5. One touch calibration

6. Weighing capacity up to 120g

7. Readability and repeatability 0.001g

8. Stabilization time < 5 second

9. Liquid Crystal Display (LCD) for display

10. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; FORENSIC MEDICINE & TOXICOLOGY

Distilation Plant

Distilation Plant :

1. Distillate output Production rate- 4 litres/hr

2. Conductivity, 0.2μs/cm

3. Minimum water supply pressure, 20.3kg/cm

4. Water consumption, 60litres/hr

5. Electrical supply 230/240V, 50/60Hz 13 amp

6. Spare heating elements - 1set

7. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; FORENSIC MEDICINE & TOXICOLOGY

Centrifuge

Centrifuge :

1. Should have a maximum speed of 3000 RPM with steeples regulator

2. Should be supplied with safety lid and lock. (Brushless Motor)

3. Should have digital speedometer and timer.

4. Should have imbalance detector and automatic cut off.

5. Standard mark on quality and safety is Compulsory

6. Should work on a220VAC +/- 10 %, 50Hz AC Supply.

7. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; FORENSIC MEDICINE & TOXICOLOGY

Spectroscopic Lens with adjustable Slit

Spectroscopic Lens with adjustable Slit :

1. Resolution 4nm (40 Angstroms)

2. Fraunhofer Line Range 390.0nm to 760.0nm (visible spectrum), should have graduation scale

3. Dispersion Angle 24° (approximately

4. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO /ISO Approved

Department ; FORENSIC MEDICINE & TOXICOLOGY

Cold storage for dead bodies

Cold Storage for Dead Bodies :

1. Corrosion free interior and exterior.

2. Audio visual alarm for high and low temperature.

3. Designed for long storage of cadaverous.

4. PUF insulation on all sides.

5. Special design ensuring best hygiene with washing & draining facility.

6. Reliable

7. Special loading trolley

8. Energy efficient and sturdy construction

9. Light weight.

10. Digital temperature indication.

11. Low maintenance.

12. Microprocessor based / PLC temperature control.

13. Double walled cooling units.

14. Outer body of the mortuary chamber is constructed out of thick S.S sheets. The inner chamber made of heavy gauge stainless steel sheet of SS-304 grade. The 100mm gap between the walls filled high grade poly urethane insulation, which ensures maximum thermal efficiency.

15. The doors connected by very sturdy chrome plate hinges and fitted with hard chrome plated lubricated latches for opening of the door.

Cold Storage for Dead Bodies :

16. The doors made of galvanized steel sheets, lined with stainless steel for extra protection and long life.

17. All the doors fitted with high quality neoprene rubber gaskets for airtight fittings with very sturdy casters.

18. CFC free compressors, conforming to latest international standards and guidelines. Twin compressors of which one is standby.

19. Vapor proof lamp inside.

20. Temperature range -2 to 4 deg C with temp failure alarms.

21. Suitable Voltage automatic stabilizer O/P 230 +/-10% I/P 150 – 280Volts.

22. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; FORENSIC MEDICINE & TOXICOLOGY

Weighing Machine for dead bodies

Weighing Machine for Dead Bodies :

1. Capacity should have 200 kg with Should have digital display of weight

2. Should have extra channel structure to provide equal Load distribution

3. Should have standard Mild Steel checkered Top cover and Back grill for support

4. Should have Over Load stoppers and Transport Protection

5. Rocker pin foot assembly to provide stability and durability

6. Four Load Cell system - Single point beam Load Cell

7. Optional Weighing Indicators

8. Should have easy access junction box

9.   Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; FORENSIC MEDICINE & TOXICOLOGY

Autopsy Tables

Autopsy Tables :

1. Approximate Dimension:-1820 mm X 600 mm X 900 mm (LxWxH)

2. It should be made of stainless steel (steel grade 304) with a frame made of rugged torsion- resistant stainless steel profiles.

3. It should have 4 stainless steel swivel locking castors.

4. Tabletop depth should be of approx. 15mm sloping towards the drain.

5. Large radii on all inside corners should be provided for easy cleaning.

6. 10 litre removable container with bayonet lock, mounted beneath the downspout, should be attached to a rack in the base frame.

7. Airtight compartment should be mounted beneath the table top to serve as an odour-free storage of drapes.

8. It should have stainless steel full extension drawer and a removable stainless steel tray provided with a perforated plate and a removable lid

9. System Configuration Accessories, spares and consumables Stainless Steel Bucket 50 Ltrs Headrest Body support shim

10. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; FORENSIC MEDICINE & TOXICOLOGY

Autopsy Saw with Accessories

| Autopsy Saw with Accessories :  1 Large Section Blade |  |
| --- | --- |
|  | Allen wrench |
|  | foot cord with hospital grade plug |

        Weight with power cord 3.7 lbs

        BIS / ISO/ISO Approved model should be offered.

Department ; FORENSIC MEDICINE & TOXICOLOGY

Weighing Machine for Organs

Weighing Machine for Organs :

1. Platform size 350 mm x 350 mm approx.

2. SS 304 grade construction

3. Complete ss Platform for easy cleaning and anti-staining

4. Maximum of 15 kg

5. Accuracy up to 2 gms

6. Rechargeable battery back-up pack provided for usage in power failure

7. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; FORENSIC MEDICINE & TOXICOLOGY

Weighing machine for Fetus

Weighing machine for Fetus :

1. Platform size 350 mm x 350 mm approx.

2. SS 304 grade construction

3. Complete ss Platform for easy cleaning and anti-staining

4. Maximum of 15 kg

5. Accuracy up to 2 gms

6. Rechargeable battery back-up pack provided for usage in power failure

7. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; GENERAL MEDICINE

Ophthalmoscope:

1. Available with LED/Halogen light source.

2. Magnification up to x15 from direct vision to maximum magnification.

3. Red-free, blue and polarization filters and Anti-reflection lens.

4. Should have small and large spot sizes, fixation targets, slit aperture, hemi-spot and cobalt blue filter.

5. Should be rechargeable battery with Charger / battery/ mains operated.

6. At least 3 apertures and fixation star.

7. Range of lenses not smaller than - 30D to +20D with steps not greater than 1D.

8. Dust free sealed optics and aspherical optical system.

9. Internal batteries, rechargeable preferred compatible with both 2.5 V and 3.5 V batteries or handles provided; Led display indicating the charging status.

Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; GENERAL MEDICINE

Flexible Video End Viewing Oesophago-Gastroduodenoscope

Flexible Video End viewing Oesophago-Gastroduodenoscope :

1. Direction of view should be zero degree. Minimum of 100 degree (app) of field of view. Range of observation from 5 mm to 90 mm.(app)

2. Angulations of tip not less than 200 deg(Up) and 90 deg (down) with right to left movement of minimum 100 deg.(app)

3. Insertion tube outer diameter of more than 8 mm with a working length of not less than 1000 mm. Distal end of less than 8 mm.

4. Instrument channel of more than 2.3 mm Compatible with the video system specified. Video processor with light source & Monitor Power supply 200-240 V A/C

5. PAL type video signal. The camera should be 3 chips CCD with high definition (HD) Output with provision of recording.

6. Controls for color adjustment, to enhancement and balance settings.

7. Controls to freeze images and Zoom Image

8. Operates on Xenon/LED lamp (for Xenon-150 W) Should have an Emergency changeover lamp. 19” LCD color monitor with XGA resolution

9. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO..

Department ; GENERAL MEDICINE

Flexible Video Colonoscope

Flexible Video Colonoscope :

1. Video colonoscope – insertion tube dia 13.2 mm, instrument channel 3.8 mm,angle of view 140

2. deg., focal range -4 to 100 mm, Angulation Up/Down 180/180 deg. R/L 160/160 deg. with 1700 mm workding length, total length 2023 mm

3. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; GENERAL MEDICINE

Proctoscope

Proctoscope :

1. Working Attachment, for use with Proctoscope and Telescope Proctoscope with obturator O.D.: 24 mm, working length: 8 cm.

2. Straight Forward Telescope 30°, Eyepiece 45° angled, diameter 4 mm, length 9.5 cm, autoclavable, Fiber optic light transmission incorporated

3. Sponge Holder, working length 20 cm Dressing Forceps, total length 19 cm

4. Fistula Hook, total length 19 cm

5. Proctoscopy Punch, Through-cutting, cutting width 3.4 mm, straight jaws, sheath diameter 3.5 mm, working length 20 cm

6. Injection Needle, straight, LUER-Lock, tip diameter 1.0 mm, working length 14 cm Hemorrhoid Grasping Forceps, for use with ligature instrument

7. Ligature Instrument, for treatment of hemorrhoids, working length 17 cm including: Loading Cone,

8. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; GENERAL MEDICINE

Fiber optic bronchoscope (May be shared with TB&CD)

Fiber optic bronchoscope :

1. Flexible Fibre Optic bronchoscope

2. Cold Light Source

3. Camera

4. Monitor

5. Video recording digital, either USB based video recording device or through external sorftware

6. Trolley

7. The working length of the fibre scope should be 60 cms to allow mounting of Endotracheal tube on the scope before performing Bronchoscopy.

8. The outer diameter should be less than 3.7 mm and allow mounting of tube size less than 4.5 mm to the scope, lesser diameter preferable. Suction channel is a must.

9. Range of bending at the tip should be minimum 180 degree up and 130 degree down approx.

10. Compatible light source. Automatic light adjustments to maintain optimum brightness. High efficiency halogen source should be suplied with compatible fiberoptic cable

11. Camera processor should give at least two output

12. a) S-video / DVI

13. b) Composite video

14. Compatible single chip camera system to give clear and big view of structures.

15. At least 14” colour monitor display.

16. Digital recording either directly from camera processor or through external software and computer.

17. Four section trolley for organizing fiberscope and facility to hang the fiberscope without bending.

18. Leak tester to be provided.

19. Standard accessories including brush for cleaning the bronchoscope lumen.

20. Standard USB video recorder or recording through external software and computer

21. Should work on 220VAC +/-10%, 50 Hz.

22. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; GENERAL MEDICINE

Spirometer (Ordinary)

Spirometer :

Electronic with computer attachment and print out with software for complete analysis:

1. Should be electronic with computer attachment and print out with software for complete analysis

2. The System should be an economically oriented lung function measuring system for the determination of the static and dynamic lung volumes using the classical FRC- Helium rebreathing and the Diffusion Capacity by using the single Breath technique.

3. It should also be possible to measure Diffusion Capacity (DLCO) by Single-Breath intrabreath technique ( non breath hold method)

4. Indian predictive values should be available for all measurements and both methods of estimation of lung diffusion.

5. The proper medical reference from where these values have been incorporated should be clearly mentioned.

6. The system should measure the following :

a) Slow and forced spirometry (Inspiratory and Expiratiory Flow Volume Curve)

b) Lung subvolumes - Functional residual Capacity (FRC), Residual Volume (RV) . Total Lung Capacity (TLC) by FRC-Helium multiple breath technique

c) Diffusion capacity of the lung Single Breath Real Time & Single Breath Intrabreath (Non - Breath hold technique)

d) Diffusion capacity of the lung by the multiple breath technique or equivalent.

7. The system should measure the following parameters : b) Lung Subvolumes : FRC, RV, TLC, RV% TLC etc. c)Diffusion capacity of the Lungs : Diffusion Capacity of the Lung: DLCOSB, DLCOSBc, VA, KCO, KCOc)

8. The system should have an easy to exchange, bidirectional heated pneumotach with the following specifications.

Range - Should be 0 to 20 lit/sec. Accuracy - Should be +/-2% Resistance - Should be less than 0.05 KPa/ lit/sec.

9. The system should have Multi gas analyser, He analyser and O2 Analyser with the following specifications:

a) Carbon monoxide analyser : Range - Should be from 0 to 0.4% Resolution/Accuracy should be 0.0002%/0.0003% Reproducibility should be 0.0006%

b) He Analyser :Range - Should be 0 to 9.5% Resolution/Accuracy should 0.005% /0.05 % Reproducibility should be 0.02%

c) O2 analyser Range - Should be 0 to 100%. Resolution / Accuracy should be 0.05% / 1.0% Reproducibility should be 0.1%. . .

10. The system should have a demand valve unit for direct breathing (no inspiratory bag) from pre- mix gas container, to minimize wastage of gas.

11. Branded PC with Core i3 2.8GHz or more, Min. 4GB RAM, Min. 250GB HDD, DVD R/W, 20" Display with Licenced Operating System Windows 7 Professional or Higher, Keyboard, Mouse, HP color Printer)

12. System software should have facility for entry of patient data and saving of this information in a data base system. Software should be compatible and the latest . It should be possible to configure different report out put formats

13. It should be possible to upgrade the system to the following:

a) Airway resistance by shutter method.

b) Respiratory impedance by Impulse Oscillometry system.

c) Respiratory muscle strength, Respiratory drive.

d) Compliance - Static / Dynamic system

e) Body Plethysmography. f) Aerosol Provocation system.

g) Ergospirometry & Stress test ECG.

14. The system should have fully computerised calibration procedure for flow sensor and gas analysers. The system should also have a check procedure during start-up.

15. It should be possible to integrate/connect the system in a local Area Network (LAN). The data base must be accessed in a novel authorised operating system.

16. The software for diffusion must have program for patient training of DLCO Test without gas.

17. The software must be able to set values for discard volume, Alveolar time & other parameters according to user requirement.

18. Power Supply Power input to be 220-240VAC, 50Hz Suitable UPS with maintenance free batteries for minimum one-hour back-up should be supplied with the system.

19. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

20. Should be provided with Start-up Kit/Regent/ Calibrators/ Consumables/ Accessories are required performing the machine initial 500 tests or 3 Months test which is higher.

Department ; GENERAL MEDICINE

Equipment for Cardiac Pacing

Equipment for Cardiac Pacing :

1.   Bi phasic defibrillator compact and light weight with AED facility

2. Energy selection 5J to 200J in steps.

3. Momentary energy selection access on front panel.

4. Should have adult and paediatric paddles integrated on same handle.

5. Monitor should display selected and delivered energy.

6. Charging time maximum 7 secs for 200J.

7. Should have battery backup for 50 discharges of 200J.

8. Should have ECG inputs through paddles or 5 lead cables.

9. Should have display for selected ECG input source

10. Should supply three pairs of AED pads and the prices of AED Pads should be quoted separately in financial bid.

11. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; GENERAL MEDICINE

Hemodialysis Machine

Haemodialysis machine and suitable capacity RO Water Plant :

1. Machine should have facility for Acetate, Bicarbonate and Sequential dialysis (Isolated UF)

2. The blood pump should run even in the absence of water or dialysate flow

3. Machine should have have a facility for bacterial Filter

4. Should have Na, Bicarbonate and UF profiling,Temperature, dialysate flow and heparin profile

5. Dialysate temperatures selectable between 35 degrees C to 39 deg. C

6. Dialysate temperatures selectable between 35 degrees C to 39 deg. C

7. Variable conductivity setting between 12.5 to 16

8. Heparin pump with sringe size of 10,20 and 30 ml

9. Treatment parameter should be displayed by graph and digitally both

10. Should have heat and chemical disinfection facility

11. Should have drain facility

12. Should have accurate UF control by flow measurement technique.

13. Minimum patient data should be stored so that machine can be used anytime without feeding data every time

14. Should have automatic self test facility OK

15. Should have auto ON/OFF Facility OK

16. Should have touch screen

17. Easy to service, troubleshoot and calibrate

18. Machine should have inbuilt diagnosis/trouble shooting screen

19. Blood pump rate from 50-600 ml with an incremental of 1 ml

20. Ability to monitor pulse rate and NIBP with graphic and tabulated trends.

21. Audio visual alarms on limit violation of conductivity, blood leak, air leak, transmembrane pressure alarms,, Dialysis temperature alarm, dialysis can empty alarm, end of disinfection alarm, bypass alarm and blood pump stop alarm

22. Suitable capacity of RO water plant for operating and functioning of minimum 20 haemodialysis machine in to two shift.

Accessories Spares & consumables

a. System as specified

b. Compatible Online sinewave UPS with one hour back up

c. Dialyzer -100 nos

d. Water filter-2 nos

23. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; GENERAL MEDICINE

Arterial Blood Gas Analyser (ABG)

Arterial Blood Gas Analyser (ABG) :

1. Compact system for measuring pH, pCO2, pO2, -HCO3 & four Electrolytes like Na, K, Ca+ and Cl - in blood.

2. All should be measured in a single injection / aspiration of Sample.

3. May have provision of modular platform for future up gradation to include glucose, lactate & haemoglobin the same machine with the inspiration of single sample.

4. Should be able to analyse all parameters using low blood volume directly from syringe or capillaries.

5. Fast and accurate result of test made available in about 60 seconds.

6. Automatic Calibration by liquid calibrators with flexible time mode. Instrument should have Stand-by mode facility and Economy mode.

7. It should not be cartridge-based system.

8. Should be provided with Start-up Kit/Regent/ Calibrators/ Consumables/ Accessories are required performing the machine initial 500 tests or 3 Months test which is higher.

9. All the consumables and spares should be quoted separately unit wise.

10. Compatible online UPS with battery backup of at least one hour

11. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; GENERAL MEDICINE

EMG and Nerve Conduction Velocity Machine

EMG and Nerve Conduction Velocity Machine :

a) Fully computerized latest EMG/NCS/EP system with 8-channel facility.

b) Computer

1. The system should be based latest i3 PC , a minimum hard disk capacity of atleast 120GB, 512 MB RAM or better (best available at time of purchase). It should have a built in DVD / CD- Writer as standard multimedia keyboard, 10/100 network interface

2. Operating System: The operating system software should be based on industry standard, robust Microsoft Windows 10 Professional with latest service pack for easy networking and data storage and integration.

3. Display Monitor : The display should be atleast 24 inch TFT color, flat panel, having full HD resolution of 1980 X 1080 or better. The monitor should be flexible to move closer (forward / backward) to the examiner for easy viewing during acquisition. Should be able to display variable screen formats depending on recording mode. Simultaneous display of wave forms, graphs and alphanumeric data should be possible.

4. Wave form acquisition, display and Storage: The time base range should be minimum 0.1ms/divn to 2 sec/div in multiple steps depending on test and time base type of single, dual and individual, independently selectable in specific tests. The storage of waveforms should be on hard disk.

c) Amplifiers : The number of channels must be 8 :- 1) The display modes of normal, and plus/minus with artifact rejection facility and average display sensitivities 0.5μV/divn to 20mV/divn in multiple

steps depending on test. Input impedance of >10,000Mohm with CMMR ratio of >110dB at 50Hz/60Hz should be available

2) Sensitivity from 2μV/Div - 20MV/Div, , High Filter: 20Hz -20KHz, Low Filter: 0.1 - 500Hz

3) noise of less than 0.7μVRMS from 2Hz to 10KHz built-in calibration, Notch filter. Built in Temperature monitoring facility.

d. Electrical Stimulator: Two Constant current stimulator with current ( from 0 to 100mA) with increments of 0.1mA and pulse duration varies from 50μs - 1000μs with 50μs increments. Remote controlled probes ( 2 nos ) with adjustable intensity,. The stimulator probe can controllable both from the probe and the control panel.

e. Auditory Stimulator: Auditory headphones with signal types click, tone and stimulus intensity of 32 -132 dB SPL depending on stimulus types and stimulus polarity of condensation, rarefraction and alternating with tone frequencies 125Hz to 8 KHz with readymade protocols should be available. White masking noise should be available.

f. Visual Stimulator: Visual stimulator with color LED monitor of 22” with flexible target size and position. It has pattern intensity control and a variety of patterns utilizing checks, bars or gratings with different color combination. It is possible to change the viewing angle and viewing distance automatically depending on each other value. The various Pattern color (foreground and background) has user selectable from combination of Red / Green and Blue and many other colors.

g. Application Software: Acquisition and review software should be based on Windows 10 latest operating system with easy to use menus, flexible protocols with standard settings and user definable settings. It has the packages for doing EMG acquisition and analysis, quantitative analysis of EMG (facility to record/Replay EMG on Hard Drive for unlimited duration depending on HDD size) and interference pattern analysis with QMUP, IPA, F-Wave, H-Wave and blink reflex studies, Repetitive stimulation studies should be available. The standard NCS studies with report generation.

ReferenceValues package must be included as standardfor online comparative data of NCS.

The Evoked potential studies for AEP, VEP and SEP should be available as standard protocols.

Compatibility with external stimulus devices for recording evoked potentials should be available with trigger in/out facility. All the reports should be integrated with MS word. SFEMG, SSR, R-R interval variation, Conduction Velocity Distribution and Auditory P-300 should be provided along with the powerful software for database management capable of HIS connectivity at any future time.

The software should have packages for recording Contingent negative variation (CNV) and other event related potentials like the Bereitschaft potential (with necessary accessories).

h. Accessories: Standard accessories should be supplied, viz.

1) EMG cables (minimum of two)

2) Surface electrodes (2 sets of 24 electrodes each)

3) Needle electrodes (6 nos.)

4) Conduction paste and cleaner fluid

5) LED goggles for VEP recordings in uncooperative patients and children Stimulus device for CNV

6) Printer :High end Laser B & W printer for all the printouts of patient reports.

i. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; GENERAL MEDICINE

Invasive Mechanical Ventilator

|  | 1. | Should have facility for Invasive and Non-Invasive ventilation. |  |  |
| --- | --- | --- | --- | --- |
| 2.Microprocessor Control suitable for adult ventilation. |  |  |  |  |
|  | 2. | Electromagnetic Compatible Hinged arm holder for holding the circuit. |  |  |
| 3. | Should be able to record and analyse various parameters. Breath to breath / pulmonary functions, loops to be stored in the memory with feasibility of trend analysis on a TFT touch screen. Screen size should be more than 10 to 15” so that be easily seen from distance. Compiled trend analysis at Minimum 48 hours for all measured parameters. Monitoring during mechanical ventilation includes measurement of peak and plateau pressures, intrinsic positive end-expiratory pressure, and work of breathing. |  |  |  |
| 4. | Machine should be Compressed air / oxygen driven / turbine driven. |  |  |  |
| 5. | Should have Nebulization assembly facility. |  |  |  |
| 6. | Ventilator, Compressor & Humidifier should be Same Trolley / cart mounting for easy transpo |  |  | rtation. |
| 7. | Should have Internal rechargeable battery at least 05 Hrs. backup. |  |  |  |
| 8. | Au | tomatic Self - test compliance and leakage compensation for circuit and ET Tube. |  |  |
| 9. | Should have the following modes: |  |  |  |
|  | b. | Volume Controlled, CMV, PCV or IRV, SIMV (volume cycled & pressure limited), |  | SIMV+PS |
|  |  | (volume cycled & pressure limited), CPAP, Bilevel, APRV. IMPRV, PRVC, AUTOMODE, |  | Minimum |
|  |  | minute ventilation, Mandatory rate ventilation, Proportional assist, Mandatory ventilation. |  | minute |
| 10. Should have the following setting. |  |  |  |  |
|  | f. | Tidal Volume: minimum 200ml maximum of 2000 ml in Volume control is essential. |  |  |
|  | g. | PEEP 0 to 30 cmH20 or more |  |  |
|  | h. | Oxygen Concentration 21 –100 % |  |  |
|  | i. | Inspiratory Pressure 1-80 cmH20). |  |  |
|  | j. | Respiratory rate 1 to 80 bpm. |  |  |
| 11. Alarm |  |  |  |  |
|  | c. | Adjustable Alarm. - Low/high minute volume, low/high pressure, low/high tidal volume, low/high rate, apnea time, low/high oxygen, low/high SpO2 |  |  |
| d. |  |  | Special alarm - O2 cell Failure , flow sensor, battery, power supply, gas supply, oxygen concentration, |  |
| 12. Compressor |  |  |  |  |
| h. |  |  | Should be supplied with External Compressor. |  |
| i. |  |  | The compressor has been designed to supply the ventilator with dry, filtered compressed air. |  |
| j. |  |  | Compressor should be oil-free. |  |
| k. |  |  | Portable & fitted with ventilator cart. |  |
| l. |  |  | Air filtration 5 microns. |  |
| m. Noise level dB 40 – 50. |  |  |  |  |
| n. |  |  | Peak flow of 200 lpm. |  |
| 13. Humidifier |  |  |  |  |
| f. |  |  | Servo controlled heated Respiratory Humidifier. |  |
| g. |  |  | Display Should be of LED /LCD. |  |
| h. |  |  | Temperature control settings & Temperature range: 28-40 deg. |  |
| i. |  |  | Temperature should be adjustable. |  |
| j. |  |  | Jar should be autoclavable |  |
| 14. Standard Accessories / spare & Consumable. |  |  |  |  |
| l. |  |  | Silicon breathing circuit (Adult reusable) - 5 complete set. |  |
| m. Nebulization assembly compatible circuit 5 complete set. |  |  |  |  |
| n. |  |  | Humidifier - 1 No. |  |
| o. |  |  | O2 Pressure Regulator with hose - 1 No. |  |
| p. |  |  | Hose for O2 connection with connector - 5 mts. |  |
| q. |  |  | Hose for compressed air with connector - 5 mts. |  |
| r. |  |  | Test lung - 1 No. |  |
| s. |  |  | Non-invasive ventilator mask reusable for adult (3sizes) - each size 5 No. |  |
| t. |  |  | ET tube cuff pressure monitor and HME filter – 10 no |  |
| u. |  |  | Inbuilt / integrated nebulizer-1 No. |  |
| v. |  |  | All sensors and other non-consumable items (other than reusable silicon ventilator circuits) should be free of cost during warranty and CMC. |  |
| 15. Ventilator, Humidifier & Compressor Power Supply input to be 200-240VAC, 50 Hz fitted with Indian conditions plug. |  |  |  |  |
| 16. Suitable online UPS with commensurate capacity for all ventilators including compressor & Humidifier with maintenance free batteries for minimum one-hour back-up should be supplied. |  |  |  |  |
| 17. Ventilator, Humidifier & Compressor Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO. |  |  |  |  |
| NOTE: |  |  |  |  |
| 1) Reusable consumables (other than reusable silicon ventilator circuits) should last during the warranty period. |  |  |  |  |
| 2) Ventilator & Humidifier any additional reusable consumables are required during the warranty period those will be supplied free of charge by the supplier. |  |  |  |  |
| 3) The life expectancy of the reusable consumable is expected to be of at least one year from the date of installation of the same. The reusable consumables will be procured at the prices accepted as per the contract. |  |  |  |  |

4) The bidders should submit all reusable consumable items price & their authorised local office/

distributor name in the financial bid.

Department ; GENERAL MEDICINE

Upper GI Endoscope (Optional)

Upper GI endoscope :

Video Gastroscope HD with NBI

1. Field of view-140 degree

2. Direction of View-Forwarding View

3. Depth of field -2-100mm

4. Distal End Outer Diameter-9.2mm

5. Insertion Tube Outer Diameter

6. Working length 1030mm or more

7. Instrument channel inner diameter 2.8 mm

8. Instrument minimum visible distance 3.0 MM

9. Cauterization Treatment Should be compatible

10. Angulation range Up: 210°, Down : 90°, Right: 100°,Left : 100°

11. Total length 1350mm or more

12. Special Light Should have NBI

13. HDTV Should be available

14. Video Bronchoscope with NBI

15. Field of view 120°

16. Direction of View Forward Viewing

17. Depth of field 3-100mm

18. Distal End Outer Diameter 5.9 mm

19. Insertion Tube Outer Diameter 6.0mm

20. Working length 600 MM

21. Instrument channel inner diameter 2.88mm

22. Instrument minimum visible distance 3.0 MM

23. Angulation range Up: 180°, Down : 130°

24. Special Light Should have NB

25. SEMI RIGID PLEURAVIDEOSCOPE

26. Distal End Outer Diameter 7.0 mm

27. Insertion Tube Outer Diameter 7.0mm

28. Field of view 120°

29. Depth of field 3-100mm

30. Angulation range

31. Working length 270 MM

32. Total length 520MM

33. Instrument channel inner diameter 2.8 mm

34. High Frequency Compatibility

35. Laser compatibility YAG, 810NM Diode

36. HD VIDEO PROCESSOR WITH NBI

37. Voltage 100-240 V AC

38. TYPE OF PROTECTION AGAINST ELECTRIC SHOCK CLASS 1

39. EXAMINATION LAMP LED/300 WATT XENON

40. OUTPUT HD-SDI.SD-SDI,DVI

41. WHITE BALANCE ADJUSTMENT FROM THE FRONT PANEL

42. COLOUR TONE ADJUSTMENT RED, BLUE, CHROM +/- 8 STEPS

43. AUTOMATIC GAIN CONTROL (AGC) SHOULD HAVE AGC

44. CONTRAST ADJUSTEMENT SHOULD HAVE ADJUSTMENT

45. NOISE REDUCTION SHOULD HAVE NOISE REDUCTION

46. IRIS SHOULD HAVE PEAK AND AVERAGE MODE

47. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; GENERAL MEDICINE

Sigmoidoscope (Optional)

Sigmoidoscope :

1.   Video sigmoidoscope – insertion tube dia 13.2 mm or less, instrument channel 4.2 mm or less, field of view – 140 Deg., Focal range - 3 to 100 mm, Angulation UP/Down 180/180 Deg. R/L 160/160 Deg. with 700 mm or less working length, total length - 1100 mm or less

2.   Only Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; GENERAL MEDICINE

Colonoscope (Optional)

Colonoscope :

1. Video colonoscope – insertion tube dia 12.9 mm or less, instrument channel 3.7 mm or more ,angle of view 140 -deg., focal range -4 to 100 mm, Angulation Up/Down 180/180 deg. R/L 160/160 deg. With 1600 mm or more workding length, total length 2000 mm or more

2. deg., focal range -4 to 100 mm, Angulation Up/Down 180/180 deg. R/L 160/160 deg. with 1680 mm workding length, total length 2000 mm

3. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; GENERAL MEDICINE

Haemocytometer

Haemocytometer :

1. Hemocytometer counting chamber should be able to determine the number of particles per volume of blood. Particles are counted under the microscope

2. It should contain improved neubauer chamber mercury coated cover slips, RBC & WBC pipette.

3. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; GENERAL MEDICINE

Light Microscope

Light Microscope :

1. Binocular microscope with universal infinity corrected optical system

2. Halogen light source illumination

3. Rigid frame with ergonomics design

4. Binocular observation tube with inclination of 45/30 degrees

5. Built in torque adjustable focusing knob

6. Square mechanical stage with rigid hand coaxial control

7. Abbe condenser, Iris diaphragm

8. Revolving Quadruple nose piece (for objectives)

9. Plan achromat objectives 4X, 10X, 40X, 100X (Oil)

10. 40X, 100X objective should be spring loaded

11. Eye piece 10X (FOV 20)

12. Antifungal treatment should be applied to the observation tube, eyepiece and objective

13. Accessories, dust cover and power cord

14. Power requirement 220 V/50 Hz

15. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; GENERAL MEDICINE

Tread Mill Test Machine

Tread Mill Test Machine (TMT) :

1. System should be Dedicated microprocessor console based cardiac workstation simultaneously 12 lead acquisition combines resting & exercise ECG in one unit.

2. Should have digital Acquisition module to acquire diagnostic quality ECG data.

3. Each wire of patient cable set should be detachable, so that each cable can be changeable in case of one cable faulty.

4. The ECG acquisition sampling rate should be 1,000 Samples/seconds channel or more.

5. System should have 22” display for easy access. Monitor should adjusts to any angle for better visualization.

6. Should have facility to Review, edit an add ECG from full discloser storage post-exam.

7. Should have facility to hide Zoom ECG, context ECG view and Trends at any time.

8. System should have 50mm sweep speed selection for ECG display.

9. Should have full disclosure of all 12 leads for beat to beat analysis.

10. The final report should include information on blood pressure, heart rat MET,s treadmill speed/Grade, ST trends relating to stage wise & recovery phase and duke treadmill score etc.

11. Report should be user-definable and can be selectable at final step of reporting.

12. Automatic calculation & display of METs.

13. System should support Time and METs ramped protocol.

14. System should show recovery elapsed time in %.

15. System should support left to right work flow.

16. System should provide online printing of ECG prints on High Quality Thermal printer manually and automatically during stress testing.

17. Treadmill soft stop option for stopping the treadmill after 20 second in recovery mode.

18. Facility to get system generated auto statement report.

19. System should support editing of final report in review phase.

20. System should support user defined ST measurement points.

21. System should have special filters to reduce noise artefacts, motion artefacts, baseline artefacts during stress test.

22. System should be capable to store full disclosure ECG data for later review using page review mode.

23. System should support multi login password protected access.

24. System should be supplied with US-FDA approved stress automatic BP measurement device with interface cable to measure automatically the patient NIBP during stress test per the programming done at stress system.

25. The stress testing system should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; GENERAL MEDICINE

Portable Suction Machine

Portable Suction Machine :

1. Vaccum /LPM : - 700 MM Hg , 50 Litres/Min

2. Pump Type- Double rotary vane type

3. flutter free vacuum control knob,

4. Wide mouthed 2 x 2 LITRE (light weight, unbreakable and clear) with self-sealing bungs and mechanical over flow safety device.

5. Noise (in dBA)- 50 dB

6. Collection container & its cap, suctions tube tips, a vacuum gauge, two sets of moisture & microbial filters and control knob

7. SiliconeTubing:8 mm ID x 2 mtr (PVC), 2x2 lt jar (one set

8. It should be Mobility, portability.

9. Equipment shall operate on 220V-240V, 50 Hz, single phase electric supply.

10. ISO/ BIS approve model should be offered.

Department ; GENERAL SURGERY

Proctoscope

Proctoscope :

1. Working Attachment, for use with Proctoscope and Telescope Proctoscope with obturator O.D.: 24 mm, working length: 8 cm.

2. Straight Forward Telescope 30°, Eyepiece 45° angled, diameter 4 mm, length 9.5 cm, autoclavable, Fiber optic light transmission incorporated

3. Sponge Holder, working length 20 cm Dressing Forceps, total length 19 cm

4. Fistula Hook, total length 19 cm

5. Proctoscopy Punch, Through-cutting, cutting width 3.4 mm, straight jaws, sheath diameter 3.5 mm, working length 20 cm

6. Injection Needle, straight, LUER-Lock, tip diameter 1.0 mm, working length 14 cm Hemorrhoid Grasping Forceps, for use with ligature instrument

7. Ligature Instrument, for treatment of hemorrhoids, working length 17 cm including: Loading Cone,

8. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; GENERAL SURGERY

Operation Theatre Table

Operation Theatre Table :

1. Should be multi purpose powered electro hydraulic OT table, C- Arm Fluoroscopic compatible, suitable for all major urological endoscopic surgical, open surgical procedures and vascular transplantation surgery, complete with a corded handset with battery level indicators and moulded, anti-static, seamless mattress.

2. Table top should have feature of movement with a traverse of minimum of 250 mm or more, either cranially or caudally

3. Should have full length X-ray translucent top with removable & interchangeable head and leg sections with an auto-locking mechanism.

4. Table must allow for unrivalled C-arm access and kidney break positioning without the need to move the patient.

5. Trendelenburg/ antitrend-: 20-25 degree, Lateral tilt- 15-20 degree, FLEX/Reflex- 220 degree/ 100 degree, Slide- 300 mm, Height- 600-1030 mm with 80 mm mattress

6. The brakes, wheels and castors should be controlled by two foot pedals (single foot pedal) /Remote control

7. The table should feature an integrated stand by panel for controlling the movements in case of handset loss or battery failure

8. The table stem should be located under the middle of the back section making the tabletop eccentric.

Operation Theatre Table :

9. Table should be able to carry heavy patients and have a capacity of up to 225 kgs.

10. Table should also be suitable for tall patients and have a length of at least 2150 mm

11. Table should offer low minimum height enabling the surgeon to operate even when seated

12. The table should have divided leg section with mattresses, arm board & universal clamp

13. Should have facilities for manual operations in case of power failures.

14. System Configuration Accessories, spares and consumables

15. The table should be supplied with following necessary accessories:

16. Padded leg crutches with lithotomy holders: 1 pair

17. Arm supports – 2 nos

18. Gel heel pads – 1 pair

19. Body strap along with separate gel strap

20. Hand Surgery Board – 1

21. Elevated Arm Support – 1

22. Padded head, shoulder and arm rest – 1 set each

23. Padded lateral support and shoulder supports – 1 set

24. A square or L shaped anaesthesia screen frame with appropriate clamps / locks – 1 No.

25. Appropriate accessories’ clamp.

Operation Theatre Table :

26. Stainless steel drainage tray with sieve and drain tube for cystoscopy 1No.

27. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; GENERAL SURGERY

Pedestal lights

Pedestal Lights :

1. Light, for medical examination, on mobile stand.

2. Arm: 105 cm articulated, spring loaded arm, arm with on/off switch and incorporated electronic transformer.

3. Mobile stand with 5 swivel castors.

4. Power supply: 220 V.

5. Bulb: approx.12V/20W, halogen, light intensity: approx 20.000 Lux at 40 cm.

6. Lamp to emit natural white light: colour temperature 4000 K.

7. Reflector adjustable for positioning.

8. Free cord: length approx. 3 m.

9. To be supplied with: 1 spare bulb and 1 spare fuse.

10. Should have antiskid castor on/wheels for easy mobility.

11. Light head – should be shadow free light. Should have sterilizable focusing handle.

12. Depth of illumination – 70 cm minimum

13. Should be disconnectable from the stand

14. Bulb life should be minimum 1000 hrs.

15. Replacement of bulb should be easy

16. Heat dissipation should be efficient to keep the air at the surgeons head level cool.

17. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; GENERAL SURGERY

Item Sche Anesthesia Equipment

| Anesthesia Equipment : |  |
| --- | --- |
| 1. Powder coated structure. |  |
| 2. Active Anti Hypoxic Electronic device - to ensure no nitrous flow with loss of O2 pressure. |  |
| 3. Ratio Control – 28% mixed ratio between O2 & N2O Gas Flow |  |
| 4. Should have Electronic Flow meter for O2, N2O & Air |  |
| 5. Should have Electronic display of Gas pressure in the user interface display |  |
| 6. Cylinder Pressure and operating pressure should be indicated electronically in user interface display. |  |
| 7. Oxygen Driven Nitrous Supply through Electronically |  |
| 8. Inner connection of tubing’s should be made of alloy metal or PPU tubes. |  |
| 9. Audio Visual alarms should be available for Oxygen Failure warning Device |  |
| 10. Electronically measured cylinder pressure, quick-release system for docking and release. |  |
| 11. Electronically /Pneumatically controlled technology valve for High precision pressure relief cum non return valve should be available |  |
| Anesthesia Equipment : |  |
| 12. The vaporizer should be electronic / Pneumatic with an injector. The container of the vaporizer chamber should be minimum 300 ml (10,1 oz.). Refilling of inhalation agent should be possible during running mode. The liquid level in the vaporizer should be displayed both on vaporizer and control display and an alarm should be given when the level becomes too low. The desflurane vaporizer should not require to be warmed up before. |  |
| 13. Provison to connect two electronic controlled vapouriser simultaneously with electronic switching. |  |
| 14. Switching between manual and automated ventilation should be possible. The system shall contain a continuously adjustable pressure relief safety valve settable between 0-70 cm H20. Total system volume should be as low as possible and not exceed 3 liter including absorber. |  |
| 15. Gas Inlet port for O2 & N2O, pin- index type 1 each & O2,N2O & Air non-interchangable type hose 1 each |  |
| 16. Direct Central Pipe Line connectivity to Machine. |  |
| 17. Standard Maggils circuits, all rubber antistatic tested. |  |
| 18. Alloy metal tray for working / writing table. |  |
| 19. Castor wheel with Individual breaks. 4 wheels with individual locking brake. |  |
| 20. Double Oxygen outlet available for circuirs |  |
| Anesthesia Equipment : |  |
| 21. Operating gas pressure available on the user interface |  |
| 22. 2 - B” type cylinder carrying facility at back side. |  |
| 23. Circle Absorber shall contain a continuously adjustable pressure relief safety valve settable between 0-70 cm H20. Total system volume should be as low as possible and not exceed 3 liter including absorber. |  |
| 24. Ventilator Specifications: |  |
| a) The equipment shall contain functions for volume- and pressure regulated ventilation. |  |
| b) The ventilator shall contain the functions “Spontaneous Breathing” and” Manual Ventilation”. |  |
| c) The ventilator shall have the following modes: VCV, PCV, PRVC, SIMV |  |
| d) The ventilator shall have Pressure Support mode with backup functionality. |  |
| e) The ventilator should not contain any moving parts |  |
| f) The equipment shall handle high- to low flow anesthesia during both non-rebreathing and partial rebreathing conditions. Switching between different breathing systems shall be a simple operation. |  |
| g) | It shall be possible to regulate the I:E ratio between 4:1 to 1:8 |
| h) | The equipment shall include an integrated continuous PEEP-function ranging from 0 to at least 20 cm H2O or more |
| Anesthesia Equipment : |  |
| i) | The set PEEP level should be visible on the control display. |
| j) | The set pressure should be adjustable between 5 - 70 cm H20. |
| k) | The set tidal volume should be adjustable between 20-2000 ml (0,67- 67 oz.) |
| l) | The breathing frequency should be adjustable between 3-80 breaths/min or more |
| m) The maximum flowrate should be approximately 200 l/min |  |
| n) | The equipment should provide a fast rise time in Pressure Controlled mode without overshoot. |
| o) | The equipment should have an easy accessible timer displaying hours, minutes and seconds. |
| p) | The equipment should have tools for performing a lung recruitment maneuver |
| q) | The lung recruitment maneuver should be automatic |
| r) | It should be possible to measure and visualize dynamic compliance breath by breath |
| s) | Predictive body weight settings should be available for the ventilator settings |
| t) | It shall be possible to pause ventilation and fresh gas flow for a defined time period." |
| 25. Supplied with all standard accessories including Bains Circuit. |  |
| Anesthesia Equipment : |  |
| 26. Closed circuit system , Sodalime canister, Single/Double Chamber |  |
| 27. Pediatric Circuit (Jackson Rees) |  |
| 28. Auxiliary O2 Supply point |  |
| 29. Mask of Different size (2 each- 1 to 5) |  |
| 30. Reservoir bag- 3 litre/5 litre/0.5 litre- 2 each |  |
| 32. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO. |  |

Department ; GENERAL SURGERY

Autoclave

| Autoclave : |  |
| --- | --- |
| 1. Horizontal rectangular or vertical cylindrical high-pressure steam sterilizer |  |
| 2. Chamber cpacity: 75 litre |  |
| 3. Electrical Power : 18/36 KW or Sufficient wattage of Industrial immersion type water heater to generate steam with in a reasonable period of time on 3 phase 440 v 50 HZ ac supply |  |
| 4. Working pressure and Temperature : 1.06 Kgf / sq.cm -1.2 Kgf /sq.cm at 121 Deg C |  |
| 5. Material of construction |  |
| a. Inner chamber, Jacket, Door : SS 316.(5mm-10mm) |  |
| b. Outer Chamber : SS 304 (insulated properly) |  |
| c. | Steam generator : Non corrosive SS /Chromium plated Brass |
| d. Heater Plate : Brass/Stainless steel |  |
| e. Pipe Line : Complete with SS |  |
| f. | Stand : Stainless Steel/High quality non corrosive Steel Instrumentation |
| 6. Temperature, Pressure and vacuum gauges, steam traps , vacuum driers, water level indicator on steam generator Safety devices. |  |
| 7. Pressure switch and safety valve, self-locking of door when chamber is under pressure: vacuum breaker for jacket. |  |
| 8. Steam generator with gauge glass valves and Low water protection with audiovisual indicator. |  |
| Autoclave : |  |
| 9. Entire cycles of sterilization should be controlled by a microprocessor unit and it should also have manual operation facility. |  |
| 10. Accessories |  |
| a. Heating element 2 set extra with each machine. |  |
| b. Necessary Pipeline works for water inlet and steam outlet should be done. |  |
| c. | Cable for connecting to wall socket should also be provided |
| 11. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO. |  |

Department ; GENERAL SURGERY

Non-invasive Multipara Monitors

Multipara Monitor :

1.   Should be able to monitor 5 lead ECG, SPO2, NIBP, Respiration rate and Temperature.

2.   Should be portable with carrying handle.

3. Should have touch screen TFT display with at least 10 inches or higher with at least 6 waveforms and numeric display simultaneously.

4. Should have Lithium ion battery with 4 hours battery backup.

5. Should have keys for quick access to main functions.

6. Should have adult, pediatric and neonatal modes.

7. Should provide prominent prioritised audio, visual alarms for high, low heart rate, Spo2, RR, low battery, lethal arrhythmia recognition and ST Analysis.

8. Should have separate volume control for beep sound for QRS and alarm sound.

9. Should provide following accessories:

Reusable adult 3 lead ECG cable set – 2 nos.

Reusable adult and pediatric SPO2 finger probes – 1 each.

Adult and pediatric NIBP cuff of different size

10. Should work on 200-240V AC/50Hz with inbuilt rechargeable battery.

11. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

12. Electrical safety conforms to standards for electrical safety IEC-60601/IS-13450.

Department ; GENERAL SURGERY

Cystoscope & Resectoscope

| Cystoscope & Resectoscope : |  |
| --- | --- |
| Cystoscope with Light source, camera & Monitor |  |
| 1. Sheath with obturator outer Diameter- 20 Fr |  |
| 2. Working channel- 04-05Fr |  |
| 3. Working length- 22cms |  |
| 4. Self-dilating tip for ease of entry |  |
| 5. Compatible 30degree an 0-degree HD telescope – 01 each |  |
| 6. Telescope bridge – 01 |  |
| 7. Stone crushing forceps compatible with above sheath and telescope – 01 in no. |  |
| 8. Resectscope sheath compatible with above mentioned telescope – 01 in no., with |  |
| Moveable thumb ring (working element) – 01 in no. |  |
| Spring loaded action with electrodes inside the sheath in rest position |  |
| a. Telescope Bridge – 01 in no. |  |
| b. Cutting loop – 10 in nos. |  |
| c. Coagulating hook electrode – 10 in nos. |  |
| d. Monopolar high frequency cord – 05 in no. |  |
| e. Protection tube for electrodes – 05 in no |  |
| 9. Fiber optic cold light source: 1no. 300 watt Xenon bulb /LED (Bulb life 500 hrs minimum for xenon & LED Life 10000 hrs. |  |
| 10. Should have 3 cheap camera |  |
| 11. Should have 21” Medical grade HD Monitor. |  |
| 12. Power supply 220VAC +/- 10 %, 50Hz. |  |
| 13. On-line UPS with 60 minutes back up Entire system. |  |
| 15. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO. |  |

Department ; GENERAL SURGERY

Flexible Video Colonoscope

Flexible Video Colonoscope :

Video colonoscope – insertion tube dia 3.7 or more mm, instrument channel 3.8 mm, angle of

view 140 deg., focal range - 3 to 100 mm, total length 2000 mm.

Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; GENERAL SURGERY

Flexible Video Gastrodudenoscopy for ERCP

Flexible Video Gastrodudenoscopy for ERCP :

a. Field of view 140 degree

b. Direction of view is side Viewing

c. Depth of field 4 to 100mm

d. Total length - 1340 to 1365mm

e. Working length more than 1000mm

f. Insertion tube outer diameter 10 mm to 12mm.

g. Distal end diameter 10 to 12mm

h. Bending section tip deflection

i. Up – 210 degrees

j. Down – 90 to 120 degrees.

k. Left - 90 to 120 degrees.

l. Right – 100 to 120degrees.

m. Instrument channel - Diameter – more than 3.5mm

Video Colonoscope :

a. Field of view 140 degree

b. 2.Direction of view Forward viewing

c. 3.Depth of field 3 to 100mm

d. 4.Total length - 1600mm to 2100mm

e. 5.Working length more than 1300 mm to 1750mm

f. 6.Insertion tube outer diameter 13mm or more

g. 7.Distal end diameter 13 to 14mm

h. 8.Bending section tip deflection

i. Up – 180 degrees.

j. Down –180 degrees

k. Left - 160 degrees.

l. Right – 160 degrees.

m. Instrument channels - Diameter – First channel more than 3.6mm/ second channel more than 2.6mm.

VIDEO PROCESSOR WITH LIGHT SOURCE COMPATIBLE WITH THE ABOVE MENTIONED VIDEOSCOPE.

a. Colour system Video HD Processor

b. Xenon lamp 100W

c. Out put for high video image transfer using External Software and computer

d. Convenient digital to digital recording of both still and moving images. using external software

e. Picture and picture display for any combination of endoscopic images.

f. Convenient index display for documentation.

g. Scope ID function for endoscopy suite management.

h. Power supply : 220 – 240V AC

i. Should include Air feeding & water feeding system with detachable water container.

High Definition Lcd Colour Monitor ( Flat Panel):

a. Full digital High Definition TV compatible high resolution LCD monitor

b. LCD Screen 19 or more inches in size

c. Power supply 220 – 240 V AC

d. Computer connector

Hardware for recording, archiving and printing

State of the art hardware for recording, storing and printing reports compatible with the system Furniture

Mobile cart with suitable compartments to house all the above equipment with built in plugs switches System Configuration Accessories, spares and consumables System as specified

a. Biopsy forceps :3 each

b. Foreign body grasper (basket type) 2

c. Polypectomy snare:2

d. Standard tip canula:2 types – 10 each

e. Polypectomy cautery system :1 No

f. Endoscopic washer -1 no

Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; GENERAL SURGERY

Flexible Video End viewing Oesophago-Gastroduodenoscope

Flexible Video End viewing Oesophago-Gastroduodenoscope :

SPECS OF SCOPE: (GI SCOPE)

1. Direction of view should be zero degree

2. Minimum of 140 degree (app) of field of view

3. Range of observation from 4 mm to 90 mm.(app)

4. Angulations of tip not less than 210-120 deg(Up) and 120 deg( down)with right to left movement of minimum 100 deg.(app)

5. Insertion tube outer diameter of less than 8 mm with a working length of not less than 1000 mm

6. Distal end of less than 7.8 mm

7. Instrument channel of more than 2.3 mm

8. Compatible with the video system specified

9. Video processor with light source & Monitor

10. Power supply 200-240 V A/C

11. PAL type video signal. The camera should be 3 chips CCD with high definition (HD) Output with provision of recording

12. Controls for color adjustment, to enhancement and balance settings

13. Controls to freeze images enhance a portion of frozen

14. Patient and physician data input keyboard

15. Operates on Xenon lamp aapprox.150 W

16. Should have an Emergency changeover lamp

17. 19” LCD color monitor with XGA resolution

System Configuration Accessories, spares and consumables

Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; GENERAL SURGERY

Flexible Video Sigmoidoscope

Flexible Video Sigmoidoscope :

1.   Video Sigmoidoscope – insertion tube dia 13.2 mm, instrument channel 3.8 mm,angle of view 140 deg., focal range -4 to 100 mm, Angulation Up/Down 180/180 deg. R/L 160/160 deg. with 1700 mm workding length, total length 2023 mm

2.   Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; GENERAL SURGERY

Flexible Video Bronchoscope

Flexible Video Bronchoscope :

1. OPTICAL SYSTEM :

a. Field of View-Minmum 120 degre

b. Direction of View- 0 degree Forward

c. Depth of field: 3 mm to 100 mm maximum

d. Insertion Tube:

e. Outer diameter (distal end):6.0mm(+-0.1mm)

f. Outer diameter (Insertion tube):6.0mm(+-0.5mm)

g. Working Length:600 mm

h. Total Length:870 mm

i. Instrument Channel

j. inner diameter: 2.8mm or more

k. Minimum Visible Distance:3mm from distal end

l. Angulation Range

m. Up: 180 degree

n. Down: 130 degree

2. Scope should be laser and electro cautery compatible LASER COMPATIBILITY: ND: YAG, 810 nm diode

3. VIDEO PROCESSOR AND LIGHT SOURCE:

a. COLOUR SYSTEM: High Definition

b. CCD should produce quality images to provide outstanding clarity, sharpness, mucosal structure and patterns of Bronchial tree

c. SCOPE should have remote switches to adjust Processor function.

Flexible Video Bronchoscope:

d. Lamp: 150 w Xenon bulb with emergency lamp provision for Illumination Video Signals:

4. Video Signals

a. Analog Output : RGBS-1 Set

b. Y/C-2 Set

c. Composite-1 set

d. Scope image spectrum with minimum three steps to get the vascular and mucosal pattern. Should display natural mucosal Color tone.

e. Should have provision for ZOOM ADJUSTMENT.

f. Suitable software for video recording and JPEG (Picture capture) Patient documentation should be possible.

g. Should have provision for air pump.

h. KEYBOARD FACILITY should be available for the data inputs.

5. Monitor

1. 14 inches LCD colour monitor along with system (Medical Grade High Definition Monitor) Should have provision for accepting Y/C, BNC and RGB signal.

2. COMPUTER WITH NECESSARY SOFTWARES WITH JPEG / VIDEO RECORDING FACILITY WITH LASER (COLOR) PRINTER

3. Core 2 Duo Processor with Intel chip mother board 915. 4GB RAM with 500GB HDD

4. DVD writer and COM PORT 4-5 USB ports,One free PCI slot

Flexible Video Bronchoscope:

5. 15 inches LCD monitor (FHD – Full High Definition)

6. Windows Vista/XP professional Service pack – 2 (Operating system) LASER colour photo printer.

7. UPS with BACK UP of 5 Hrs. Suitable mouse and KEY BOARD

8. Fully loaded with MS office, Adobe reader, Photo Editing software and media players (suitable to play recorded bronchoscope videos) with DVD/CD Burning software (Latest versions)

9. Suitable computer table ACCESSORIES FOR SCOPE

10. Cleaning Brushes (1 No.) Cylinder Cleaning Brushes (1 No.)

11. Needles for Trans bronchial Needles FNAC/Aspiration-4 nos Biopsy forceps with windows (4 nos.)

12. Cytology brushes (4 nos.) with suitable handles

13. Protected specimen brush (PSB) with sheaths (20 nos.) with suitable handles Alligator biopsy forceps ( 2 nos. )

14. Vertical stand to be provided for storing and drying the scope after use. Fogarty Balloon Catheter- (Inflatable) 2 Nos-

15. Stabilizer mobile trolley

16. sterilization tray for scope and acessories

17. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; GENERAL SURGERY

Harmonic Scalpel

Harmonic Scalpel :

1. System should have a universal connector to connect Ultrasonic energy and Advanced RF energy instruments.

2. Ultrasonic and bipolar energy in the system must work separately and at no point in combination.

3. System should have automatic instrument recognition

4. System should have a touch screen display for fast and setup, operation and on-screen diagnostics.

5. System should have a high-resolution display with wide viewing angles

6. System should have the ability for software updates via USB memory stick

7. System should be a single generator that provides Ultrasonic energy and Advanced RF energy technology for soft tissue dissection and vessel sealing

8. System should have a potential equalization terminal for compatibility with other medical systems requiring such connections

9. System should conform to the following international standards EN (IEC) 60601-1, EN (IEC) 60601- 1-2, EN (IEC) 60601-2-2, EN (IEC) 60601-1-8

10. System should provide Class 1 protection against electric shock

11. System should provide Class 1 protection against electric shock

12. System should have the ability to select hand switch or footswitch activation or both for Ultrasonic and advanced RF energy instruments and the ability to change selection during use

13. System should have English language as default

14. System should not have minimal lateral thermal spread more than 1 mm.

15. System should not have an auto switch off mechanism.

16. System should have standby mode to ensure safety.

17. System should come equipped with system diagnostics and troubleshooting guide to pin point any problems in the systems.

18. System should have onscreen warning display system for generator overheating, generator software upgrade, handpiece errors and instrument errors

19. System should be able to power ultrasonic energy instruments with 55.5 KHz frequency and can power ultrasonic energy instruments in the frequency range of 30-80 KHz in future

20. The hand piece for the system should come with an inbuilt transducer

Harmonic Scalpel :

21. System should be compatible for open surgery and for laparoscopic surgery

22. System should be compatible with both 5mm and 10mm instruments.

23. System should have atleast 5 power settings levels with power level display for ultrasonic energy instruments

24. System should be able to power energy instruments with microprocessor controlled bipolar electrosurgical radiofrequency technology with a quasi-sinusoidal forced impedance output.

25. System should be equipped with smart advanced RF energy technology to measure the tissue impedance and control the power delivery.

26. System should be equipped with advanced RF energy technology that can simultaneously seal and transect vessels up to and including 7mm, large tissue pedicles and vascular bundles

27. System should be equipped with advanced RF energy technology that provides temperature- controlled energy delivery which should maintain tissue temperature approximately at 100 degree Celsius.

28. System should have Advanced RF Energy hand instruments with a unique electrode configuration to minimize the lateral thermal spread.

Harmonic Scalpel :

29. System should have Advanced RF Energy hand instruments with technology to deliver high compression uniformly across seal area.

30. System should have Advanced RF Energy hand instruments that provide tissue / vessel seal strength to withstand bursting pressure of 7 times the systolic pressure.

31. All hand probes for open and lap procedures should be able to simultaneously cut and coagulate tissues.

32. System should be able to power advanced RF energy hand instruments of 5mm shaft diameter for both open & laparoscopic procedures with round trip (5mm tip width) in the following shaft lengths (14cm, 25cm , 35cm & 45cm) and should be both hand & foot activated.

33. Systems should be able to power ultrasonic energy hand instruments of 5mm shaft diameter for both open & laparoscopic procedures with the following specifications

System should comprise of the following Hardware:

1. Generator

2. Footswitch & Cable Accessories:

3. Handpiece (Transducer)

4. Handpiece (Blue)

5. Generator Cart

6. Adaptors for ultrasonic

7. advanced RF energy instruments

Harmonic Scalpel :

Open Surgery Instruments (Ultrasonic cutting and coagulation device):

1. 9cm shaft, curved, tapered tip for precise dissection, seals 5 mm vessels, as well as lymphatic with16 mm active blade & 240-degree activation, triggers support multiple hand positions.

2. 17cm shaft, curved, tapered tip for precise dissection, seals 5 mm vessels, as well as lymphatic with16 mm active blade & 240-degree activation, triggers support multiple hand positions.

3. 5mm Hand Activated Curved Coagulating Shears capable of sealing blood vessels upto 5mm in diameter, 23 cm shaft length, ergonomic handle

4. Curved Blade having telescoping shaft (10cm-14cm) with integrated hand activation control buttons.

5. Dissecting Hook having telescoping shaft (10cm-14cm) with integrated hand activation control buttons.

Open Surgery Instruments (Bipolar vessel sealing device):

1. Hand probes with 5mm shaft diameter, 14cm long with 5mm tip width.

2. Hand probes with 5mm shaft diameter, 25cm long with 5mm tip width.

3. Hand probe with, 22cm long shaft and 40mm jaw length

Laparoscopic Surgery Instruments (Ultrasonic cutting and coagulation device):

1. 5mm Lap Hand Activated Curved Coagulating Shears capable of sealing blood vessels upto 5mm in diameter, 36 cm shaft length, ergonomic handle.

Harmonic Scalpel :

2. 5mm Lap Hand Activated Curved Coagulating Shears capable of sealing blood vessels upto 5mm in diameter 45 cm shaft length, ergonomic handle.

3. 5mm Lap Dissecting Hook, 32 cm long

Laparoscopic surgery instrument (Bipolar vessel sealing device):

1. Laparoscopic probe probes with 5mm shaft diameter, 35cm long with 5mm tip width.

2. Laparoscopic probes with 5mm shaft diameter, 45cm long with 5mm tip width.

3. Articulating Laparoscopic probe probes with 5mm shaft diameter, 35cm long with 5mm tip width, Probe should have ability to articulate 45- 50* both sides.

Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Rates for all consumables items should be quoted separately in price price-bid.

Department ;    GENERAL SURGERY

Portable Suction Machine

Portable Suction Machine :

1. Vaccum /LPM : - 700 MM Hg , 50 Litres/Min

2. Pump Type- Double rotary vane type

3. flutter free vacuum control knob,

4. Wide mouthed 2 x 2 LITRE (light weight, unbreakable and clear) with self-sealing bungs and mechanical over flow safety device.

5. Noise (in dBA)- 50 dB

6. Collection container & its cap, suctions tube tips, a vacuum gauge, two sets of moisture & microbial filters and control knob

7. SiliconeTubing:8 mm ID x 2 mtr (PVC), 2x2 lt jar (one set

8. It should be Mobility, portability.

9. Equipment shall operate on 220V-240V, 50 Hz, single phase electric supply.

10. ISO/ BIS approve model should be offered.

Department ; MICROBIOLOGY

Analytical Balance : upto 200g/1gm increment

Analytical Balance : upto 200g/1gm increment :

1. Microprocessor based single pan Analytical Balance with High accuracy & precision is required. Reading of the weight by digital display.

2. Weigh accurately up to 3rd decimal place

3. Auto self-calibration facility

4. Auto zero Setting, TARE setting

5. One touch calibration

6. Weighing capacity up to 200g

7. Readability and repeatability 0.001g

8. Stabilization time < 5 second

9. Liquid Crystal Display (LCD) for display

10. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department : MICROBIOLOGY

BOD Incubator

1. Double walled modular structure with 3” thick PUF insulation

2. Interior illumination

3. Inside full length observation glass door with secure gasket

4. A key lock function is provided so that settings may not be changed unintentionally

5. Racks & trays – Stainless Steel with adjustable height- 3 ADJUSTABLE SHELVES

6. Compressor with CFC Free R 134 A / R 404 (Eco Friendly) refrigerant having state of art by- pass system for continuous run of compressor- CFC FREE REFRIGERANT

7. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department : MICROBIOLOGY

Centrifuge Machine

1. Should have a maximum speed of 5000 RPM with steeples regulator

2. Should be supplied with safety lid and lock. (Brushless Motor)

3. Should have digital speedometer and timer.

4. Should have imbalance detector and automatic cut off.

5. Standard mark on quality and safety is Compulsory.

6. Should work on a220VAC +/- 10 %, 50Hz AC Supply.

7. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department : MICROBIOLOGY

CO2 Incubator / Candle Jar

1.   Capacity: 165 L

2.   Interior chamber: Stainless steel for easy cleaning and decontamination

3. Unique Six-Sided Direct Heating to ensure stable temperature control, excellent uniformity, and rapid recovery with no overshoot. Fanless convection circulation to provide chamber homogeneity, eliminate vibration & reduces sample evaporation.

4. Hepa Filters (99.98% efficient) at the inlet to minimize contamination should be provided at the inlet.

5. Timer: 1 min. to 100 hours and hold position-Dual heat sterilisation utilises the incubator’s two heaters during the 180°C sterilisation process, which takes 11 hours.

6. Temperature range: AT +5 to +50, ±0.1

7. Temp Accuracy ±0.1 of required temp, with inbuilt

8. Temp. Sensor - Dual IR

9. Audiovisual Alarm to Indicate when temperature deviates more than 1°C from set point, and when program or time has finished. Alarm may be muted.-( power failure, Temperature Deviation, High temperature , CO2 deviation, Door open)

10. There should be a Membrane Keypad with LCD/LED to set and display operating parameters, current status, running time and alarm conditions for time and temperature.- Colour LCD touchscreen

11. Internal glass door for the observation

12. Dual decontamination cycle at 95 deg C (wet decontamination) and at 145 deg C (-Dry heat sterilisation, 180°C, 11 hours) should be available

13. CO2 Range - 0 to 20, CO2 Accuracy: ±0.15; CO2 Inlet pressure 1.5 bars (app)

14. Compensation: Temperature compensation @ 0.5 deg C / min and CO2 Compensation up to AT +5 to +50, ±0.1. Temperature uniformity, ±0.25

15. High Humidity – Chamber to achieve Humidity 95 %, ±5, minimizing sample evaporation. Independent door heater to eliminate condensation on inner glass surfaces should be available.

16. A colour LCD touchscreen panel allows full control, even with gloved hands. Transfer of data is easy via a USB port. The easy-to-clean incubator interior features fully rounded corners and integrated shelf supports

17. USB port available to transfer the data

18. 4 x Stainless steel copper-enriched alloy

19. Interior lighting facility, insulated door fitted with heavy hinges handles locking, mechanical door lock/Electric door lock with password.

20. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department : MICROBIOLOGY

Deep Freeze -20° C & Deep Freezer

| 1. Minimum Capacity should be 300 litres |  |
| --- | --- |
| 2. Digital display of set and actual temperature, with audiovisual alarm |  |
| 3. No condensation on storing material with automatic electric defrost |  |
| 4. Construction - Solid rust free cabinet to prevent corrosion and lockable castor wheels. |  |
| 5. Refrigeration System - Heavy Duty refrigeration system, maintenance free, below -20°C (+ 10°C) cascaded connection with hermetically sealed refrigeration compressors and reliable refrigeration to minimize noise and vibration, air cooled with security lock to prevent unintentional switch off should be supplied. It should have cooling time within hours at maximum ambient temperature of 33°C. The equipment should be of continuous duty and frost free. |  |
| 6. Alarm- It should also have audio visual Electronic Alarm System for |  |
| a) | Door open |
| b) | Power failure |
| c) | Other mal functions |
| 7. Insulation - High density polyurethane or equivalent Gaskets - Double seal silicon. |  |
| 8. Should have automatic high and low voltage compensator |  |
| 9. Temperature homogeneity between top and bottom shelves should be maintained + 1 to 2°C of set temperature |  |
| 10. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO. |  |

Department : MICROBIOLOGY

Distilled water Plant

1. Distillate output Production rate- 4 litres/hr

2. Conductivity, 0.2μs/cm

3. Minimum water supply pressure, 20.3kg/cm

4. Water consumption, 60litres/hr

5. Electrical supply 230/240V, 50/60Hz 13 amp

6. Spare heating elements - 1set

7. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department : MICROBIOLOGY

Elisa Reader

1.   Should have 96 wells and should reading capability of 1 to 96 wells individually.

2.   0 – 3 Abs, ± 2%, 96-well plate, fast mode 0 – 4 Abs, ± 2%, 96-well plate, normal mode 0 – 2.5 Abs,

± 2%, 384-well plate, fast mode

3. 0 – 3 Abs, ± 2%, 384-well plate, normal mode"

4. Wavelength range 340 – 850 nm

5. Photometric accuracy (405 nm) ± 1% (0.3 – 3 Abs), ± 2% (3 – 4 Abs)

6. Resolution 0.001 Abs

7. Linear shaking with three modes: slow, medium and fast

8. 8-position filter wheel, the instrument is delivered with the following standard filters installed:

9. 405 nm; 450 nm; and 620 nm. Additional filters can be ordered separately."

10. It has provision to add 4 additional filters.

11. Should have automatic filter selection.

12. Should have automatic calibration before each reading.

13. It can storage calibration curves datas which can be exported with USB memory stick.

14. Capable of doing multi standard tests and controls.

15. Should have different types of blanking facility like air wise and well wise.

16. Should have capable of reading U,V and flat type wells.

17. Light source Quartz-halogen lamp 6 V / 10 W

18. USB for external printer

19. Laser printer with computer will be provided

20. "During start up all major functions of the instrument, such as plate position,

21. measurement stability, lamp functionality, filters, optical system, incubation and electronic operation, are checked to ensure reliable operation."

22. Adjustable 8 channel micropipette (40-300 μl) 1 no. will be supplied

23. System work with input 200 to 240Vac 50 Hz supply.

24. High contrast color display (480 x 272 dots) and storage of plates readings.

25. Should have graphical display of plate layout for specifying controls, standards blanks etc.

26. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO..

27. UPS of suitable rating with voltage regulation and spike protection for 30 minutes back up will be provided

28. Should be provided with Start-up Kit/Regent/ Calibrators/ Consumables/ Accessories are required performing the machine initial 500 tests or 3 Months test which is higher.

Department : MICROBIOLOGY

Elisa Washer

It should have the capability to wash flat, U or V bottomed micro plates or 8 strip plates.

It has 8 way manifold.

Should have 25 wash program software memory or more.

Should have programmable washing time, volume and soaking time.

Should have minimum 6 wash cycle.

Should have residual volume less than 2 ul with proper demo during technical demonstration.

It has peristaltic pump to ensure accurate and quiet washing.

Should have waste bottle with full bottle alarm or sufficient mechanism to avoid spillage and damage to equipment.

Should have suction based washed buffer intake.

Should work with input 220 to 240 Vac 50 Hz supply.

Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

UPS of suitable rating with voltage regulation and spike protection for 30 minutes back up will be provided

Capability for adjustable dispensing & aspirating pressure rate settings. (Optional)

Calibrated at every 6 months from the date of installation will be done at free of cost.

Should be provided with Start-up Kit/Regent/ Calibrators/ Consumables/ Accessories are required performing the machine initial 500 tests or 3 Months test which is higher.

Department : MICROBIOLOGY

Laminar flow

1. Type of Flow: Vertical – Re-circulatory

2. HEPA FILTER : Face dimensions: 4ft (L) X 2ft (W) X 6 ft The HEPA filter should have rated efficiency of 99.97% (or better) at 0.3 microns to provide product protection of Class 100 or exceeding Class 100 requirements of Federal Standards 209E or

3. Equivalent ISO within the work. Area

4. PRE Filter with Synthetic, non-woven polyester fibers having casing of namel painted CRCA frame with Retention of 10 - 15 Micron and 90 % Efficiency. Washable with an arrestance of 90% or better

5. Material of construction: Main body and rear panel: Electro-galvanized steelor Mild Steel, oven baked epoxy powder coated finish. Side window (panels): UV stabilized transparen’t Perspex or polycarbonate. Worktable (surface): SS304 or SS316

6. Working area should be 24 cuft.

7. Blower Assembly: DIDW type blower system with high RPM motor, enclosed in an powder coated MS casing suitably suspended in a pair springs & connected to the filter chamber through flexible canvas duct.

8. Front Windows Acrylic, fixed by clamps

9. Illumination with Fluorescent tubes with diffusers. Light Intensity at Work Surface :800- 1000 lux/75-90 foot candles

10. Laminar Airflow Velocity: Approx. 90 feet per minute (fpm)+/-10% average velocity measured 50 mm from the filter face. Uniformity +/-20% of average or better.

11. Additional Requirement: Vibration free Gas burner facility on working bench .Air pressure indicator with manometer (Differential Pressure Gauge with Scale display in cms of water). Drain valve with smooth drainage arrangement. Exhaust ducting as per site requirement

12. Noise level

13. UV Germicidal lamp intensity > 40 microwatt/sq. cm. over the entire work surface

14. Switched and indicators: Individual switches and indicator lamps for blower motor, florescent lamp and UV lamp.

15. System Configuration Accessories, spares and consumables

16. System as specified -

17. Spare HEPA Filters and PRE Filters- 2 SETS EACH

18. Other fitting required for attaching auxiliary services are

19. Electrical outlet socket (5 ampere rating) qty: 2 nos. 2.Valves for gas service-one each for gas and vacuum.

20. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department : MICROBIOLOGY

~ Automated Blood Culture System

1. Fully automated, technology with ability to take patient I.D. by barcode.

2. Should process blood samples, other sterile body fluids both aerobic and anaerobic systems.

3. Sample capacity more than 200 samples

4. Besides pyogenic, system should have facility of detection for yeasts and fastidious organisms.

5. Capacity to include pediatric and adult samples.

6. Media in bottles should have agents for neutralization of antibiotics.

7. Continuous agitation system to allow better organism growth

8. Should analyze each sample separately as per ID, time of entry, incubation period, growth etc.

9. Should have built in calibration check and alarms/ reminders for the same.

10. Decontamination facility should be available for the system as well as individual rack

11. System should have high sensitivity & specificity with continuous monitoring of all samples.

12. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO..

13. Should be provided with Start-up Kit/Regent/ Calibrators/ Consumables/ Accessories are required performing the machine initial 500 tests or 3 Months test which is higher.

Department : MICROBIOLOGY

- Colony Counter

1. Accepts 6 cm, 9 - 15 cm petry dish

2. Maximum counts 999 large display

3. Display showing the present and 3 previous count values at the same time

4. With Pressure sensor system, illumination LED device

5. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department : MICROBIOLOGY

Binocular Microscope (Student)

1. Binocular microscope with universal infinity corrected optical system

2. Halogen light source illumination

3. Revolving Quadruple nose piece (for objectives)

4. Binocular observation tube with inclination of 45/30 degrees

5. Built in torque adjustable focusing knob

6. Square mechanical stage with rigid hand coaxial control

7. Abbe condenser, Iris diaphragm

8. Revolving Quintuple nose piece (for objectives)

9. Plan achromat objectives 4X, 10X, 40X, 100X (Oil)

10. 40X, 100X objective should be spring loaded

11. Eye piece 10X (FOV 20)

12. Antifungal treatment should be applied to the observation tube, eyepiece and objective

13. Accessories, dust cover and power cord

14. Power requirement 220 V/50 Hz

15. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department : MICROBIOLOGY

Refrigerated Microcentrifuge

| 1. | The Centrifuge should be a Microprocessor controlled Bench Top Model. |
| --- | --- |
| 2. | Should have Speed of 15,000 -16,000 rpm and RCF 20,000xg or more. |
| 3. | Temperature Range: -10degreeC to +40degree C. |
| 4. | Should have preselection of time 10 sec to 99 hours and continuous run. |
| 5. | Should have a LCD display for set & current run parameters. |
| 6. | Should have a Maintenance-free / brushless induction drive |
| 7. | The lead lock should be motorized. |
| 8. | The instrument should have an automatic Rotor and Imbalance Identification system. |
| 9. | Should be equipped with an over speed protection. |
| 10. Should have a program memory. 11. At least 9 acceleration & 9 deceleration rates. |  |
| 12. Should have a key for short spin |  |
| 13. Pre-cooling function should be there. |  |
| 14. Noise level should be ≤ 60dB |  |
| 15. Selection of speed in both RPM & RCF should be there. |  |
| 16. Suitable voltage stabilizer should be provided for the equipment. |  |
| 17. Equipment should be supplied with following rotors and adaptors: |  |
| a. Fix Angle /swinging-bucket Rotor for 24x1.5/2.0 ml with speed of 15,000-16,000 rpm and RCF 20,000xg or more |  |
| 18. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO. |  |

Department : MICROBIOLOGY

Autoclave

| Horizontal rectangular or vertical cylindrical high-pressure steam sterilizer |  |
| --- | --- |
| Chamber capacity: 75 litre |  |
| Electrical Power : 18/36 KW or Sufficient wattage of Industrial immersion type water heater to generate steam within a reasonable period of time on 3 phase 440 v 50 HZ ac supply |  |
| Working pressure and Temperature : 1.06 Kgf / sq.cm -1.2 Kgf /sq.cm at 121 Deg C |  |
| Material of construction |  |
|  | Inner chamber, Jacket, Door : SS 316.(5mm-10mm) |
|  | Outer Chamber : SS 304 (insulated properly) |
|  | Steam generator : Non corrosive SS / Chromium plated Brass |
|  | Heater Plate : Brass / Stainless steel |
|  | Pipe Line : Complete with SS |
|  | Stand : Stainless Steel/High quality non corrosive Steel Instrumentation |
| Temperature, Pressure and vacuum gauges, steam traps, vacuum driers, water level indicator on steam generator Safety devices. |  |
| Pressure switch and safety valve, self-locking of door when chamber is under pressure: vacuum breaker for jacket. |  |
| Steam generator with gauge glass valves and Low water protection with audiovisual indicator. |  |
| Entire cycles of sterilization should be controlled by a microprocessor unit and it should also have manual operation facility. |  |
| Accessories |  |
|  | Additional Accessories required: |
| 1. Heating element 2 set extra with each machine. |  |
|  | Requirements at installation: |

1.    Necessary Pipeline works for water inlet and steam outlet should be done.

        Cable for connecting to wall socket should also be provided

Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department : MICROBIOLOGY

Biosafety Cabinet Type - 2A

1. Protection for operator, environment and the product, from aerosols and microorganisms

2. Microprocessor / Microcontroller / Microcomputer controlled system.

3. External Antimicrobial powder coated

4. ULPA Filter with efficiency 99.999%

5. Automatic speed compensation system against clogged main HEPA filter Pre-filtration unit with retention of 10 to 15 micro meter

6. Air Circulation to vertical with 30% exhaust and 70% recirculation

7. Single stainless steel perforated working platform

8. Alarms for power failure and door opening

9. Should be fitted with UV light > 800 lux

10. High-speed centrifugal blower with lifetime lubricated

11. Noise level <58dBA, Elapsed hour counter

12. DOP test outlet

13. Fluorescent lamp to obtain powerful glare-free lighting

14. On site installation and appropriate certificate to be provided

15. On/Off switch with key lock.

16. Gas connection should be provided in the cabinet

17. Quote for BOP tested HEPA filters separately

18. Shall meet IEC-60601-1-2 : 2001(Or Equivalent BIS / ISO) General Requirements of Safety for Electromagnetic Compatibility.

19. Accessories, consumables and miscellaneous:

20. One filter set replacement should be included in CMC once in a year

21. All consumables required for installation and standardization of system to be given free of cost

22. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department : MICROBIOLOGY

Hot Air Oven

1.   Double Walled, inner wall made of aluminum and outer wall Mild Steel sheet.

2.   Shelves: 3 to 5 stainless steel, perforated shelves, could be adjusted at any levels.

3. Heating Elements: placed at the bottom or sides.

4. Temperature control: Thermostatic control from 40 to 250 o C with Regulator and digital temperature-indicator.

5. Automatic timer 0 - 60 minutes.

6. Air circulating fan.

7. Volume : 200 L

8. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department : MICROBIOLOGY

Incubator

1. Capacity: 120 L

2. Interior chamber: SS 316 for easy cleaning and decontamination

3. Timer: 1 min. to 100 hours and hold position

4. Minimum turbulence and no cross contamination

5. Adjustable safety thermostat for temp setting at 1 deg C increment

6. Temp Accuracy : +/-1%

7. Internal glass door for the observation

8. With minimum two adjustable shelfs

9. Audiovisual Alarm to Indicate when temperature deviates more than 1°C from setpoint, and when program or time has finished. Equipment should have provision to mute the alarm.

10. Peltier heating with continuous air circulation and Heating by natural/forced convection for homogenous temperature distribution

11. Temperature range: +5° C to 60°C

12. There should be a Membrane Keypad with LCD/LED Display to set and display operating parameters, current status, running time and alarm conditions for time and temperature.

13. The Equipment should have Interior lighting facility, and insulated door fitted with heavy hinges

14. Power Supply: 230V, 50 Hz

System Configuration Accessories, spares and consumables

1. Incubator as per specification

2. All consumables required for installation and standardization of system should be provided free of cost

3. Should be supplied with all the accessories required for the functioning of the equipment.

4. Power Cord

Standard, Safety

Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department : MICROBIOLOGY

Water bath with variable temperature

1. Should have a double walled construction.

2. The inner chamber and top lid should be made of stainless steel.

3. The space between the two walls should be packed with thick glass wool.

4. Should provide with a thermostat control.

5. Working temperature should be from ambient+5 ºC to 80ºC having an accuracy of +/- 1ºC

6. Should have an approximate inner chamber dimension of 350 mm x 250 mm x 125 mm.

7. Should be operated on 230V, 50 Hz single phase AC supply

8. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department : MICROBIOLOGY

- 80-degree deep freezer with UPS

1. Should have galvanized steel body with epoxy paint/ HDGI Sheet -Powder coated and vacuumed polyurethane foam panels, outer double/single door with locking facility.

2. Should have alarm for audible and visual fault acknowledgement, high and low temperature audio visual alarms, open door alarm and power failure alarm.

3. Castor wheels and leveling adjustors should be provided for adjustment and installation.

4. Capacity:

Total storage capacity of 320 bags of 250 ml plasma bags

5. Refrigerant:

a. CFC free or HCFC free refrigerants with Biodegradable oil Compressor

6. Cooling System:

a. Cascade Cooling system

b. It should have 2 stage compressors with standard refrigerants.

c. Should have two compressors sealed with high grade seal.

7. Doors:

Single outer door foaming insulated with patient technology, left hinged, main door with CAM lock

8. Inner Compartment :

Minimum 3 compartments with independent doors (304 SS)

9. Temperature:

a. LED Display, temperature sensor probe (PT100)

b. Programmable temperature range at least minimum -80 degree Celsius in increment of 1 degree Celsius

c. Range -55 to -80 degree Celsius, uniformity ±3ºCelcius

| 10. Additional accessories: |  |
| --- | --- |
| a. SS Racks and Boxes (Plastic, cryo with racks) |  |
| b. System monitoring and reporting technology systems built-in for diagnosis or set point variants |  |
| c. | System should have provision for software to control and monitor up to 30 freezers simultaneously. |
| d. Noise level < 50 DB |  |
| e. Stabilizer with enough capacity should be provided |  |
| f. | Continuous temperature monitoring facility and temperature chart that can be replaced weekly. |
| g. | Temperature chart paper with pen should be provided. |
| 11. Power Supply: Should operate from AC 220 VAC ± 10 %, 50 Hz & with Suitable capacity Servo Stabilizer. |  |
| 12. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO. |  |

Department : OBSTETRICS & GYNAECOLOGY-DEPARTMENT

Cryo / electro cautery apparatus

1. Should have interchangeable gun probes of six different sizes.

a. Dia 25 mm – 8 mm

b. 2 mm – 15 mm

c. 22 mm – 25 mm

d. 19 mm – 13 mm

e. 19 mm – 10 mm

f. 8 mm – 10 mm

2. With a carrying case for keeping the gun.

3. Should be hand operated , light weight & have easy maneuverability

4. Should also be equipped with N2O / CO2 gas cylinder, with connecter with a filter, pressure regulator & trolley for cylinder and cylinder key.

5. Mains Power : 230 V,50 Hz.

6. USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO Approved model should be offered.

Department : OBSTETRICS & GYNAECOLOGY-DEPARTMENT

NST machine

1.   Color LCD Fetal Monitor

2.   7” TFT Color LCD Display

3. Saving patient data of more than 1500

4. 1 MHZ Sensitivity

5. Battery Backup should be present

6. Twin Monitoring should be made available

7. Built in thermal printer for fetal heart trace

8. Tocodynamometer for recording uterine contractions

9. Light Weighted – to be carried with hand.

10. Thermal paper - 30 rolls to be provided along with the machine.

11. Should be provided with 2 USB probes & 1 Toco probe.

12. Central Network monitoring facility

13. Alarms – Fetal Heart rate – Maximum and Minimum , Unit function, No paper

14. Should be able to record the data in digital form in computer

15. Should work continuously for CTG for 10 hours.

16. One central station with monitor and Laser jet printer facility to archive 15,000 patient data. Wiring of individual foeatal monitor to central station to be included in the scope of supply.

17. Cart to keep the foetal monitor

18. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO..

Department : OBSTETRICS & GYNAECOLOGY-DEPARTMENT

Cystoscope & Resectoscope

Cystoscope with Light source, camera & Monitor

| 1. Sheath with obturator outer Diameter - 20 Fr |  |
| --- | --- |
| 2. Working channel- 04 - 05Fr |  |
| 3. Working length - 22cms |  |
| 4. Self - dilating tip for ease of entry |  |
| 5. Compatible 30 degree an 0-degree HD telescope – 01 each |  |
| 6. Telescope bridge – 01 |  |
| 7. Stone crushing forceps compatible with above sheath and telescope – 01 in no. |  |
| 8. Resectoscopes sheath compatible with above mentioned telescope – 01 in no., with |  |
| a. Moveable thumb ring (working element) – 01 in no. |  |
| b. Spring loaded action with electrodes inside the sheath in rest position |  |
| c. | Telescope Bridge – 01 in no. |
| d. Cutting loop – 10 in nos. |  |
| e. Coagulating hook electrode – 10 in nos. |  |
| f. | Monopolar high frequency cord – 05 in no. |
| g. | Protection tube for electrodes – 05 in no |
| 9. Fiber optic cold light source: 1no. 300 watt Xenon bulb / LED (Bulb life 500 hrs minimum for xenon & LED Life 10000 hrs. |  |
| 10. Should have 3 cheap camera |  |
| 11. Should have 21” Medical grade HD Monitor. |  |
| 12. Power supply 220VAC +/- 10 %, 50Hz. |  |
| 13. On-line UPS with 60 minutes back up Entire system. |  |
| 14. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO. |  |

Department : OBSTETRICS & GYNAECOLOGY-DEPARTMENT

High End suction machine

1. The Suction pump should be oil immersed fitted on Motor shaft

2. Suction pump should have line grinding internally.

3. To facilitate maintenance the cover of machine should be easily to open from the top & sides

4. The suction machine should be capable of producing minimum vacuum of 500 approx mm Hg. Which should be adjustable and monitored by vacuum gauge of suitable range. The suction capacity should be 15 litres per minute and can be regulated.

5. It should have two bottles of 1 or 2 liters (As per requirement) with synthetic rubber lids. The bottle shall be fitted with the arrangement to prevent overflow of fluid.

6. ON / OFF Switch and Power indicator should be available

7. Body material: Base, top & panel made of rust proof and corrosion resistant moulded ABS/Stainless Steel. Jar/Bottle material: Autoclavable polycarbonate.

8. Inbuilt maintenance free battery. Battery backup up to 60 minutes on full charge. Should be with built in charger.

9. Should operate on 230V,50 Hz supply.

10. Accessories, Spares and Consumables

11. System as specifies

12. Powercord 2 metre long with 15A three pin plug top

13. One litre bottle with over flow protection- 02 nos ( one spare set)

14. Lid set – 02

15. Seal for bottle -02 nos

16. Suction tubing set - 04 set

17. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department : OBSTETRICS & GYNAECOLOGY-DEPARTMENT

Vacuum Extractor and suction machine

Table Top Powerful Surgical suction Unit for Fetal delivery

Suction pressure can be adjustable from 0 to 750 mm Hg by finely adjustable knob (not merely pressure release screw) for slow build of vacuum for delivery of fetus

Mains Power : 230 Volts

Suction jar with valves for overflow cut off

40 mm metal suction cup (good quality) – 2 No

50 mm metal suction cup (good quality) – 2 No to be provided (Occipito anterior and Occipito Posterior Cup)

60 mm metal suction cup (good quality) – 2 No

Bacterial Filter – 20 Pcs

40 mm silicone suction cup (good quality) – 1 No.

50 mm silicone suction cup (good quality) – 1 No

60 mm silicone suction cup (good quality) – 1 No

Suction Hose of 2 meters in length

Suction Hose can be easily detachable from the cups

Should also be provided with foot switch for on & off and control of pressure.

With portable trolley (Inbuilt)

Overflow protection device to prevent liquid & solid particles from entering the intermediate tubing.

Light weight machine

Automatic control (Microprocessor based)

Single button control for building vacuum gradually.

Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department : OBSTETRICS & GYNAECOLOGY-DEPARTMENT

~ Fetal Monitor for Antepartum Surveillance

Fetal Monitor for Antepartum Surveillance

1. Should be portable.

2. Should have the measurement of FHR, and Uterine activity.

3. Should be supplied with event marker cable. Automatic identification of events possible

4. Should have an alert indication for probe contact, high/low FHR.

5. Should have a volume control for FHR.

6. Should have large format LED/LCD/TFT display.

7. Should be provided with thermal printer for printing FHR, uterine activity pressure.

8. Thermal paper should be available locally.

9. Should provide alarm limit settings for high, low FHR.

10. Should provide belt for both ultrasound and toco transducer.

11. Should be supplied with SS trolley.

12. Should operate on mains (230v, 50Hz).

13. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

14. Accessories:

i. Gel – 4 bottle

ii. Thermal paper- 10Nos

iii. Voltage Stabilizer

15. Should have built in rechargeable Li –ion battery with backup of at least one hour or suitable UPS need to be supplied .

Department ; OBSTETRICS & GYNAECOLOGY-DEPARTMENT

Suction Machine

Suction Machine :

1.   Vacuum /LPM: - 700 MM Hg , 50 Liter’s/Min.

2.   Pump Type- Double rotary vane type

3. flutter free vacuum control knob,

4. Wide mouthed 2 x 2 LITRE (light weight, unbreakable and clear) with self-sealing bungs and mechanical over flow safety device.

5. Noise (in dBA)- 50 dB A

6. Collection container & its cap, suctions tube tips, a vacuum gauge, two sets of moisture & microbial filters and control knob.

7. SiliconeTubing:8 mm ID x 2 mtr (PVC), 2x2 lt jar (one set extra)

8. It should be Mobility, portability.

9. Equipment shall operate on 220V-240V, 50 Hz, single phase electric supply.

10. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

11. ISO/ BIS approved model should be offered.

Department ; OPHTHALMOLOGY

Automated Perimeter

Automated Perimeter :

1. Specification of Perimeter (Goldman Type) should have following: - High quality Goldman standard Imported automated full field perimeter with bowl size 30cm.Computer monitor should be inbuilt with the perimeter.

2. Maximum intensity 10,000Asb, Bowl illumination 31.5Asb

3. Internal hard disk drive, USB with future upgradation to MOD

4. Stimulation duration 200ms, wavelength Broad band visible light

5. Stimulus/Background colour White on White

6. Maximum temporal range 90Deg.Suitable for central 30 as well as full field testing

7. Central field test patterns 30-2,24-2,10-2,Macula

8. Peripheral field test pattern 60-4,Nasal Step

9. Threshold test strategies full threshold, FastPac, SITA, SITAFast, SITA Standard

10. Screening field test P-60,FF-80,FF-120,FF-240,Nasal Step for periphery

11. Screening test strategies Two zone, Three Zone and Quantify Defects

12. Stimulus Size I-V as per Goldman standards

13. Glaucoma hemifield test ,Heijl –Krakau blind spot monitor

14. Video eye monitoring , Trial Lens Holder,

15. Touch screen monitor as well as Keyboard & Mouse

16. Motorized chinrest, Motorized table with Laser Jet Printer

17. Glaucoma progression analysis software

18. Power supply to be 220-240VAC, 50Hz fitted with Indian plug.

19. Suitable UPS with maintenance free batteries &Back up time 30 minutes.

20. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; OPHTHALMOLOGY

Ophthalmoscope (Direct)

Ophthalmoscope (Direct) :

1. Available with LED/Halogen light

2. Magnification up to x15 from direct vision to maximum magnification.

3. Red-free, blue and polarization filters and Anti-reflection lens.

4. Should have small and large spot sizes, fixation targets, slit aperture, hemi-spot and cobalt blue filter.

5. Should be rechargeable battery with Charger / battery/ mains operated.

6. At least 3 apertures and fixation

7. Range of lenses not smaller than - 30D to +20D with steps not greater than 1D.

8. Dust free sealed optics and aspherical optical system.

9. Internal batteries, rechargeable preferred compatible with both 2.5 V and 3.5 V batteries or handles provided; Led display indicating the charging status.

Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; OPHTHALMOLOGY

Streak Retinoscope

Streak Retinoscope :

1. Should have an external focusing sleeve which is easy to grip and manipulate.

2. Should have crossed-linear polarizing filter.

3. Should allow one-hand operation for streak focus and 360º streak rotation.

4. Should be interchangeable to plane mirror and concave mirror mode by sleeve movement.

5. Should use halogen/Xenon streak lamp.

6. Should have 100% dust proof housing and multi-coated optics.

7. Should have detachable brow rest for spectacle wearer

8. Should be battery/ rechargeable battery operated.

9. Should have a carrying case.

10. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

11. Should be supplied with the following accessories.

Bulb 5 nos.,bulb Holder and bulb cover

Department ; OPHTHALMOLOGY

Indirect Ophthalmoscope

Indirect Ophthalmoscope :

1. Binocular Indirect ophthalmoscope with precision viewing up to 1.0 mm pupil size.

2. Spot size: 3 integrated spot size small spot, medium spot & large spot.

3. Filters: 3 or more integrated filters to choose form red filter, cobalt blue filter and diffuser.

4. Vertical adjustment, +/-4˚

5. Headband with Rheostat and Articulating Hinge to provide vertical adjustment of the rear band.

6. integrated flip up adjustment optics, which can be flipped, and locked at 0˚, 12.5˚, 47.5˚, 60˚

7. Aperture and filter adjustment levers: can be locked to the desired position required.

8. P.D. range from 46-75 mm.

9. 6V halogen Xenon bulb/LED.

10. Transformer compatible with voltage system of AC 220 -240 volts

11. Carrying case

12. Head Band Mounted Rechargeable battery

13. Accessories +20D lens.

14. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; OPHTHALMOLOGY

Keratometer

| Keratometer : |  |
| --- | --- |
| 1. Radius of curvature:5.5 to 10 mm (Increments: 0.01mm) |  |
| 2. Corneal power:33.75 to 61.25 D when cornea equivalent refractive index is 1.3375 (Increments: 0.05D, 0.12D, 0.25D) |  |
| 3. Axis:1˚ to 180˚ (Increments: 1˚) |  |
| 4. Corneal periphery: Measurement area: 30˚ |  |
| 5. when radius of curvature is 8mm |  |
| CORNEAL DIAMETER MEASUREMENT |  |
| a. Corneal diameter:2.0 to 16.0 mm (Increments: 0.1mm) |  |
| b. Retroillumination: Available; |  |
| c. | 2 images can be stored in memory |
| d. Built-in printer |  |
| e. Should work on 230 V,50 Hz |  |
| f. | Power-saving system should be available |
| Accessories , Spare and Consumable |  |
| g. | Motorized table |
| h. Keratometry model eye (with contact lens holder) |  |
| i. | Printing paper |
| j. | Chin rest paper |
| k. | Blower brush |
| l. | Dust cover |
| 6. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.. |  |

Department ; OPHTHALMOLOGY

Schiotz’s  Tonometer

Schiotz’s Tonometer :

1. High quality agate bearing for an extremely long service life.

2. Precision measurement on a scale of 0 to 20 subdivision and 0 to -1 subdivisions.

3. Reading of the scale with red pointer

4. Complete in a case with velvet-look inserts

5. Supplied with three weights (5.5 g, 7.5 g, 10 g) and a conversion table.

6. It should be Autoclavable.

7. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; OPHTHALMOLOGY

Keratometer

| Keratometer : |  |
| --- | --- |
| 1. Radius of curvature:5.5 to 10 mm (Increments: 0.01mm) |  |
| 2. Corneal power:33.75 to 61.25 D when cornea equivalent refractive index is 1.3375 (Increments: 0.05D, 0.12D, 0.25D) |  |
| 3. Axis:1˚ to 180˚ (Increments: 1˚) |  |
| 4. Corneal periphery: Measurement area: 30˚ |  |
| 5. when radius of curvature is 8mm |  |
| 6. CORNEAL DIAMETER MEASUREMENT |  |
| a. Corneal diameter:2.0 to 16.0 mm (Increments: 0.1mm) |  |
| b. Retroillumination: Available; |  |
| c. 2 images can be stored in memory |  |
| d. Built-in printer |  |
| e. Should work on 230 V,50 Hz |  |
| f. | Power-saving system should be available |
| Accessories , Spare and Consumable |  |
| g. | Motorized table |
| h. | Keratometry model eye (with contact lens holder) |
| i. Printing paper |  |
| j. | Chin rest paper |
| k. Blower brush |  |
| l. | Dust cover |
| 7. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO. |  |

Department ; OPHTHALMOLOGY

Schiotz’s  Tonometer

Schiotz’s Tonometer :

1. High quality agate bearing for an extremely long service life.

2. Precision measurement on a scale of 0 to 20 subdivision and 0 to -1 subdivisions.

3. Reading of the scale with red pointer

4. Complete in a case with velvet-look inserts

5. Supplied with three weights (5.5 g, 7.5 g, 10 g) and a conversion table.

6. It should be Autoclavable.

7. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; OPHTHALMOLOGY

Cryo Unit

| Cryo Unit : |  |
| --- | --- |
| 1. Should have interchangeable gun probes of six different sizes. |  |
| 2. Should have interchangeable gun probes of six different sizes. |  |
| 3. Diameter |  |
|  | 25mm – 8 mm |
|  | 2 mm – 15 mm |
|  | 22 mm – 25 mm |
|  | 19 mm – 13 mm |
|  | 19 mm – 10 mm |
|  | 8 mm – 10 mm |
| 4. With a carrying case for keeping the gun. |  |
| 5. Should be hand operated , light weight & have easy maneuverability |  |
| 6. Should also be equipped with N20 / CO2 gas cylinder, with connecter with a filter, pressure regulator & trolley for cylinder and cylinder key. |  |
| 7. Mains Power : 230 V,50 Hz |  |
| 8. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO. |  |

Department ; OPHTHALMOLOGY

Operation Theatre Table

Operation Theatre Table for Ophthalmology :

1. Length Should be 1840 mm-1880 mm

2. Width should be 600 mm – 700 mm

3. It should be motorized height adjustment with average speed of 8 -10 mm/ sec

4. Height range 580 mm – 810 mm

5. Manual controlled up & down motion head rest

6. Motor should be CE Make

Department; OPHTHALMOLOGY

Anesthesia Machine and Accessories like Laryngoscope, all size Endotracheal Tubes, Nasal and Oral

Airways ,  Maill ’s  Forceps

Anesthesia machine and accessories like laryngoscope, all size endotracheal tube, nasal and oral airway, Magill’s forceps:

1. Powder coated structure.

2. Active Anti Hypoxic Electronic device - to ensure no nitrous flow with loss of O2 pressure.

3. Ratio Control - 25-28% mixed ratio between 02 & N2O Gas flow

4. Should have Electronic Flow meter for O2, N2O & Air

5. Should have electronic display of Gas pressure in the user interface display

6. Cylinder Pressure and operating pressure should be indicated electronically in user interface display.

7. Oxygen Driven Nitrous Supply through Electronically

8. Inner connection of tubing’s should be made of alloy metal or PPU tubes

9. Audio Visual alarms should be available for Oxygen Failure warning Device

10. Electronically measured cylinder pressure, quick-release system for docking and release.

11. Electronically/pneumatically controlled technology valve for High precision pressure relief cum non return valve should be available

| 12. The vaporizer should be electronic / Manual with an injector. The container of the vaporizer chamber should be minimum 300 ml (10,1 oz.). Refilling of inhalation agent should be possible during running mode. The liquid level in the vaporizer should be displayed both on vaporizer and control display and an alarm should be given when the level becomes too low. The desflurane vaporizer should not require to be warmed up before. |  |
| --- | --- |
| 13. Provision to connect two electronic / manual controlled vaporizer simultaneously. |  |
| 14. Switching between manual and automated ventilation should be possible. The system shall contain a continuously adjustable pressure relief safety valve settable between 0-70 cm H20. Total system volume should be as low as possible and not exceed 3 liter including absorber. |  |
| 15. Gas Inlet port for O2 & N2O, pin- index type 1 each & O2, N2O & Air non-interchangeable type hose 1 each |  |
| 16. Direct Central Pipe Line connectivity to Machine. |  |
| 17. Standard Maggils circuits, all rubber antistatic tested. |  |
| 18. Alloy metal tray for working / writing table. |  |
| 19. Castor wheel with Individual breaks. 4 wheels with individual locking brake. |  |
| 20. Double Oxygen outlet available for circuirs |  |
| 21. Operating gas pressure available on the user interface |  |
| 22. 2 - B” type cylinder carrying facility at back side. |  |
| 23. Circle Absorber shall contain a continuously adjustable pressure relief safety valve settable between 0 - 70 cm H20. Total system volume should be as low as possible and not exceed 3 liter including absorber. |  |
| 24. Ventilator Specifications: |  |
|  | The equipment shall contain functions for volume- and pressure regulated ventilation. |
|  | The ventilator shall contain the functions” Spontaneous Breathing” and” Manual Ventilation”. |
|  | The ventilator shall have the following modes: VCV, PCV, PRVC, SIMV |
|  | The ventilator shall have Pressure Support mode with backup functionality |
|  | The ventilator should not contain any moving parts |
|  | The equipment shall handle high- to low flow anesthesia during both non-rebreathing and partial rebreathing conditions. Switching between different breathing systems shall be a simple operation. |
|  | It shall be possible to regulate the I:E ratio between 4:1 to 1:8 |
|  | The equipment shall include an integrated continuous PEEP-function ranging from 0 to at least 20 cm H2O or more. |
|  | The set PEEP level should be visible on the control display |
|  | The set pressure should be adjustable between 5 - 70 cm H20 |
|  | The set tidal volume should be adjustable between 20 - 1500 ml. |
|  | The breathing frequency should be adjustable between 3 - 80 breaths/min or more |
|  | The maximum flow rate should be approximately 200 l/min |
|  | The equipment should provide a fast rise time in Pressure Controlled mode without overshoot. |
|  | The equipment should have an easy accessible timer displaying hours, minutes and seconds. |
|  | The equipment should have tools for performing a lung recruitment maneuver |
|  | The lung recruitment manoeuvre should be automatic |
|  | It should be possible to measure and visualize dynamic compliance breath by breath |
|  | It shall be possible to pause ventilation and fresh gas flow for a defined time period." |
| 25. Should have single / Double absorbent chamber canister. Easily removable for changing. Supplied with all standard accessories including Bains Circuit. |  |
| 26. Closed circuit system, Sodalime canister, Single / Double Chamber |  |
| 27. Pediatric Circuit (Jackson Rees) |  |
| 28. Auxiliary O2 Supply point |  |
| 29. Mask of Different size (2 each- 1 to 5) |  |
| 30. Reservoir bag- 3 litre/5 litre/0.5 litre- 2 each |  |
| 31. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO. |  |

Department : ORTHOPAEDICS

Set for Hip Replacement

| 1. | BASIC INSTR. SET, HIP – 1 |
| --- | --- |
| 2. | FORCEPS, GROSS-MAIER, CVD., 26.5 CM – 2 |
| 3. | TOWEL FORCEPS, BACKHAUS, SHARP, 11 CM – 6 |
| 4. | TOWEL FORCEPS F. PAPER DRAPES, 11.5 CM – 6 |
| 5. | SCALPEL HANDLE, NO. 4, 13.5 CM-1 |
| 6. | SCALPEL HANDLE, NO. 4L, LONG, 20.5 CM-1 |
| 7. | TC-DISS. SCISSORS, FINE, CVD., 20.5 CM – 1 |
| 8. | TC-DISSECTING SCISSORS, CVD., 23 CM – 1 |
| 9. | TC-DISS. SCISSORS, CVD., SERR., 18 CM – 1 |
| 10. | TC-SCISSORS, MAYO-LEXER, CVD., 16 CM – 1 |

| 11. | OPERATING SCISSORS, SH/BL, STR., 14.5 CM – 1 |
| --- | --- |
| 12. | TISSUE FORCEPS, 1X2 T., 16 CM – 2 |
| 13. | TISSUE FORCEPS, 1X2 T., 20.5 CM – 2 |
| 14. | TISSUE FORCEPS, 1X2 T., SLIM, 20.5 CM – 1 |
| 15. | ATR. FORCEPS, DE BAKEY, 2.7 MM, 20 CM – 2 |
| 16. | ATR. FORCEPS, DE BAKEY, 2.7 MM, 25 CM – 1 |
| 17. | HAEM. FORCEPS, MOSQUITO, CVD., 12 CM – 4 |
| 18. | HAEMOSTATIC FORCEPS, PEAN, STR., 16.5 CM – 2 |
| 19. | FORCEPS, KOCHER, 1X2 T., STR., 16.5 CM – 4 |
| 20. | HAEM. FORCEPS, MOSQUITO, CVD., 18 CM – 2 |
| 21. | FORCEPS, KOCHER, 1X2 T., STR., 22.5 CM – 2 |
| 22. | FORCEPS, GROSS-MAIER, STR., 26.5 CM – 2 |
| 23. | DISS. FORCEPS, OVERHOLT, NO. 4, 22 CM - 1 |
| 24. | TC-NEEDLEHOLDER, MAYO-HEGAR, 20 CM – 2 |
| 25. | TC-NEEDLEHOLDER, HEGAR, 20 CM – 1 |
| 26. | RETRACTOR, DOUBLE, ROUX, NO. 3, 17 CM – 2 |
| 27. | RETRACTOR, KOCHER, 61X20 MM, 23 CM – 2 |
| 28. | RETRACTOR, MIKULICZ, 90X35 MM, 25 CM – 2 |
| 29. | RETRACTOR, KOCHER, SEMISH., 6-PR., 22 CM – 2 |
| 30. | RETRACTOR, KOERTE, SHARP, 8-PR., 24.5 CM – 2 |
| 31. | RETRACT., ISRAEL, BL., 50X60 MM, 24.5 CM – 2 |
| 32. | RETRACTOR, VOLKMANN, SH., 1-PR., 21.5 CM – 1 |
| 33. | BONE LEVER, HOHMANN, 18 MM, 23.5 CM – 2 |
| 34. | BONE LEVER, HOHMANN, 22 MM, 25.5 CM-2 |
| 35. | BONE LEVER, HOHMANN, 43 MM, 23.5 CM – 2 |
| 36. | BONE LEVER, HOHMANN, 70 MM, 25 CM – 1 |
| 37. | SCHUMACHER BONE RETRACTOR, 35X18MM, 29CM – 1 |
| 38. | BONE LEVER, HOHMANN, 43 MM, 23.5 CM – 1 |
| 39. | UTERINE SCOOP, SIMON, NO. 2, 23.5 CM – 1 |

| 40. | UTERINE SCOOP, SIMON, NO. 4, 24 CM – 1 |
| --- | --- |
| 41. | BONE CURETTE, 37x30MM, 32,5CM – 1 |
| 42. | BONE HOLDING FORCEPS, LANGENBECK, 21 CM - 1 |
| 43. | BONE RONGEUR, RUSKIN, CVD., 24 CM-1 |
| 44. | ELEVATOR, LANGENBECK, 10 MM, 19.5 CM - 1 |
| 45. | RASPATORY, LANGENBECK, CVD., 19 CM - 1 |
| 46. | RASPATORY, SEMB, NO. 5, 15 MM, 23 CM - 1 |
| 47. | BONE LEVER 24CM, 17 MM - 1 |
| 48. | REAMER, PERTHES, 21.5 CM - 1 |
| 49. | OSTEOTOME, LEXER, 20 MM, 22 CM - 1 |
| 50. | OSTEOTOME, LEXER, 30 MM, 22 CM - 1 |
| 51. | OSTEOTOME, LAMBOTTE, 30 MM, 24 CM - 1 |
| 52. | LEXER OSTEOTOME 10 MM ANGULAR, 27,5 CM-1 |
| 53. | LEXER OSTEOTOME 15 MM ANGULAR, 27,5 CM-1 |
| 54. | SWAN-NECK GOUGE, 16MM - 1 |
| 55. | SWAN-NECK GOUGE, 25MM - 1 |
| 56. | CASPAR TAMPERS 7 MM 20 CM - 1 |
| 57. | TAPPET 12MM DIAMETER, 20 CM - 1 |
| 58. | MALLET, RELPASE FREE, 620 GR., 26.5 CM - 1 |
| 59. | STILLE OSTEOTOME CURVED 20 CM 20 MM - 1 |
| 60. | LUXATIONSHEBEL Z. ABHEB. D.FEMURKOPF - 1 |
| 61. | CARTILAGE SCISSORS, RESANO, ANGLED, 25CM - 1 |
| 62. | CART. FORCEPS, BIRCH.-GA., CVD., 20.5 CM - 1 |
| 63. | IMPACTION INSTR. F. HIP PROSTHESES/HEADS - 1 |
| 64. | NEEDLE CASE, ROUND, PERF., F. 55-309-65 - 1 |
| 65. | BOWL, METAL, H = 55, Ø 128 MM, 0.35 L - 1 |
| 66. | KIDNEY DISH, 250X140X40 MM - 1 |
| 67. | GUIDE NEEDLE, ANG., KNIFE SHAPE, 10 CH - 1 |
| 68. | GUIDE NEEDLE, CVD., KNIFE SHAPE, 14 CH - 1 |

| 69. | GUIDE NEEDLE, CVD., KNIFE SHAPE, 16 CH - 1 |
| --- | --- |
| 70. | CONTAINER MS, 60X30X16 CM, HANDLE GREY - 1 |
| 71. | Tray DIN, 480x255x33 mm – 1 |
| 72. | Tray DIN, 480x255x73 mm – 1 |
| 73. | COLOR-TAG, RED – 2 |
| 74. | CODING LABEL, WITH TEXT, WITHOUT HOLE - 2 |
| 75. | Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO. |

Department : ORTHOPAEDICS

Set for Knee Replacement

1.   BASIC INSTRUMENT SET, KNEE REPLACEMENT

2.   FORCEPS, GROSS-MAIER, CVD., 26.5 CM - 2

3.   TOWEL FORCEPS, BACKHAUS, SHARP, 11 CM - 4

4. TOWEL FORCEPS F. PAPER DRAPES, 11.5 CM - 4

5. SCALPEL HANDLE, NO. 4, 13.5 CM - 2

6. SCALPEL HANDLE, NO. 3, 12 CM - 2

7. TC-DISS. SCISSORS, FINE, CVD., 18 CM - 1

8. TC-DISS. SCISSORS, FINE, CVD., 20.5 CM - 1

9. TC-DISS. SCISSORS, CVD., SERR., 18 CM - 1

10. TC-SCISSORS, MAYO-LEXER, CVD., 16 CM - 1

11. OPERATING SCISSORS, SH/BL, STR., 14.5 CM - 1

12. TISSUE FORCEPS, 1X2 T., 14.5 CM - 2

13. TISSUE FORCEPS, 1X2 T., 20.5 CM - 2

14. ATR. FORCEPS, DE BAKEY, 2 MM, 16 CM - 2

15. HAEM. FORCEPS, MOSQUITO, CVD., 12 CM - 6

16. FORCEPS, KOCHER, 1X2 T., STR., 16.5 CM - 6

17. DISS. FORCEPS, OVERHOLT, NO. 3, 21.5 CM - 1

18. TC-NEEDLEHOLDER, CRILE-WOOD, 15 CM - 1

19. TC-NEEDLEHOLDER, MAYO-HEGAR, 18.5 CM - 1

20. TC-NEEDLEHOLDER, HEGAR, 20 CM - 1

21. RETRACTOR, DOUBLE, ROUX, NO. 1, 14.5 CM - 2

22. RETRACTOR, LANGENBECK, 50X11 MM, 22 CM - 2

23. RETRACTOR, KOCHER, 61X25 MM, 23 CM - 2

24. RETRACTOR, VOLKMANN, SH., 1-PR., 21.5 CM - 1

25. RETR., VOLKMANN, SEMISH., 4-PR., 22.5 CM - 2

26. WOUNDSPREADER, SHARP, 3X4 T., 16 CM - 1

27. ELEVATOR, LANGENBECK, 7 MM, 19.5 CM - 1

28. WAGNER BONE ELEVATOR, BLUNT, 17CM - 2

29. LANGE-HOHMANN LEVER MOD, 280MM - 2

30. BONE CURETTE, VOLKM., OVAL, NO. 1, 17CM - 1

31. WELLER MENISCUS FORCEPS, CVD., 20CM - 1

32. BONE HOLDING FORCEPS, LANGENBECK, 21 CM - 1

33. BONE RONGEUR, RUSKIN, CVD., 19 CM - 1

34. BONE RONGEUR, RUSKIN, CVD., 24 CM - 1

35. BONE CUT. FORCEPS, RUSKIN, CVD., 18.5 CM - 1

36. RASPATORY, LAMBOTTE, 15 MM, 21.5 CM - 1

37. OSTEOTOME, LEXER, 10 MM, 22 CM - 1

38. OSTEOTOME, LEXER, 15 MM, 22 CM - 1

39. OSTEOTOME, LEXER, 25 MM, 22 CM - 1

40. MALLET, RELPASE FREE, 620 GR., 26.5 CM - 1

41. CASPAR TAMPERS 5 MM, 20 CM - 1

42. CASPAR TAMPERS 7 MM 20 CM - 1

43. FLAT NOSE PLIERS, TRANSM., WIDE, 18 CM - 1

44. X-RAY RULER, PLASTIC, FLEXIBLE, 50 CM - 1

45. PROBE, BUTTON END, Ø 2.0/2.0 MM, 18 CM - 1

46. NEEDLE CASE, ROUND, PERF., F. 55-309-65 - 1

47. BOWL, METAL, H = 40, Ø 80 MM, 0.14 L - 1

48. BOWL, METAL, H = 55, Ø 128 MM, 0.35 L - 1

49. KIDNEY DISH, 250X140X40 MM - 1

50. GUIDE NEEDLE, ANG., KNIFE SHAPE, 10 CH - 1

51. GUIDE NEEDLE, CVD., KNIFE SHAPE, 14 CH - 1

52. CONTAINER MS, 60X30X16 CM, HANDLE GREY - 1

53. Tray DIN, 480 x 255 x 33 mm – 1

54. Tray DIN, 480 x 255 x 73 mm – 1

55. COLOR-TAG, RED – 2

56. CODING LABEL, WITH TEXT, WITHOUT HOLE - 2

57. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department : ORTHOPAEDICS

Interlock nailing sets

1. Material Should be titanium / stainless steel (Instruments)2. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

3. Drill Bit, 3.2mm dia., L 225/200 mm for quick coupling – 2 Nos

4. Holding Sleeve, large, L 120 mm -1 Nos

5. Screwdriver, hexagonal, large, L 280 m - 1Nos

6. Pin Wrench, 4.5mm, L 120mm - 1 nos

7. 5Combination Wrench, 11/14mm, L 150mm - 1 no’s

8. Awl, small, L 210 mm 17 Tissue Protector, L 140 mm -1 nos

9. T-Handle with quick coupling, L 85mm -1nos

10. Removing Tool for Syn Ream -1 nos

11. Hand Reamer, 8.0mm dia., for predrilling in pseudarthroses --1 nos

12. Reaming Rod, 2.5mm dia., L 950mm, 3.5mm olive--1 nos

13. Flexible Shaft, 7.0mm dia., reaming depth to 470mm -1 nos

14. Cleaning Brush for 3.6mm Flexible Shaft, L 600mm -1 nos

15. Reduction Head, straight -1 nos

16. Reduction Head, displacement 2.5mm - 1 nos

17. Reamer Head, 8.5 mm dia. -1 nos

18. Reamer Head, 9.0 mm dia. -1 nos

19. Reamer Head, 9.5 mm dia. -1 nos

20. Reamer Head, 10.0 mm dia. -1 nos

21. 20Reamer Head, 10.5mm dia. -1 nos

22. Reamer Head, 11.0mm dia. -1 nos

23. Reamer Head, 11.5mm dia. -1 nos

24. Reamer Head, 12.0mm dia. -1 nos

25. Reamer Head, 12.5mm dia. - 1 nos

26. Reamer Head, 13.0mm dia. -1 nos

27. Reamer Head, 13.5mm dia. -1 nos

28. Guide Rod, 3.0mm dia., with flat tip -1 nos

29. Socket Wrench 11mm, cannulated, L 180mm - 1 nos

30. Driving Piece, curved, L 120mm -1 nos

31. Driving Head – 1 nos

32. Guide Rod, cannulated, L 455 mm - 1 nos

33. Ram - 1nos

34. Grip, flexible, L 170mm -1nos

35. Insertion Handle, for Tibial Nails 9.0 to 14.0mm dia.-1nos

36. Threaded Bolt, conical, for Tibial Nails 9.0 to 14.0mm dia.-1nos

37. Knurled Nut for Tibial Nails 9.0 to 14.0mm dia. - 1 Nos

38. Insertion Handle, for Femoral Nails 9.0 to 12.0mm dia. – 1 Nos

39. Threaded Bolt, conical, for Femoral Nails 9.0 to 12.0mm dia.-1Nos

40. Knurled Nut for Femoral Nails, 9.0 to 12.0mm dia -1nos

41. Protection Sleeve 11.0/8.0, L 96mm - 1 nos

42. Drill Sleeve 8.0/4.5 - 1nos

43. Insert Drill Sleeve 3.2 – 1 nos

44. Trocar 8.0 mm dia., L 110 mm – 1 Nos

45. Depth Gauge for Locking Bolts -1nos

46. Drill Bit, 4.0/4.5mm dia., L 225/200mm, for quick coupling - 2nos

47. Holding Forceps for Reaming Rod 2.5 mm - 1 nos

Department : ORTHOPAEDICS

Fracture Reduction OT Table

| 1. Operating tables provide an elevated surface that supports the patient’s body during surgical procedures, stabilizing the patient’s position and providing optimal exposure of the surgical field. C-arm compatibility with electro-hydraulic operation table. |  |  |  |  |
| --- | --- | --- | --- | --- |
| Essential technical specifications: |  |  |  |  |
| 2. Minimum six section table-top, which should be X-Ray translucent for fluoroscopy with 'c' arm and with radiolucent mattress. |  |  |  |  |
| 3. The table should have a provision to be operated in the following modes: |  |  |  |  |
| a. Electrically by remote control and should have a removable standby handset with rollable cord located on the table column as well as an additional manual auxiliary mode facility. |  |  |  |  |
|  | b. It should have integrated batteries with a capacity for approx. |  | 80 operations and an |  |
|  | integrated battery charger. |  |  |  |
| 4. It should have at least the following electrical functions: |  |  |  |  |
| (a) High / low approx. 600 / 1020 – 1050 mm |  |  |  |  |
| (b) Side tilt / lateral left and right approx. 15-30° |  |  |  |  |
| (c) rendelenburg and Reverse Trendelenburg approx. 20-25/30-35° |  |  |  |  |
| (d) Back section approx. +70° / -40° |  |  |  |  |
| (e) Leg section approx. +70° / -90° |  |  |  |  |
| 5. | The leg section must be removable and the leg section interface attachment of electrically adjustable leg holders or an optional section for shoulder arthroscopy |  | must allow the alternative electrically adjustable Back |  |
| 6. | Manual adjustments (hydraulic / gas spring supported). It should functions: |  | have at | least the following |
| (a) Head section up and down approx. +25° / -45° |  |  |  |  |
| (b) Divided leg section up & down approx. +20° / -90° |  |  |  |  |
| (c) Swivelling of leg sections approx. 70° |  |  |  |  |
| (d) Back section up & down approx. +70° / -40° |  |  |  |  |
| (e) High / low approx. 600-1050mm |  |  |  |  |
| (f) Side tilt / lateral left and right approx. 15-30° |  |  |  |  |
|  | (g) Trendelenburg and Reverse Trendelenburg approx. 25/35°. The |  |  | table can be |
| operated manually, electrically on AC mains power or on batteries. |  |  |  |  |
| 7. | The remote control should have a clearly labelled control panel for main adjustments such as height, lateral, Trendelenburg / Reverse Trendelenburg, back section, leg section. |  |  |  |
| 8. | It must have an indication of the load control of the batteries. |  |  |  |
| 9. | The control unit should have a backlit display and control panel ensuring a safe and easy handling in the darkened operating room when doing minimally invasive surgery. |  |  |  |
| 10. Optional provision of Bluetooth / Infrared remote control unit . |  |  |  |  |
| 11. Should preferably have Integrated software information system to simplify maintenance and service of the table |  |  |  |  |
| 12. Radiolucent six section table top in head section, short back section, back section extension, seat section with perineal cut, divided leg section. With quick release facility to remove or interchange head and leg sections (Reverse Mode). |  |  |  |  |
| 13. Should preferably have antistatic and liquid-tight mattresses with shock absorbing foam for Decubitus Prophylaxis and for reducing the cooling of the patient during surgery. The mattresses must be free from Latex material. Mattress must preferably have a thickness 50-80mm. |  |  |  |  |
| 14. All metal components of the table should be made of corrosion resistant and disinfectant-proof stainless steel. |  |  |  |  |
| 15. Mobile table-base on antistatic heavy-duty swivel castors with central brake ensuring a stable base. |  |  |  |  |
| 16. Should also be suitable for obese patients. Maximum tolerable weight should be 450 kgs. |  |  |  |  |
| 17. Should have stainless steel accessory rails on both sides of the table-top to hold various accessories. |  |  |  |  |
| 18. Should be supplied with the accessories for orthopaedic traumatology, spinal surgery and shoulder and knee arthroscopy. Orthopaedic accessories: |  |  |  |  |
| a. Extension Device with following Pair of Adapters Counter traction Post for femur Telescopic bar long Telescopic bar short Screw tension device Foot Plate support Side rail Ext Supporting bars Radial Setting clamp Transport Cart |  |  |  |  |
|  | b. | Foot Plate (Pair) 3.Counter traction post for femur |  |  |
|  | c. | Rotation & tilting clamp |  |  |
|  | d. | Traction Stirrup clamp |  |  |
|  | e. | Countertraction Post tibia |  |  |
|  | f. | Condyle Fixation |  |  |
|  | g. | Universal support for positioning lower Leg |  |  |
|  | h. | Pad for disc operations |  |  |
|  | i. | Fixture for body support |  |  |
|  | j. | Lateral Support |  |  |
|  | k. | Shoulder support |  |  |
| l. Back buttocks support |  |  |  |  |
| m. PuBIS / ISO-sacrum-sternum support |  |  |  |  |
| n. Allen Arm/Hand table (Radiolucent) |  |  |  |  |
| 19. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO. |  |  |  |  |

Department : ORTHOPAEDICS

Simple OT tables

| 1. | Motor driven with remote / Electrohydralic table |
| --- | --- |
| 2. | 3 - part carbon top with the following functions |
| 3. | Height: 600 - 1050 mm ± 100 mm |
| 4. | Side tilt: +15 - 20 degrees |
| 5. | Back section adjustment: - 15 degrees to 70 degrees |
| 6. | Foot section adjustment: - 90 to 0 degree, detachable |
| 7. | Trendelenburg: 20 – 25 degree |
| 8. | Anti -Trendelenburg: 30 – 35 degree |
| 9. | Head section adjustment: - 40 to - 30 degree, detachable |
| 10. Maximum width: 540 mm without side rails |  |
| 11. Length: 1950 mm |  |
| 12. Standard accessories: |  |
| 13. Padded arm rest with straps - pair with damps |  |
| 14. Anesthesia screen with clamps |  |
| 15. Side supports: pair with clamps |  |
| 16. Shoulder supports: pair with clamps |  |
| 17. Knee crutches: pair with damps |  |
| 18. X- ray cassette tray |  |
| 19. Kidney bridge position to be achieved by Flex movement of the table |  |
| 20. Infusion rod with clamp |  |
| 21. Adapter for orthopaedic lower limb attachments |  |
| 22. Telescopic extension bars |  |
| 23. Turn over screw tensioner |  |
| 24. Adjustable foot support with rotation |  |
| 25. Traction: foot apparatus Adult ; Paediatric |  |
| 26. Pelvic support with central , right and left post positions |  |
| 27. Carbon top hand operating table: 80- 90cm x 35-40cm with support rod and side rail |  |
| 28. Stirrup clamp |  |
| 29. Knee support: radiolucent and padded; height adjustable with telescopic support post attachable to both sides of the table |  |
| 30. Knee support: radiolucent and padded; height adjustable with telescopic support post attachable to extension bar |  |
| 31. Femoral counter traction post including pads should be radiolucent |  |
| 32. Calf support with clamp and pad |  |
| 33. Accessories stand : mobile with frames and baskets of stainless steel 3 level for storing small parts and accessories |  |
| 34. Cushions and soft frame for spinal frames |  |
| 35. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO. |  |

Department : ORTHOPAEDICS

Pneumatic Drill set reamer Set

| 1. Cannulated Pneumatic Drill handpiece -Compatible with existing attachments |  |  |  |
| --- | --- | --- | --- |
| 2. Cannulation with 3.2 mm diameter Operating pressure : 6 -7 bars (maximum 10 bars)Weight of 5 handpiece 600 – 800 grams without any attachments Power 120 w Variable Speed from 0-900 rpm |  |  |  |
| 3. Noise Level of max 75 db Separate forward and reverse triggers Safety Device to cut off air supply to drill on handpiece is compatible with radiolucent drive Instant change between clockwise and counterclockwise rotation Offers reliable protection of soft tissues with oscillating drill attachment Fully Autoclavable Fully machine washable |  |  |  |
| 4. All Attachments can be fitted on single handpiece |  |  |  |
| 5. The reverse trigger automatically locks when the oscillating saw and the reduction drive attachments are attached to handpiece Adapter for Lubrication For oiling of hand piece |  |  |  |
| 6. Supply, Installation and Commissioning of Equipment required in Orthopedics |  |  |  |
| 7. Autoclavable Should be made of Stainless Steel3Double Air Hose Length 5 meters. |  |  |  |
| 8. Autoclavable Should have concentric inlet and outlet pipes |  |  |  |
| 9. | Radiolucent Drive Precise aiming and drilling under image intramedullary nails |  | intensifier control for locking |
| 10. Drill Bit diameter 2.0 to 4.5 mm, Length 100 to 150 mm, Usable length 80 to 120 mm |  |  |  |
| 11. Reduced exposure to x-rays 5 Jacob's Chuck attachment huck capacity up to 0 to 6.5 mm Cannulation of 3.2 mm diameter |  |  |  |
| 12. Maximum Speed of 900 rpm Torque of 4-5 Nm6Quick Coupling attachment Cannulation 3.2 mm |  |  |  |
| 13. Maximum Speed: 900 rpm Torque of 4-5 Nm7Reduction Drive for Intramedullary / Acetabular |  |  |  |
| 14. Reaming with reverse option Reaming Speed of 300-350 rpm Reaming |  |  |  |
| 15. Torque of 12-14 Nm Option of attachment with reverse rotation |  |  |  |
| 16. | Quick Coupling for K-wire Continuous adjustment facility for | wire | diameter from 0.6 to 3.2 mm |
| 17. Speed up to 900 rpm 9 Oscillating Saw attachment with key |  |  |  |
| 18. It can operate on an oscillating frequency of 0 to 14,000 osc/min. |  |  |  |
| 19. | Attachment can be locked in 8 different positions 10 Quick | Coupling for drill bits |  |
| 20. | Speed: 0 – 900 rpm Torque: 0 – 4.7 Nm Cannulation: 1.3 triple reamers | mm11Quick coupling for DHS / DCS |  |
| 21. Oscillating Saw Blades (All Sizes)For Trauma For Joint replacement 613 Aluminum case for Pneumatic Drill system, Perforated, Autoclavable 114 Aluminium Box for accessory attachment, Perforated, Autoclavable1 Technical Specifications for Pneumatic Drill Set –1. |  |  |  |
| 22. Company should have relevant experience in successful execution of similar work at least in five Institutes of national importance and central government Institutes |  |  |  |
| 23. Company should be at least in its 5 years of operations at the date of Submission of tender. |  |  |  |
| 24. Physical demo may preferably be arranged at the time of requirement. |  |  |  |
| 25. Instruments quality should meet the international standard. |  |  |  |
| 26. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.. |  |  |  |
| 27. Company should provide material certificates. |  |  |  |
| 28. Principle Company may preferably have certificate for passivation process of instrument. |  |  |  |

Department : ORTHOPAEDICS

Electric Drill & Reamer Set

1 Description of Function

1.1 Drilling machines are used in several orthopedic surgical procedures, for example, in making holes in bones for bone screws and in drilling out the medulla or marrow areas of bones.

2 Operational Requirements

2.1 Electric driven, autoclavable, versatile, forward & reverse mode with oscillating saw handpieces

3 Technical Specifications

3.1 1. Heavy duty electrically operated console/controller

Input voltage: 220-240V/domestic power supply.

Frequency: 40 - 70 Hz

Fuses: 3.5 - 4.5mA / commonly used for such device.

With min 12 ft autoclavable quick coupled light weight cord.

Specification for Oscillating Saw Handpiece

Pistol type

Autoclavable

Speed: min. 0-10,000cpm

Wrench free blade placement

Variable speed control with trigger

Plane of Blade movement: 360 degree

• Specification of Oscillating blade:

Thickness : 0.6 mm -12 mm

Cutting Depth : 60 mm – 95 mm

Width : 10 mm – 25 mm

ACL Blades : 5 mm – 9 mm

4. Specifications for Handpiece for Drill & Reamer

Autoclavable

Pistol Grip & Toolless Assemble of attachment

Cannulated Quick Coupling Connection

Drill Speed Range: 0 – 750 rpm

Torque: 30 - 35 in IB

Speed Control: Through trigger

Forward Reverse Mode: Yes

Oscillating mode for drilling/reaming: 270 deg forward &Reverse.

Screw Mode

Tap Mode

Reaming Attachment: 0-250 rpm 100 in lb torque, Cannulated

5. Drill Reamer handpiece attachments

Jacob chuck for drill 5/32”

Jacob chuck for drill 1/4 “

Wire driver attachment

AO attachment for drilling

AO attachment for reaming

Zimmer / Hudson Attachment for reaming

Reciprocating Saw attachment

Radiolucent Drive adapter

System Configuration Accessories, spares and consumables As specified

5 Environmental factors

5.1 1. The unit shall be capable of being stored continuously in ambient temperature of 0-50 deg C and relative humidity of 15 - 90%.

2. The unit shall be capable of operating continuously in ambient temperature of 10-40 deg C and relative humidity of 15-90%

6 Power Supply

6.1 Power input to be 220-240 VAC, 50Hz fitted with Indian plug

7 Standards, Safety and Training

7.1 1. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

2. Manufacturer should have ISO certification for quality standards.

3. Comprehensive training for lab staff and support services till familiarity with the system on site.

4. Certified to be meeting Electrical safety standards for medical equipment

7.2 Should have local service facility. The service provider should have the necessary equipment recommended by the manufacturer to carry out preventive maintenance test as per guidelines provided in the service/maintenance manual.

8 Documentation

8.1 1. User/Technical/Maintenance manuals to be supplied in English.

2. Certificate of calibration and inspection.

3. List of Equipment available for providing calibration and routine Preventive Maintenance Support as per manufacturer service/ maintenance manual.

4. List of important spare parts and accessories with their part number and costing.

5. Log book with instructions for daily, weekly, monthly and quarterly maintenance checklist. The job description of hospital technician and company technician clearly spelt out.

6. Compliance Report to be submitted in a tabulated and point wise manner clearly mentioning the page/para number with authenticated catalogue/manual, without which it will not be considered.

Department : ORTHOPAEDICS

Arthroscope

1. 4mm 30 D autoclavable scopes

2. Sheath

3. 4.All Attachments can be fitted on single handpiece

4. 4 mm 70 D scope

5. Light Source LED

6. Camera Controller

7. CAMERA HEAD

8. HD LED Monitor

9. 10. Drill Bit diameter 2.0 to 4.5 mm, Length 100 to 150 mm, Usable length 80 to 120 mm

10. Autoclavable Light Guide

11. Irrigatiob Control

12. 15.Torque of 12-14 Nm Option of attachment with reverse rotation

13. SHAVER CONSOLE

14. SHAVER HANPIECE

15. FOOTHSWITH

16. RF GENERATOR

17. MOBILE TROLLEY

18. Hand instruments

19. ACL / PCL Set

20. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO..

Department ; OTORHINOLARYNGOLOGY

Suction Apparatus.

Suction Machine :

1. Vaccum /LPM : - 700 MM Hg , 50 Litres/Min

2. Pump Type- Double rotary vane type

3. Flutter free vacuum control knob,

4. Wide mouthed 2 x 2 LITRE (light weight, unbreakable and clear) with self-sealing bungs and mechanical over flow safety device.

5. Noise (in dBA)- 50 dB

6. Collection container & its cap, suctions tube tips, a vacuum gauge, two sets of moisture & microbial filters and control knob

7. SiliconeTubing:8 mm ID x 2 mtr (PVC), 2x2 lt jar (one set

8. It should be Mobility, portability.

9. Equipment shall operate on 220V-240V, 50 Hz, single phase electric supply.

10. ISO/ BIS approve model should be offered.

Department : OTORHINOLARYNGOLOGY

Otoscope

1. 2.5 V (White Led Light) Pocket Scope

2. Auriscope With Aa Battery Handle

3. Wide - Angle Viewing Lens Allows For Instrumentation Under Magnification.

4. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; OTORHINOLARYNGOLOGY

Rigid Nasal Endoscope 0 Degree.

FESS Set :

Rigid nasal endoscope 0 degree, 4 mm and 2.7 mm size - 01

Department ; OTORHINOLARYNGOLOGY

Rigid Nasal Endoscope 30 Degree.

FESS Set :

Rigid nasal endoscope 30 degree, 4 mm and 2.7 mm size -01

Department ; OTORHINOLARYNGOLOGY

Light Source and Light Cable.

Light source and light cable -01

Department ; OTORHINOLARYNGOLOGY

Endoscopic Camera with suitable display with recording & archiving facility. (with High Definition Camera) .

| Endoscopic Camera with suitable display with recording & archiving facility. (with High Definition Camera) : |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Otoendoscope, Nasalendoscope, Laryngeal telescope with camera, monitor and light source for teaching purposes (1 Unit of the above may be placed either in Minor OT or Endoscopy room or in the OPD in the Consultant Chamber) Oto - Endoscopes |  |  |  |  |  |
| 1. Straight forward Telescope 0 ° 3 mm 14cm length |  |  |  |  |  |
| 2. Straight forward Telescope 30 ° 3 mm 14cm length-1 |  |  |  |  |  |
| 3. Straight forward Telescope 0 ° 4 mm 6 cm length-1 |  |  |  |  |  |
| 4. Straight forward Telescope 30 ° 4 mm 6 cm length-1 |  |  |  |  |  |
| 5. Pneumatic adapter for Oto-Endoscope With all the attachments-1 |  |  |  |  |  |
| NASAL ENDOSCOPES |  |  |  |  |  |
|  Nasal Rigid Endoscope. – One |  |  |  |  |  |
|  Straight forward telescope, 0 degree enlarged view, size: 4.0 MM |  |  |  |  |  |
|  |  |  |  | Rod lenses system, Length:18-19 cms, Autoclavable, Fiber optic light transmission incorporated. |  |  |
|  |  |  | Nasal Rigid Endoscope. – One |  |  |  |
|  |  |  |  | Forward oblique 30 degree enlarged view, size: 4.0 MM rod lenses system, Length:18- 19 cms, Auto clavable, Fiber optic light transmission incorporated. |  |  |
|  |  |  | Nasal Rigid Endoscope. – One |  |  |  |
|  |  |  |  | Wide Angle Forward- Oblique Telescope 45°, enlarged view, diameter 4 mm rod lenses system, length 18 - 19 cm, autoclavable. Fiber optic light transmission incorporated. |  |  |
|  | LARYNGEAL ENDOSCOPE |  |  |  |  |  |
|  |  |  |  | Strabo-Laryngoscope with integrated lateral telescope 70 deg, oval sheath, 7.2 x 9.3mm diameter, length 17 cm, autoclavable, fiber optic light transmission incorporated. |  |  |
|  | TECHNICAL SPECIFICATION OF DESKTOP COMPOSITE DOCUMENTATION TERMINAL WITH INBUILT CAMERA, MONITOR, LIGHT SOURCE AND STROBOSCOPE |  |  |  |  |  |
|  |  |  |  | The system should be compact and portable suitable for a variety of endoscopic applications from physician’s offices to operating rooms in a variety of specialties. |  |  |
|  |  |  |  | The powerful all-in-one unit should consist of everything needed for endoscopic imaging, Video recording, and viewing of saved Videos, the monitor, camera, and light source. |  |  |
|  |  |  |  | IT should have USB ports and a SD card slot for documentation purpose. |  |  |
|  |  |  |  | It should also have an inbuilt Stroboscope for larynx examination a modern device that generates light with a high-performance LED |  |  |
|  |  |  |  | For stroboscopic examination (stroboscopemode) |  |  |
|  |  |  |  | For normal viewing (continuous light-mode = pulsating light with high frequency) |  |  |
|  | Should also be Suitable for larynx examination, |  |  |  |  |  |
|  | Consisting of: |  |  |  |  |  |
|  |  |  |  | Mains Cord |  |  |
|  |  |  |  | Microphone set |  |  |
|  |  |  |  | One USB pedal footswitch with integrated activation for Stroboscope function |  |  |
|  |  |  |  | DISPLAY: Crystal clear display |  |  |
|  |  |  |  | 15" LCD display • LED backlight display technology for extended service life, enhanced image brightness and reduced Image rotation power consumption |  |  |
|  |  |  |  | DVI video output for brilliant transmission quality |  |  |
|  |  |  |  | LED Light Source: • Highperformance LED light source: Light output similar to Power LED |  |  |
|  |  |  |  | Color temperature of 6000 K - similar to daylight - guarantees color fidelity |  |  |
|  |  |  |  | Long lamp life - with an average lamp life of 50,000 hours - Cost Effective |  |  |
|  |  | Easy, extremely reliable control: |  |  |
|  |  | Membrane keyboard included, suitable for wipedown disinfection |  |  |
|  |  | Hot keys for rapid and direct manipulation |  |  |
|  |  | Arrow buttons for intuitive control |  |  |
|  |  | Connection socket for pedal control without lag time |  |  |
|  |  | Stroboscopy mode can be activated via a special footswitch |  |  |
|  |  | Should have flexible storage possibilities |  |  |
|  |  | SD card slot for high storage capacity |  |  |
|  |  | USB ports for external hard drives, USB flash drives |  |  |
|  |  | MPEG4 video recording with sound via microphone input |  |  |
|  |  | Compatible with medical grade USB Color Printer |  |  |
|  |  | Picture gallery for records |  |  |
|  |  | Playback of saved videos |  |  |
|  |  | Patient information input and reports |  |  |
| Technical Specification of Camera Head:  Image sensor: 1/4’’ CCDChip. |  |  |  |  |
|  |  | Resolution:> 450 lines (horizontal). |  |  |
|  |  | Pixels - 752(H) x 582(V) |  |  |
|  |  | Signal-to-noise ratio: >= 60dB. |  |  |
|  |  | AGC: Microprocessor controlled |  |  |
|  |  | LENSE: Integrated optical zoom lens system 25 - 50 mm |  |  |
|  |  | Min. sensitivity:3 Lux (f 1.4). |  |  |
|  |  | Power input:100 W Power supply |  |  |
|  |  | Power supply:100- 240 VAC |  |  |
|  |  | Weight: Approx 7 kg |  |  |
|  |  | Interfaces |  |  |
|  |  | a Video interface |  |  |
|  |  | b DVI-D (in/out) |  |  |
|  |  | c- Audio: 3.5 mm jack |  |  |
| Interfaces  d-(rear), Line-in, Line-out |  |  |  |  |
|  |  | e-Footswitch port: 5-pin |  |  |
|  |  | f-socket for footswitch |  |  |
|  |  | g-LEMO socket (side) |  |  |
|  |  | e-Printer port: USB |  |  |
|  |  | Light Source |  |  |
|  |  | Lamp: LED |  |  |
|  |  | Color temperature: Light source 6400 K |  |  |

| - Average life span:  approx. 30,000 hours Imageformat |  |
| --- | --- |
|  | Image Format:JPEG |
|  | Video codec:MPEG-4 |
|  | Video format:PAL/NTSC |
|  | USB 2.0; Storage interface |
|  | Storage interface: SD memory card |
|  | (SDHC-compatible) |
| - Screen size: 15" LED backlight  Resolution: monitor 1024 x 768 |  |
|  | contrast:1:700 |
|  | Loudspeaker:Power: 2 W |
|  | Scope of supply should include of : |
|  | Complete Integrated Desktop Documentation Terminal with accessories |
|  | Stroboscope Kit |
|  | One USB of minimum 8GB capacity |
|  | Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO. |

Department ; OTORHINOLARYNGOLOGY

Bronchoscopes different sizes for different age groups.

| Bronchoscopes different sizes for different age groups : |  |  |
| --- | --- | --- |
| Rigid Bronchoscope |  |  |
| Digital Single Chip Endo vision Camera: one (compatible with telescopes of Bronchoscope, Oesophagoscope , and Cystoscope |  |  |
| Color system PAL / NTSC, power supply: 100 – 240 VAC, 50/60 Hz compatible with Pal, NTSC Automatic white balance with control on base unit and also on camera. |  |  |
| Integrated zoom lens for manual and automatic control for exposure of fog. |  |  |
| Compatible with VHS and Comp and DVI. |  |  |
| Minimum sensitivity 3 lux. Instrument coupling for all rigid endoscope. |  |  |
| Long camera cable 300 cm. |  |  |
| Two pre set function keys on camera for control of camera function , printer, computer, VCR and other peripherals Medical Color Monitor 19” (fully flat screen): one |  |  |
| Special 19 “monitor, medical grade, Trintron colour tube, PAL , NTSC & SECAM color system. For Bronchoscopy, Oesophagoscopy and cystoscopy |  |  |
| Telescope |  |  |
| Straight forward telescope , 0 degree , diameter 2.9 mm , length 36 cm , autoclavable, fiber optic light transmission incorporated( one) |  |  |
| Straight forward telescope 0 degree , diameter 2.9 mm , length 30 cm , autoclavable, fiber optic light transmission incorporated ( one) |  |  |
| Straight forward telescope 0 degree , diameter 2.7mm , length 18cm autoclavable, fiber optic light transmission incorporated ( one) |  |  |
|  | Bronchoscope Sheath |  |
|  | Bronchoscope sheath , size 4.5 outer diameter 7.3 mm, inner diameter 6.6 mm , length :30 cm ( one |  |
|  | Bronchoscope sheath , size 4 outer diameter 6.7mm, inner diameter 6mm , length :30 cm ( one ) |  |
|  | Bronchoscope sheath , size 3.7 outer diameter 6.4 mm , inner diameter 5.7 mm , length :30 cm ( one ) |  |
|  | Bronchoscope sheath , size 3.5 outer diameter 6.4 mm, inner diameter 5.7 mm , length :30 cm ( one ) |  |
|  | Bronchoscope sheath , size 4.0 outer diameter 6.7 mm, inner diameter 6.4 mm , length :26 cm ( one ) |  |
|  | Bronchoscope sheath, size 3.7outer diameter 6.4 mm ,inner diameter 6.4 mm , length :26 cm ( one ) |  |
|  | Bronchoscope sheath , size 3.5 outer diameter 5.7 mm, inner diameter 5.0 mm , length :26 cm ( one ) |  |
|  | Bronchoscope sheath , size 3.0 outer diameter 5.0 mm, inner diameter 4.3 mm , length :26 cm ( one ) |  |
|  | Bronchoscope sheath , size 3.5 outer diameter 5.7 mm, inner diameter 5.0 mm, length :18.5cm( one ) |  |
|  | Bronchoscope sheath , size 3.0outer diameter 5.0 mm, inner diameter 4.3 mm, length :18.5m ( one ) |  |
|  | Bronchoscope sheath , size 2.5 outer diameter 4.2 mm, inner diameter 3.5 mm, length :18.5m ( one ) |  |
|  | Optical forceps (compatible with above mentioned telescopes and bronchoscope sheaths |  |
|  | Optical alligator forceps, 2x2 teeth , with spring action handle , for controlled removal of flat foreign bodies ( such as coins) ;(two) |  |
|  | Optical forceps, with spring action handle for controlled removal of soft foreign bodies ( such as peanuts) : ( two) |  |
| Optical alligator forceps, with forced controlled handle, for removal of hard foreign bodies ( two ) |  |
| Forceps (one each) |  |
| Forceps alligator, grasping, double action jaw, sheath diameter 1.5 mm , working length: 35 cm |  |
| Forceps alligator, grasping, double action jaw, sheath diameter 1.5 mm , working length: 35 cm , pointed serrated for coins and foreign bodies. |  |
| Forceps alligator, grasping double action jaw, sheath diameter 1.5 mm , working length 35 cm , pointed for peanuts and soft foreign bodies. |  |
| Biopsy Forceps 35 Cm |  |
| Sponge holder: |  |
| Spring handle, working length 35 cm ( one) |  |
| Rigid suction tube with rubber tip ( one each ) |  |
| Rigid suction tube, straight length 35 cm, diameter: 3 mm |  |
| Rigid suction tube, straight length 25 cm. diameter: 3 mm |  |
| Foreign body basket with rigid handle, working length 35 cm |  |
| Telescope bridge for fixed position between telescope and bronchoscope. |  |
| Compatible with above mentioned telescopes and bronchoscope sheaths ( Three) |  |
| Prismatic light deflection with connection for fiber optic light cable, autoclave (Two) |  |
| Glass Window Plug (Two) |  |
| Rubber Telescope Guide for use with telescopes or optical forceps (Two) |  |
| FLUVOG Adaptor with sliding glass window plug, sealing cap, notched lens and keyhole opening , moveable (Two) |  |
| Injection canula with Leur lock outer diameter – 3.5 mm, for use with bronchoscope tubes (for positive pressure assisted ventilation (Two) |  |
| Injection canula with Leur lock outer diameter – 2.7 mm, for use with bronchoscope tubes (for positive pressure assisted ventilation (Two) |  |
| Instrument guide for suction catheter (Two) |  |
| Adaptor from bronchoscope to any type of pediatric respiration equipment (Two) |  |
| Sealing plug for ventilation attachment of bronchoscopes (Two) |  |
| Adjustable magnifier, swing – away type, autoclave (One) |  |
| Fiber optic light cable size 3.5mm, length 180 cm (Two) |  |
| Cold light fountain twin bulb Halogen 150 watts Power supply 220 VAC, 50 Hz. (One) |  |

Facility to capture images and record videos of the procedure should be available. Storage facility should be adequate for at least 100 videos.

Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; OTORHINOLARYNGOLOGY

Nasal Endoscopy Trolley with 0 degree & 30 degree 4mm Endoscope with

Light Source, Cable, Monitor & Camera.

Nasal endoscopy trolley :

Nasal endoscopy trolley with 0 degree & 30 degree 4mm endoscope

with light source, cable , monitor & camera

Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; OTORHINOLARYNGOLOGY

Basic OT Table

| Basic OT Table : |  |
| --- | --- |
| 1. Electro- Hydraulic operation table |  |
| 2. heavy weight of the patient. Complete with all accessories. |  |
| 3. Electro-hydraulically powered for movements of height, tilting and Trendelenburg. |  |
| 4. The base of stable construction with large castors for easy travel and maneuvering |  |
| 5. Table top 500 mm x 1940 mm, Height range 600- 950 mm |  |
| 6. Trend/Antitrend-20-25, latel tilt- 15-20 degree |  |
| 7. The Operation table provided with corded hand control and override – control panel for additional safety. |  |
| 8. The tabletop divided in 5 parts i.e., head plate (removable), back plate, seat plate and leg plates (removable). |  |
| 9. Guided rails for X-Ray cassettes inserted from the head end. |  |
| 10. Reverse positioning of patients on a special shoulder plate (three section) that can be connected at the foot end of the tabletop. |  |
| 11. Back plate is motorised for easy movements |  |
| 12. The Operation table should have safe positioning and adjustment due to optimal covering of moving joints. |  |
| 13. The Operation table should provide optimal patient positioning in a supine position. The longitudinal shift feature of 290mm should give unobstructed imaging of vertebrae. |  |
| 14. Mobile Operating Table, Electrohydraulic, SFC padding |  |
| 15. Standard Accessories: |  |
|  | Head Rest, Leg plates, pair, hinged and abductable, SFC padding |
|  | Arm board, Radial Setting Clamp, Geopel knee crutch, Body Strap |
|  | Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO approved. |

Department ; OTORHINOLARYNGOLOGY

Puretone Audiometer.

| Puretone Audiometer : 1. Light weight with suitable carrying case to accommodate the main unit and accessories. |  |
| --- | --- |
| 2. Channels:2 separate channels with independent attenuators. |  |
| 3. Test frequencies 250Hz to 8000Hz |  |
| 4. Hearing level range : |  |
| 5. air-conduction 10db to 120 db |  |
| 6. bone -10db to 80db(apprx) |  |
| 7. masking 10db to 90 db (apprx) narrow band, wide band |  |
| 8. Attenuator: 2.5db and 5 dB steps |  |
| 9. Stimulous:continousn pulsed or warable tone |  |
| 10. Total Hormonic distortion: |  |
| 11. air: less than 2.5 % / bone: less than 5% |  |
| 12. Power supply : suitable mains adapter: 230 VAC |  |
| 13. Standard accessories |  |
|  Microphone –Headphone |  |
|  | Patients' response switch required by the system |
|  | Power cord-Bone vibrator |
|  | Carrying case to accommodate the main unit & accessories |
|  | Headphone TDH 39 and TDH 49 |
| 14. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO. 15. sound Proof booth/ room |  |

Department ; OTORHINOLARYNGOLOGY

Brainstem Evoked Response Audiometer with ASSR.

| Brainstem Evoked Response Audiometer with ASSR : |  |
| --- | --- |
| 1. Should have ABR, MLR, LLR, ASSR |  |
| 2. Transducers: insert phone –Ear Tone 3 A, Head phone, Bone oscillator |  |
| 3. Should have minimum 2 channels. |  |
| 4. Polarity: Rarefaction, condensation, alternating |  |
| 5. Should have Intensity: 0-99 nHL or above |  |
| 6. Stimulus: Clicks, Tone burst |  |
| 7. Should have user friendly repletion rate |  |
| 8. Should have Notch filter 50 or 60 Hz. |  |
| 9. Should have lpsi and contra noise masking |  |
| 10. Should be Portable & built in isolation and shielding and test should be carry out in NICU or any locations in hospital |  |
| 11. ASSR: up to 4 stimulus frequency per ear, intensity 0 to 125 dp SPL, frequency range 250 Hz to 8 KHz, users defined modulation, |  |
| 12. Simultaneous testing of both ears |  |
| 13. Upgradable to: Diagnostic DPOA, TEOAE – Adjustable pass criteria, frequency dominant spectrum vies |  |
| 14. Ear probe tips – 2 packets (assorted) |  |
| 15. Reusable cup electrodes – 6 nos |  |
| 16. T 20 conduction gel – 3 nos |  |
| 17. Nuprep gel 114 gm - 6 nos. |  |
| 18. Desk Top Computer |  |
| 19. Processor: Intel core i3 |  |
|  | RAM : 4 GB |
|  | Hard Disk: 1TB |
|  | CD/ DVD |
|  | Monitor : LED/LCD 18.5 inch or above |
|  | Laser Printer |
| 20. Should be provided with table/trolley (Wooden/Ply/ MDF) to keep BERA machine, computer, UPS and printer |  |
| 21. Only UPS of suitable capacity with 30 minutes back up for BERA, Desktop and Printer shall be provide |  |
| 22. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO. |  |

Department ; OTORHINOLARYNGOLOGY

OAE Impedance Audiometer (With Sound Treated Air-Conditioned Room for Audiometry)

| OAE Impedance Audiometer (With Sound Treated Air-Conditioned Room for Audiometry) : |  |
| --- | --- |
| 1. Impedance Audiometer / Tympanometry |  |
| 2. Audiometer impedance with contra ear testing facilities |  |
| 3. Multifrequency |  |
| 4. Probe Frequency- 226Hz, 678Hz,800Hz,1000Hz |  |
| 5. Pressure Range- +200 to – 400 dapa |  |
| 6. Volume Range - 0.1 ml to 6.0 ml |  |
| 7. Accuracy - ±5% to ±10 dapa |  |
| 8. Test Time- < 3 Seconds |  |
| 9. Reflex Mode |  |
| 10. Test Frequencies- 500, 1000, 2000, 4000 Hz ± 2% |  |
| 11. Test Method- IPSI Lateral, Contralateral |  |
| 12. Noise (Band) - WN/HP/LP |  |
| 13. Intensities IPSI Lateral-70 to 110 dbHz |  |
| 14. Intensities Contra Lateral- 70 to 120 dbHz (with TDH39 ) |  |
| 15. Intensity Setting- Automatic or Manual |  |
| 16. Eustachian Tube Function - Intact and Perforated mode |  |
| 17. ETF Pressure Range-+ 300 to – 400 dapa |  |
| 18. Test -IPSI Lateral Reflex Test with AGC, Reflex Delay |  |
| 19. Test Programme- Reflex Test selectable |  |
| 20. Memory- Test Result of both ears |  |
| 21. Probe - Light weight, adjustable, Hand Held , With Built in control light & switch |  |
| 22. Printer- Silent Thermal Printer , (with paper printer facility) |  |
| 23. Display-Graphic LCD with adjustable contrast |  |
| 24. Power Supply- Mains 100-240 Volts, 50/60 Hz 25 VA |  |
| 25. PC Interface- USB Cable |  |
| 26. Automatic self calibration |  |
| 27. Oto Acoustic Emission (Screening unit) OAE (DP and TE) |  |
| 28. TEOAE |  |
| 29. 1.5 to 4 kHz |  |
| 30. Sample Rate - 16 kHz |  |
| 31. Stimulus Level- ca. 80 dB SPL peak Stimulus Type- Nonlinear click |  |
| 32. Statistical stop criterion (TE Quick) or user defined stop criterion (SNR: 3, 6 or 9 dB) in 3, 4, or 5 out of frequency bands (1, 1.5, 2, 3, 4 kHz) (TE Diag) Window of analysis- 5-13 ms post stimulus |  |
| 33. DPOAE |  |
| 34. DP 2 to 5 kHz Sample Rate - 24kHz |  |
| 35. Frequency Ratio f2/f1- 1.2 |  |
| 36. Level Ratio L12/L1- Scissor Paradigm Measurement Interval- 512 samples |  |
| 37. Frequencies f2- 1.5, 2, 3, 4, 6, 8, kHz (single & multiple selections possible) Stimulus Levels L2- 35 to 65 dB HL (in steps of 5dB) |  |
| 38. Also battery operated Multiple test methods |  |
| 39. Database for at least 1000 tests |  |
| 40. Data transfer to PC via USB or wireless Printing via PC/ Printer |  |
| 41. Stimulus intensity: 40 to 70 dB SPL (DPOAE). 83 dB |  |
| 42. SPL (TEOAE). |  |
| 43. Maximum output (Protection): 90 dB SPL. Microphone system noise: -20 dB SPL @ 2 kHz (1 Hz bandwidth). |  |
| 44. -13 dB SPL @ 1 kHz (1 Hz bandwidth). |  |
| 45. Power supply: (4) AA/UM-3/R6 - alkaline (6V total) Battery life: Approximately 300 tests. |  |
| 46. Display: LCD-display 4 line x 10 character. |  |
| 47. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO. |  |

Department ; OTORHINOLARYNGOLOGY

Flexible Nasopharyngolaryngoscope

Flexible Nasopharyngolaryngoscope :

1. Should be able to view nasopharynx, larynx, trachea and hypopharynx

2. Should be able to take biopsy with flexible forceps

3. Should have suction and irrigation facility

4. Should have own LED/Halogen light

5. Power Supply 220VAC +/- 10 %, 50Hz fitted with Indian plug.

6. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; OTORHINOLARYNGOLOGY

Electronystagmography.

Video Nystagmography :

1. High quality ISO certified with sensitivity of 105 images per second binocular, 174 images/sec monocular

2. Goggle with one camera and goggle with 2 cameras (non-occluded and occluded view)

3. Able to perform all vestibular test including smooth pursuit test (tracking)

4. Compatible with latest window software

5. Laptop with minimum 1.8GHz, 2GB DDR3 RAM, 160GB hard disc with resolution of 1024x768 resolution or better

6. Rotatory chair, irrigator for water and air included

7. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; OTORHINOLARYNGOLOGY

Brainstem Evoked Response Audiometer.

| Brainstem Evoked Response Audiometer : |  |
| --- | --- |
| 1. Should have ABR, MLR, LLR, ASSR |  |
| 2. Transducers: insert phone –Ear Tone 3 A, Head phone, Bone oscillator |  |
| 3. Should have minimum 2 channels. |  |
| 4. Polarity: Rarefaction, condensation, alternating |  |
| 5. Should have Intensity: 0-99 nHL or above |  |
| 6. Stimulus: Clicks, Tone burst |  |
| 7. Should have user friendly repletion rate |  |
| 8. Should have Notch filter 50 or 60 Hz. |  |
| 9. Should have lpsi and contra noise masking |  |
| 10. Should be Portable & built in isolation and shielding and test should be carry out in NICU or any locations in hospital |  |
| 11. ASSR: up to 4 stimulus frequency per ear, intensity 0 to 125 dp SPL, frequency range 250 Hz to 8 KHz, users defined modulation, |  |
| 12. Simultaneous testing of both ears |  |
| 13. Upgradable to: Diagnostic DPOA, TEOAE – Adjustable pass criteria, frequency dominant spectrum vies |  |
| 14. Ear probe tips – 2 packets (assorted) |  |
| 15. Reusable cup electrodes – 6 nos |  |
| 16. T 20 conduction gel – 3 nos |  |
| 17. Nuprep gel 114 gm - 6 nos. |  |
| 18. Desk Top Computer |  |
| 19. Processor: Intel core i3 |  |
|  | RAM : 4 GB |
|  | Hard Disk: 1TB |
|  | CD/ DVD |
|  | Monitor : LED/LCD 18.5 inch or above |
|  | Laser Printer |
| 20. Should be provided with table/trolley (Wooden/Ply/ MDF) to keep BERA machine, computer, UPS and printer |  |
| 21. Only UPS of suitable capacity with 30 minutes back up for BERA, Desktop and Printer shall be provide· |  |
| 22. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.. |  |

Department ; PAEDIATRICS

Digital Weighing Machine (for New Born)

Digital Weighing Machine for Baby :

1. Tabletop, light and portable.

2. Built in rechargeable battery.

3. Easy to clean baby tray (acrylic).

4. Zero weight adjustment facility.

5. Quick, clear digital read outs.

6. Measurement does not change with position of baby on the pan.

7. Provision to measure the height of the baby in its laying position.

8. Accuracy: 5g, resolution: 1g, limit: 10gm to 15kg.

9. Auto setting to 0.00 once a the machine is switched on or when no external weight has been put on.

10. 4XAA battery(rechargeable) or equivalent; one hour backup.

11. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

12. Should have model approval from Legal Metrology Dept., Govt. of India.

Department; ANAESTHESIOLOGY

Multipara Monitor- with P, NIBP, ECG facility, SpO2 Etco2, IBP

| Multipara Monitor: |  |
| --- | --- |
| 1. Should be able to monitor 5 lead ECG, SPO2, NIBP, EtCo2, 2 IBP, Respiration rate and Temperature. |  |
| 2. Should be portable with carrying handle. |  |
| 3. Should have touch screen TFT display with at least 10 inches or higher with at least 6 waveforms and numeric display simultaneously. |  |
| 4. Should have Lithium-ion battery with 1 hour’s battery backup. |  |
| 5. Should have keys for quick access to main functions. |  |
| 6. Should have adult, pediatric and neonatal modes. |  |
| 7. Should provide prominent prioritized audio, visual alarms for high, low heart rate, Spo2, RR, low battery, lethal arrhythmia recognition and ST Analysis. |  |
| 8. Should have separate volume control for beep sound for QRS and alarm sound. |  |
| 9. Should provide following accessories: |  |
|  | Reusable adult 3 lead ECG cable set – 2 nos. |
|  | Reusable adult and pediatric SPO2 finger probes – 1 each. |
|  | Adult and pediatric NIBP cuff of different size |
| 10. Should work on 200-240V AC / 50Hz with inbuilt rechargeable battery. |  |
| 11. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO. |  |
| 12. Electrical safety conforms to standards for electrical safety IEC-60601/IS-13450. |  |

Department ; PAEDIATRICS

Ophthalmoscope

Ophthalmoscope :

1. Available with LED/Halogen light source.

2. Magnification up to x15 from direct vision to maximum magnification.

3. Red-free, blue and polarization filters and Anti-reflection lens.

4. Should have small and large spot sizes, fixation targets, slit aperture, hemispot and cobalt blue filter.

5. Should be rechargeable battery with Charger / battery/ mains operated.

6. At least 3 apertures and fixation star.

7. Range of lenses not smaller than - 30D to +20D with steps not greater than 1D.

8. Dust free sealed optics and aspherical optical system.

9. Internal batteries, rechargeable preferred compatible with both 2.5 V and 3.5 V batteries or handles provided; Led display indicating the charging status.

Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; PAEDIATRICS

Mechanical Ventilator (Neonatal & Child)

| Mechanical Ventilator (Neonatal & Child) : |  |
| --- | --- |
| 1. Advanced technology ventilator for use in NICU, suitable for ventilating Premature Neonates patients. |  |
| 2. Should have facility for Invasive and Non-Invasive ventilation |  |
| 3. Microprocessor controlled system with individual selection of various ventilation parameters & PEEP. |  |
| 4. 12” inch size full colour, total touvh screen operation for big display to access from long distance. |  |
| 5. Machine should be Compressed air (medical oil free air compressor of the same brand as ventilator). |  |
| 6. Should have battery backup at least 30min. |  |
| 7. It should allow the user to deliver conventional ventilation as well as HFOV. |  |
| 8. Should have the following modes of ventilation: |  |
| a. Assist/ Control |  |
| b. Volume control |  |
| c. | Pressure control |
| d. Pressure support |  |
| e. SIMV with pressure support (Pressure and volume control) |  |
| f. | PEEP |
| g. | Inverse ratio Ventilation |
| h. Non-invasive ventilation-BIPAP, CPAP and Nasal C PAP |  |
| i. | Apnea ventilation, user selectable, volume & pressure control |
| j. | HFOV. |
| 9. Should have facility to measure and display of the following parameters: |  |
| a. Airway Pressure (Peak & Mean) |  |
| b. Tidal volume (Inspired & Expired) |  |
| c. | Minute volume (Inspired & Expired) |
| d. Respiratory mechanics |  |
| e. Spontaneous Minute Volume |  |
| f. | Total Frequency |
| g. | FiO2 |
| h. PEEP |  |
| i. | Plateau Pressure |
| j. | Use selector Alarms for all measured & monitored parameters |
| k. | Occlusion Pressure |
| l. | Pressure Flow & Volume curves |
| 10. Automatic compliance and leakage compensation for circuit. |  |
| 11. Conventional ventilation& HFO Ventilation Mode Parameters: |  |
| a. BPM: 4to150 |  |
| b. Inspiratory Time: 0.1 to 3.0 sec |  |
| c. | CPAP Pressure: 0 to 35 mbar |
| d. Inspiratory Pressure: 0 to 60 mbar |  |
| e. FIO2: 21% to 100% |  |
| f. | Tidal Volume 2-300 ml with Volume Guarantee |
| g. | HFO Mode Parameters: |
| h. HFO Frequency should be wide range with 5 to 20 Hz i. MAP-0-40M Bar |  |
| i. | I: E Ratio: 1:2 to 1:1, MAP-0-40m CmH2O, Delta P- 0-100 CmH2O, RR-4 to 150 bpm, Ti-0.1- 5 |
|  | sec, P-1-50m Bar. |
| 12. Alarm |  |
| a. Alarm :- a. Adjustable Alarm. - Low/high minute volume, low/high pressure, low/high tidal volume, low/high rate, apnea time, low/high oxygen |  |
| b. Special alarm - O2 cell Failure, flow sensor, battery, power supply, gas supply, oxygen concentration, |  |
| 13. Should have inbuild Nebulization assembly facility. |  |
| 14. Ventilator, Compressor & Humidifier should be Same Trolley/cart mounting for easy transportation. |  |
| 15. Humidifier |  |
| a. Servo controlled heated Respiratory Humidifier. |  |
| b. Display Should be of LED /LCD. |  |
| c. | Temperature control settings & Temperature range: 28-40 deg. |
| d. Temperature should be adjustable. |  |
| e. Jar should be autoclavable |  |
| 16. Standard Accessories/spare & Consumable. |  |
| a. Silicon breathing circuit circuit (Neonatal reusable) - 5 complete set. |  |
| b. Nebulization assembly compatible circuit 5 complete set. |  |
| c. | Humidifier - 1 No. |
| d. Hose for O2 connection with connector - 5 mts. |  |
| e. Hose for compressed air with connector - 5 mts. |  |
| f. | Test lung - 1 No. |
| g. | HME filter – 10 no |
| h. Inbuilt / integrated nebulizer-1 No |  |
| i. | All sensors and other non-consumable items (other than reusable silicon ventilator circuits) should be free of cost during warranty and CMC. |
| 17. Ventilator, Humidifier & Compressor Power Supply input to be 200-240VAC, 50 Hz fitted with Indian conditions plug . |  |
| 18. Suitable online UPS with commensurate capacity for all ventilators including compressor & Humidifier with maintenance free batteries for minimum Two hours back-up should be supplied. |  |
| Ventilator, Humidifier & Compressor Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO. |  |
| 1) Reusable consumables (other than reusable silicon ventilator circuits) should last during the warranty period. |  |
| 2) Ventilator & Humidifier any additional reusable consumables are required during the warranty period those will be supplied free of charge by the supplier. |  |
| 3) The life expectancy of the reusable consumable is expected to be of at least one year from the date of installation of the same. The reusable consumables will be procured at the prices accepted as per the contract. |  |

Department ; PAEDIATRICS

Suction Machine

Suction Machine :

1. Vacuum /LPM: - 700 MM Hg , 50 Liter’s/Min.

2. Pump Type- Double rotary vane type

3. flutter free vacuum control knob,

4. Wide mouthed 2 x 2 LITRE (light weight, unbreakable and clear) with self-sealing bungs and mechanical over flow safety device.

5. Noise (in dBA)- 50 dB A

6. Collection container & its cap, suctions tube tips, a vacuum gauge, two sets of moisture & microbial filters and control knob.

7. SiliconeTubing:8 mm ID x 2 mtr (PVC), 2x2 lt jar (one set extra)

8. It should be Mobility, portability.

9. Equipment shall operate on 220V-240V, 50 Hz, single phase electric supply.

10. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

11. ISO/ BIS approved model should be offered.

Department ; PAEDIATRICS

Transport Incubator

Transport Incubator :

1. It should have Visual and audible alarms for: (i) Patient and air high / low temperature alarm. (ii) Air circulation / probe / system / power failure alarm.

2. It should have Battery and AC supported.

3. Built-in-sealed rechargeable batteries capable of working for at least 3 hrs when fully charged.

4. Internal noise level 60 dB.

5. Mode of operation should be properly displayed.

6. Indicator light should be provided for its ready to be in normal use.

7. Infants straps should be provided to restrict the baby movement.

8. Skin temperature probe should be small in size not more than 10mm diameter and 4mm in height to fix the probe firmly on the infant. Baby contact material should be biocompatible.

9. Infant bed should be drawable. Mattress foam density should be minimum 25kg./cm3 and infant bed mattress cover should be biocompatible material. Baby Mattress should be at least 60X30cm.

10. Examination light should be provided for inspection.

11. Warm up time 30-40 minutes and shall not differ by more than 20%.

12. Shall be equipped with a thermal cut-out. It shall be so arranged that the heater is disconnected and an audible and visual warning is given at an incubator temperature which does not exceed 40⁰ C.

13. Patient skin temperature range: 35⁰ C to 37.5⁰ C. Temperature accuracy less than ± 0.2⁰C.

14. Air temperature range: 30 ⁰C to 39⁰C. Temperature accuracy less than ± 0.2⁰C.

15. Display is to be backlit and allows easy viewing in all ambient light levels.

16. Should have Oxygen port with tubing, also mount for oxygen cylinder of 5 litre size.

17. Should have Accommodates shelves, suction unit and I/V poles.

18. Should have collapsible trolley with lockable castors.

19. Two extra sets of all sensors.

20. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

21. Power Supply 220VAC +/- 10 %, 50Hz fitted with Indian plug.

Department : PATHOLOGY

Manual Rotary Microtome

Technical Specification

1. High precision machine suitable for both delicate as well as hard tissue sectioning

2. Section thickness settings 1-60 μm with settings in 1,2,5 increment at different levels

3. Specimen advance 28 mm or more

4. Vertical stroke 60 mm or more

5. Provision of step trimming

6. Adjustable specimen clamp at atleast 50 x 45 mm with orientation in X,Y axis

7. Single disposable blade holder for accommodating both high and low profile blades

8. Lateral coarse feed

9. Integrate removable section waste tray

10. Spare low and high profile blades in dispenser pack of 50 blades: 6 packets each

11. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department : PATHOLOGY

Automated Rotary Microtome

1. Fully automated motorized rotary microtome along with manual operation having microprocessor controlled panel with provision for motorized cutting via operating panel or foot pedal control.

2. Precise micrometer feed system via stepper motor permits precision sectioning selectable at least from 2.0-40/60 micron in 0.5 micron increments.

3. Trimming section selectable from 2 micron onwards.

4. The vertical specimen stroke length of 70mm, larger specimen can be sectioned. The specimen holder should be clamp type and hold 60mm size block.

5. Suitable knife holder for high profile and low profile should be provided.

6. The specimen retraction should occur on return stroke.

7. Knife angle position locking facility should be provided.

8. Cold light source.

9. Precise specimen orientation with zero point indication, with an orientation 6º X-Y axis helps in marking perfect orientation of the sample for sectioning.

10. Motorised coarse feed in two speeds 30 micron/sec and 90 micron/sec. Variable sectioning speed adjustable from 0.5 to 420 mm/sec.

11. Disposable blade holder with lateral displacement feature that can hold both high and low profile blades and knife holder which can accommodate 16 cm C&D type knifes.

12. Knife holder should be vibration free.

13. Integrated section waste tray.

14. Accessories

15. Microtome disposable blades (high profile coated) - 20 packets (50 blades/pack). Total 1000 Nos.

16. Microtome disposable blades (low profile coated) - 2 packets (50 blades/pack).

17. C type knife 16 cm - 3 Nos

18. Cold plate (Dry type)

19. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department : PATHOLOGY

Cryostat

1.   Open top, healed sliding window, corrosion proof, stainless steel cryo chamber with good fluorescent illumination.

2.   Cooling via two separate refrigerator system.

3. Temperature of cryo chamber should be at least -30º C. Facility for integrated quick specimen freezing up to -45º C. Separate cooling should be adjustable up to -50º C.

4. Temperature of cryo chamber should be maintained within ± 2º C of set temperature and maintained by hermetically sealed compressor system.

5. Automatic programmable defrosting and manual defrosting should also be possible.

6. Fully motorized microtome - movement controlled by manual as well as foot switch.

7. Microtome should be encapsulated to support efficient spray disinfection.

8. Microprocessor/Microcontroller based touch key control panel with LCD display for all functions incrading microtome.

9. Space for other specimen rack minimum 6 blocks. Removable section waste tray.

10. Section thicknes setting must be outside the cryo chamber.

11. Disposable blade holder for low and high profile blades and knife which can hold minimum 16cm C type knife.

12. Specimen holder can hold specimen size upto 70 x 50 mm. Section thickness cutting 1-60 micrometer.

13. Specimen retraction around 50 micron.

14. Trimming in steps from 5 to 150 Microns.

15. Motorised coarse speed 500 micro meter/sec and 1000 micrometer/sec.

16. Cryo cabinet should be of appropriate size.

17. Accessories

18. Microtome knife 160μm - 4 Nos.

19. Suitable UPS

20. Two bottle of low temperature oil

21. High and Low profile blades - 5 packets each

22. Specimen holder of appropriate size - 20 Nos

23. Freezing compound at least 10 bottles

24. Glass holding device for knife and blades.

25. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department : PATHOLOGY

Paraffin embedding bath

1.   Embedding station should have 2 separate components cold plate heated paraffin dispensing module.

2.   Module should allow flexibility to arrange embedding work flow in the either direction.

| 3. Capacity of Paraffin reservoir should be 3 lit. capacity with vacuum lid, easy to clean, scratch resistant, heated wailing surface with paraffin drain system. |  |
| --- | --- |
| 4. Adjustable paraffin flow rate. |  |
| 5. Two heated removable paraffin waste tray. |  |
| 6. Paraffin outlet: manual or foot pedal operated paraffin outlet. |  |
| 7. Removable and interchangeable cassette and mould warmer. |  |
| 8. Large peltier cooling spot, even for super mega cassettes. |  |
| 9. Adjustable temperature for all modules between 55 to 70 Deg C |  |
| 10. Heated forceps holder which is easily removed for cleaning. |  |
| 11. Capacity of approximate ranging 100 to 150 cassettes. |  |
| 12. Programmable timing and temperature. |  |
| 13. Magnifier (essential) |  |
| 14. Illuminator (essential) |  |
| 15. Main:230V,50 Hz |  |
| 16. Height - 400 mm |  |
| 17. Width - 500 mm |  |
| 18. Depth - 700 mm |  |
| 19. Weight - 30 Kg |  |
| 20. Size of working surface - 2 x 175 mm |  |
| 21. Size of cold spot - 70 mm |  |
| 22. Accessories |  |
| a. Removable holder. |  |
| b. Paraffin scrapper |  |
| c. | Freezer |
| d. Replacement bulb |  |
| e. Set of fuse |  |
| f. | Set of power cords |
| g. | Cassette - 50,000 labels |
| 23. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO. |  |

Department : PATHOLOGY

Heated Paraffin Embedding Module

1.   The  instrument  should  incorporate  two  separate  systems  for  cold  plate  and  heated  paraffin embedding module.

2.   Temperature range of cold plate: -5®C, adjustable in steps of 1 deg C.

3.   At least 60 cassette molds capacity with acrylic cover.

4. Two number of Paraffin collection tray.

5. Heated embedding module should have adjustable paraffin dispenser control with paraffin flow rate adjustment.

6. Paraffin reservoir capacity of at least 3 litres.

7. Working temperature range between 50-75®C.

8. Foot switch operated.

9. Precisely metered and adjustable gravity feed paraffin dispenser to deliver the right amount of paraffin.

10. Large warm working surface on either side for keeping a minimum 10 cassettes on each side.

11. Flow adjustment in continuous mode.

12. Heated removable paraffin waste tray.

13. Six forceps holder with temperature range of 50-75®C

14. Working surface of cold plate should not be less than 300 x 350 mm

15. Should have a magnifying lens adjustable in any position and white light illumination for specimen orientation.

16. All functions of the system controlled through electronic system with digital program able on and off timer

17. The system should work on 220-240 V, 50 Hz.

18. Provide Power back up (UPS)

19. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department : PATHOLOGY

Automated Tissue Processor

1. Microprocessor controlled tissue processor with changeable duration in each station, delay start of at least 7 days, vacuum application and agitation at station.

2. It should have capacity to hold 150 to 200 tissues cassettes

3. Reagent Jars and wax impregnation bath should be of material resistant to chemicals like chloroform, xylene, alcohol etc. Metal jar preferable.

4. The wax bath temperature should be such that fumes from reagent will not be emitted out of the system and there will be minimal wastage due to evaporation. Extra two wax impregnated bath to be supplied with the instrument

5. Capsule basket holder - 2 hanging types.

6. Duration time of 60 second in each section to reduce carry over contamination.

7. Maximum safety concept with a liter beaker in case of power failure.

8. Audible alarms, error messages and warning message for maximum safety.

9.   Electronic locking facility to avoid inadvertent operation.

10. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department : PATHOLOGY

~ Ultrapure water solutions - Distilled water plant

| a) Ultrapure water purification system b) Reverse Osmosis based System. |  |
| --- | --- |
| c) Resistivity of water produced should be between 10 - 15 meg Ohms-cm. |  |
| d) Purified Water Flow Rate: Should be at least 8 Liters per hour |  |
| e) Storage tank of at least 25 Liters |  |
| f) Should have self-diagnostic control feature automatically giving alarms. |  |
| g) Filtration System: LV. 1 20" 5μm PP filter h) Resistivity MΩ-cm : 17.0 - 18.2 |  |
| i) Micro-organisms should be < 1 cfu/ml |  |
| j) | Pyrogen level should be +/< 0.001 EU/ml |
| k) Monitor system: LCD Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO. |  |

Department : PATHOLOGY

~ Centrifuge Machine

Centrifuge Machine (16 tube)

1. Should have a maximum speed of 5000 RPM with steeples regulator

2. Should be supplied with safety lid and lock & Brushless motor.

3. Should have digital speedometer and timer.

4. Should have imbalance detector and automatic cut off.

5. Standard mark on quality and safety is Compulsory.

6. Should work on a220VAC +/- 10 %, 50Hz AC Supply.

7. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department : PATHOLOGY

Fully Automated high throughput Multi-Stainer Workstation

1. High throughput robotic strainer to process up to 11 racks at one time (capacity of each of rack - so specimen slides)

2. Simultaneous staining of various staining protocols.

3. 18 reagent stations and 5 wash stations of 450 ml capacity.

4. 16 programs up to 25 programmed - permanent memory capacity.

5. Incubation time setting is from 0 second up to within 59 sec.

6. Integrated oven with temperature setting for 30º to 65º C for optimal slide drying.

7. Continues loading and unloading of sliding via rack entry and exit door.

8. Load/unload station one each.

9. Specimen slide throughput of at least 200-600 slides per hour.

10. Agitation programmable from 0 to 20 times or continuous.

11. Up and down movement of robotic arm can be programmed

12. Fume extraction fan with charcoal filter to remove hazardous fumes.

13. Vibration to slide rack during lifting to reduce carry over contamination.

14. Audible warning buzzer in case of any error during operation.

15. Staining jars / reagent stations contain lids to reduce the evaporation.

16. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department : PATHOLOGY

Fully Automated Embedding System (Heated embedding module & cold plate)

1. Embedding station should have 2 separate components cold plate heated paraffin dispensing module.

2. Module should allow flexibility to arrange embedding work flow in the either direction.

3. Capacity of Paraffin reservoir should be 3 lit. capacity with vacuum lid, easy to clean, scratch resistant, heated wailing surface with paraffin drain system.

4. Adjustable paraffin flow rate.

5. Two heated removable paraffin waste tray.

6. Paraffin outlet: manual or foot pedal operated paraffin outlet.

7. Removable and interchangeable cassette and mould warmer.

8. Large peltier cooling spot, even for super mega cassettes.

9. Adjustable temperature for all modules between 55 to 70 Deg C.

10. Heated forceps holder which is easily removed for cleaning.

11. Capacity of approximate ranging 100 to 150 cassettes.

| 12. Programmable timing and temperature. 13. Magnifier (essential) |  |
| --- | --- |
| 14. Illuminator (essential) |  |
| 15. Main:230V,50 Hz |  |
| 16. Height - 400 mm |  |
| 17. Width - 500 mm |  |
| 18. Depth - 700 mm |  |
| 19. Weight - 30 Kg |  |
| 20. Size of working surface - 2 x 175 mm |  |
| 21. Size of cold spot - 70 mm |  |
| 22. Accessories |  |
| a. Removable holder. |  |
| b. Paraffin scrapper |  |
| c. Freezer |  |
| d. Replacement bulb e. Set of fuse |  |
| f. Set of power cords |  |
| g. | Cassette - 50,000 labels |
| h. Stainless steel capsules for matismould. 23. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.. |  |

Department : PATHOLOGY

Fully Automated Flexible Coverslipping Workstation

1. Should produce slides with superior optical quality for reliable long-term storage.

2. It should allow for each single glass slide separately and without bubbles.

3. It should have active carbon filters for safety.

4. It should have the storage capacity of about 200- 350 pcs. of slides.

5. Should be capable of cover slipping 200 - 300 slides per hour

6. Should be able to handle slide racks of various manufacturers and should be adaptable to individual laboratory requirements

7. Should be used with common range of mounting media including mounting with wet mount ants. Should dispense adequate amount of mount ant for cover slipping each slide.

8. Should be equally useful for histopathology and cytopathology slides

9. Should be highly reliable, cause minimum wastage and form a fully automated walk-away system.

10. Should have an inbuilt system for fume extraction to minimize exposure of lab personnel

11. Instrument should have an operating voltage suitable for Indian plugs. Voltage supply 230 V- 50/60 Hz.

12. Should be provided with UPS 30 minutes battery backup.

13. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department : PATHOLOGY

Single Pan Digital Balance, Chemical

Balance single pan

1. Easy to read Large back light GRAPHICAL LCD display with A.E.P.(Advanced Eye protection) (NEW)

2. Standard RS 232C Interface Ps/2 output

3. Hanger for Below Balance weighing

4. Fully automatic internal Calibration with built in weight

5. Complies GLP / GMP.

6. Dye cast aluminum design for long term stability and accurate results.

7. Various weighing units like, mg ,ct. oz, dwt, mon, GN

8. User selectable Stability and filter level Spacious draft shield interior.

9. Technical data: Capacity 220 gm: Readability 0.1mg: REPEATABILITY (+/-) 0.1mg;linearity (+/-) 0.2mg; PAN Size (mm/inch) 90Ǿ; response time: 03sec, display Back light LCD graphical display;

calibration automatic external  units of measure; G,mg, ct, GN, mo,  oz,  dwt, T are range full operating; temperature 5deg. C to 40 deg.C., housing dimension (342.5 mm X 212 mm X 341 mm – WDH.

10. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department : PATHOLOGY

Balance, chemical with weights

Technical Specification

1. Microprocessor based single pan Analytical Balance with High accuracy & precision is required. Reading of the weight by digital display.

2. Weigh accurately up to 3rd decimal place.

3. Auto self-calibration facility.

4. Auto zero Setting, TARE setting.

5. One touch calibration.

6. Weighing capacity up to 120g.

7. Readability and repeatability 0.001g.

8. Stabilization time < 5 second.

9. Liquid Crystal Display (LCD) for display.

10. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department : PATHOLOGY

Microscope for Diagnostic & Research Work

Microscope for Diagnostic & Research Work Type with Attachment for Camera

1. Research microscope with digital image analysis system.

2. Magnification 40X-100X for observation. Optical system fly-eye optics (infinity corrected optical system).

3. High intensity and uniform eco-illumination

4. Trinocular eyepiece tube with variable incline at 10-20-degree angle and extension up to 30mm with for 25mm.

5. Should be antifungals type.

6. There should be provision to attach camera which should be included in the main tender. Eyepiece lens 10X (2 pcs) with both sides diopter adjustment (fov 25mm) objectives. High performance chromatic aberration free plan-objective suitable for bright field microscope.

a. Achromat 4X na 0.10, wd 30.00 mm or better

b. Achromat 10X na 0.25, wd 7.00 mm or better

| c. | Achromat 20X na 0.45, wd 3.90 mm or better |
| --- | --- |
| d. Achromat 40X na 0.65, wd 0.65 mm or better |  |
| e. Achromat 100X na 1.25, wd 0.23 mm oil or better |  |
| f. | Nose piece-sextuple revolving nosepiece. |
| g. | Condenser-swing out condenser within and vertical movement |
| 7. Adjustable course and fine focusing. |  |
| 8. Stage-ceramic coated XY stage with two slide holder capacity at a time. |  |
| 9. Illumination-high performance white LED illumination (auto intensity reproduction function)/ 100W halogen illumination. 3 spare white light illumination bulbs should be supplied with each unit at no extra cost. |  |
| 10. Accessories- Apart from the spare bulbs, lens cleaning paper and at least 2 units of branded research grade immersion oil units should be supplied premiers care at no extra cost. |  |
| 11. Digital scientific grade color CCD camera:- Digital color camera with 2/3”high density CCD chip, of at least 5 million pixel resolution. At least 14-bit dynamic output. Binning zoom, and live capability should be there. Sensitivity equivalent to ISO 60 or better. Dynamic range 1000:1 or better. |  |
| 12. Microscope camera and image analysis software should be compatible with options for up gradation to newer version as and when available within warranty period at no extra cost. |  |
| 13. Image acquisition software should be user friendly, innovative and be supplied with the system at no extra cost. Installation drivers should be supplied. The image acquisition and image analysis software should be compatible. |  |
| 14. Computer for camera photography: branded desktop with intel i5 processor, at least 4GB RAM, 1TB HDD, DVD writer, optical keyboard, mouse, 22” color high dimension TFT screen should be supplied with the instrument. |  |
| 15. Online UPS of at suitable capacity should be supplied with the instrument. All required cables and accessories should be provided. |  |
| 16. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO. |  |

Department : PATHOLOGY

Penta Head Microscope with High end Optics with HDMI Multi output

Photographic camera (> 5 MP) including Software

Penta Head Microscope with High end Optics with HDMI Multi output Photographic camera (> 5 MP) including Software

1. The instrument should be sturdy, fitted with plan achromatic objectives 2/2.5x, 4x, 10x, 20x, 40x (spring loaded and 100x (spring loaded) on a reversed sextuple nosepiece with dick stop.

2. The optical system should be color corrected for infinity with antifungals property built in transmitted Kohler illumination.

3. The microscope stand should have co-axial focusing knobs for coarse and fine adjustment with upper limit stopper.

4. Present button for automatic light intensity level for photomicrography.

5. Wide field high point eye piece 10x, 22mm with diopter adjustment (+2 to -8) and rubber eye shield (pair) with inter pupillary distance of 48 to 75mm.

6. Trinocular eye piece inclined at 30-45 Degree with 360º rotation.

7. Rectangular scratch resistant stage with right hand control with double slide holder and Vernier calipers on XY axis.

8. Plan achromatic universal type swing-out condenser (dry type) with numerical aperture 0.9-1.2.

9. Transmitted light filters for day light, green and neutral light with density filters built-in die basic stand.

10. Illumination: LED ILLUMINATION- LONG LIFE

11. Power: 220-240V, 50 Hz

12. Vinyl dust cover

13. Multi head ergonomic 1 Trinocular set (with three-way light pam selector, 100:0, 80:0,0) +4 Binocular heads (2 on each side) with complete two color pointer unit (1 pc), ac adapter (1 pc), power cord (1 pc)

14. All the necessary adaptor and power card should be provided for functioning of microscope.

15. Digital Camera: Fire wire digital camera with the following features: Recent model with 18 Megapixel CCD camera with appropriate lens system and mounted

16. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department : PATHOLOGY

Deca Head Microscope with High end Optics with HDMI Multi output

Photographic camera (> 5 MP) including Software

Microscopes for Deca Head Microscope

1. Microscope Frame: Upright microscope with Built in Koehler illumination for transmitted light LED (pre centred) uniform distribution, with lifespan of at least 50,000 hrs. or more (higher will be preferred).

2. Observation tube and Eyepieces: 3. Main Head: Trinocular observation tube with three-way light distribution (100:0, 20:80, 0:100), have 26 MM FOV or more. 4. Teaching Head: Binocular teaching heads Siedentop type for ten-person including main observer.

5. Eyepiece: Main observer: 10X magnification with F.O.V 26mm - 2 no’s or more. With both side dioptre adjustment facility. LED arrow pointer tricolour, joy stick should be provided by supplier Teaching head: for 1+9, 10X magnification with F.O.V minimum 22mm with both side dioptre adjustment facility.

6. Mechanical Stage: Ceramic coated coaxial stage with right-hand drive control with two slide holders.

7. Condenser: High quality swing out universal turret condenser for all application.

8. Nosepiece: Reversed Sextuple (Six position) revolving nosepiece.

9. Objectives: High performance objective should be suitable for bright field and polarizing application. Plan Fluor2X / 2.5X, Plan Flour 4X , Plan Fluor 10X, Plan Fluor 20X, Plan Fluor 40X , Plan Fluor 100X. Should provide required accessories for polarizer and analyser for simple polarising.

10. Teaching attachment: Teaching Head for Ten person including main observer. with eyepiece pointer 360 degree rotable and should have color variation and intensity control feature.

11. Should quote latest model as per above specification.

12. List of important spare parts and accessories should be available for next 10 years.

13. Vinyl dust cover for entire unit, immersion oil minimum100ml and other necessities for installation and commissioning of instrument should be provided by supplier only at their own cost.

14. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department : PATHOLOGY

Fully Automated Immuno-histo-chemistry Setup with Continuous supply

of Important Antibodies, Lymphoma Panel etc.

| Operational Requirements |  |
| --- | --- |
| a. Fully Automated Random-Access System: Immunoassay of more than 90 different parameters all Hormones, all Tumor Markers, all Cardiac markers. |  |
| Technical Specification |  |
| a. System should be able to perform Routine & STAT assays |  |
| b. The Equipment should have a Throughput of not less than 100 tests per hour |  |
| c. | Should have Two Precision Syringes for accurate delivery of Samples and Reagents. |
| d. System should have continuous Loading of Samples and Reagents. Must also have 12 or more reagent loading at a time & Triple Marker test capability. |  |
| e. Equipment must have a integrated Water and Probe Wash system. Centrifugal Washing technique and Automatic reagent level indication by Sensors. |  |
| f. | Audible and Visual Alarms for all error messages. |
| g. | Equipment must have facility for Automatic /computer assisted manual Dilution of Sample. |
| h. The Equipment should have flexible Windows based software, LIS interface and real time system monitoring. Optional Bar Coding & Color Coding with State-of-the-Art Software. |  |
| i. | The equipment should be managed by a Computer and have RS232 interface, software for control. Data evaluation & management. Extensive QC graphics including L-J plots, QC management. |
| j. | The Specification of the computer should be having a microprocessor of speed not less than 3.0 GHz, 8 GB RAM, 380 GB HDD, 105 key ergonomic key Board, scroll mouse, 56 kbps modem, 32 MB AGP Card, 52x CD-RW Drive, with 17" TFT Digital Color Monitor with Operating system and compatible Laser printer for documentation having minimum 600 DPI resolution, not less than 12 pages per minute speed. |
| k. | Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO. |

Department : PATHOLOGY

Five Part Haematology Analyzer.

1.   It should include 24 parameter including histogram for RBC, WBC and platelet.

2.   Should be capable of processing 60 samples/hour.

3. Should have minimum 3 sampling mode. Manual mode, manual closed mode and capillary mode.

4. Sample volume should have manual mode for 85 ml and closed mode for 150 ml. or less than that.

5. Should have dual channel measurement.

6. Double dilution chamber.

7. Should have a data storage of 10,000 sample inclusive of Scattergrams and histograms

8. Should have a facility to store atleast 5000 patient information.

9. Should have flourescent base flow cytometry technology using semi conductor laser / Optical laser light scattering with impedance for WBC and WBC 5 differential measurement.

10. Should have DC detection method for measuring RBC and PLT.

11. Should have automatic floating thresholds for seperation of RBCs and Platelets during overlab of RBC and platelets abnormal lies.

12. Should have the facility to operate the instrument, QC data management and maintenance through dedicated software with computer (which should be supplied along with instrument)

13. QC data management should have L-J, XbarM and radar chart with real time QC.

14. Should have facility to classify normal and abnormal sample with colour code.

15. Should have facility to display pathological abnormalities and flagging.

16. Should have patient cumulative results and related flag system.

17. Should have impedance principle for counting and photometer for hemoglobin.

18. Should have dual channel measurement.

19. Double dilution chamber.

20. Sample volume less than 20 microlitres in whole blood and pre-dilute mode.

21. It should have various types of discrete mode and real time random access analysis to save reagent consumption and analysis time.

22. Sampling needle should have automatic wash from inside and outside.

23. LCD graphical user interface (GUI) for easy operation.

24. Large illuminated colored VGA/ LCD should display the result of all parameters and histogram together.

25. Should have capacity to store at least 20000 numeric patient results and 5000 graphics.

26. Should have inbuilt graphic printer.

27. Should have RS 232 serial and parallel port can be connected with LAN and laser printer.

28. Should have a membrane keyboard for routine operations and maintenance with option to attach external key board for patient demographic entry at instrument operation.

29. Should have extended analysis time for cytopenic sample.

30. Should have zero routine maintenance with automated slide maker and strainer.

31. Should have zero routine maintenance with automatic electronic aperture cleaning and back flush after each sample.

32. Instrument should accept all types of vacutainer tubes.

33. The instrument should have option for auto sampler, bar code.

34. Reagent cost per cycle including start up and shut done if 200 & 500 samples are processed at a time should be submitted separately in the financial bid.

35. There should be automatic storage of calibration data and extensive quality control programmed with LJ plot for at least 8 control and at least 25 runs per lot.

36. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department : PATHOLOGY

Three Part Fully Automated Cell Counter.

| 1. Should be fully automated three part haematology Analyzer providing 20 parameters including a 3- part differential |  |
| --- | --- |
| 2. The system should give a differential count as Lymphocytes, mid population and Granulocytes. |  |
| 3. System should be capable of processing samples at 70 samples/hour & storage memory result capacity 10000. |  |
| 4. The system should be based on “closed, maintenance free Sample Rotary Valve (SRV)” for precise sample all quoting for dilution. |  |
| 5. System should have auto Probe wiper to clean the sample probe automatically after sample aspiration and with predilution mode. |  |
| 6. The system should use non cyanide based reagents for Hb estimation. |  |
| 7. System for the reliability of the results should have "electrical Impedance" method of cell counting with an integrated temperature sensor for monitoring and compensating for shift in room temperature |  |
| 8. The system should use to proven and approved "Volumetric & time Metering" of cell counting, for WBC, RBC and PLT for high precision of the results and stability of the calibration with close measuring chamber. |  |
| 9. The system should have a system of count and aperture monitoring every 30 sec for precision and reliability of counts. |  |
| 10. The system should have automatic floating discriminator of RBC /PLT. |  |
| 11. The system should have Open mode as well as prediluted mode of sample aspiration. |  |
| 12. The system should use high Intensity LED for HB estimation. |  |
| 13. All reagents required should be available locally from the company or its authorized distributor |  |

14. System should be user friendly with colour touch screen and should option for external printer as well as data inter facing.

15. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department : PATHOLOGY

Coagulometer (Fully automated).

1. Viscosity / Optical based detection system

2. Should be Capable to run Clotting assay, Chromogenic assay & Innmunoturbidimetric assays

3. Parameters- PT, APTT, Fibrinogen, Factor VIII, IX, VII, X, XI, XII, XIII,TT, ATIII, Heparin, Protien C and S

4. Minimum Throughput should be 50 PT test/hour & time to first patents result - PT/APTT/FIB = 6 mts

5. Should be minimum 80 test Methodologist available

6. Should be minimum 15 Reagent positions. All should be temperature controlled & and maintained from 15 to 19 deg C. Each reagent can be placed in several positions and analyzer will detect automatically without manual intervention.

7. Should be Positive barcode identification for reagents. Should be Barcode identification for reagent name, Lot number, Expiry, on board stability.

8. Should be minimum 20 samples carrousels for primary tubes. Capable of adapting pediatric tube & Microtainers

9. Should be Positive barcode identification for samples

10. Continues sample loading with ail positions for STAT mode.

11. Should be at least 12 programmable tests for one sample.

12. Should be Single probe for Samples & Reagents with LLD ( Liquid level detection)

13. Should be minimum 200 cuvettes on board with continuous loading capability.

14. Should be Pre calibrated assays for PT, APTT, FIB & D. Dimer.

15. Should be Storage capacity for Calibration curves. Automatic dilution for calibrations.

16. Should be Multi tasking software with touch screen with LCD display.

17. Should be at least 600 patients' results & One year IQC results in memory.

18. Should be QC management with Levy Jennings chart & Wesdgard QC alarms.

19. Should be facility of Alarm for QC out range message.

20. Should be availability of internal printer or capability to adapt external printer.

21. Should be LIS capability with Bi directional transferring capacity.

22. Should be facility to auto rerun with operator's rules.

23. Should have Windows 10 operating system for updated cyber security with touch screen facility

24. Should have same manufacturer for instrument, reagents and consumables

25. Should have insensitive to icteric , lipemic and haemolysed samples.

26. System Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department : PATHOLOGY

Autoclave

| Horizontal rectangular or vertical cylindrical high-pressure steam sterilizer |  |  |  |
| --- | --- | --- | --- |
| Chamber capacity: 75 litre |  |  |  |
| Electrical Power : 18/36 KW or Sufficient wattage of Industrial immersion type water heater to generate steam within a reasonable period of time on 3 phase 440 v 50 HZ ac supply |  |  |  |
| Working pressure and Temperature : 1.06 Kgf / sq.cm -1.2 Kgf /sq.cm at 121 Deg C |  |  |  |
| Material of construction |  |  |  |
|  | Inner chamber, Jacket, Door : SS 316.(5mm-10mm) |  |  |
|  | Outer Chamber : SS 304 (insulated properly) |  |  |
|  | Steam generator : Non corrosive SS / Chromium plated Brass |  |  |
|  | Heater Plate : Brass / Stainless steel |  |  |
|  | Pipe Line : Complete with SS |  |  |
|  | Stand : Stainless Steel/High quality non corrosive Steel Instrumentation |  |  |
| Temperature, Pressure and vacuum gauges, steam traps, vacuum driers, water level indicator on steam generator Safety devices. |  |  |  |
| Pressure switch and safety valve, self-locking of door when chamber is under pressure: vacuum breaker for jacket. |  |  |  |
| Steam generator with gauge glass valves and Low water protection with audiovisual indicator. |  |  |  |
| Entire cycles of sterilization should be controlled by a microprocessor unit and it should also have manual operation facility. |  |  |  |
| Accessories |  |  |  |
|  | Additional Accessories required: |  |  |
| 1. Heating element 2 set extra with each machine. |  |  |  |
|  | Requirements at installation: |  |  |
|  | 1. | Necessary Pipeline works for water inlet and steam outlet | should be done. |
|  | Cable for connecting to wall socket should also be provided |  |  |
| Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO. |  |  |  |

Department ; PATHOLOGY

Water Baths

Water Bath :

1. Should have a double walled construction.

2. The inner chamber and top lid should be made of stainless steel.

3. The space between the two walls should be packed with thick glass wool.

4. Should provide with a thermostat control.

5. Working temperature should be from ambient+5 ºC to 80ºC having an accuracy of +1 ºC

6. Should have an approximate inner chamber dimension of 350mm x 250mm x 125mm.

7. Should be operated on 230V, 50 Hz single phase AC supply

8. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department : PHARMACOLOGY

Manikins for demonstration of intravenous injection, enema, local, intramuscular injections, intracardiac injection and other routes of drug administration

| 1. The Manikin should be a full body manikin designed to run advance life support course. |  |
| --- | --- |
| 2. The manikin should allow to teach patient assessment and physical examination skills including |  |
| a. Head to toe examination |  |
| b. Pupillary reaction with multiple scenarios |  |
| c. | Tongue swelling (angioedema) |
| d. Carotid pulse |  |
| e. Visible chest rise |  |
| f. | Lung auscultation including normal and abnormal breath sounds |
| g. | Pneumothorax |
| h. Heart sound including normal and abnormal |  |
| i. | Bowel sound |
| j. | Blood pressure measurement |
| k. | ECG monitoring |
| 3. The manikin should allow to teach airway technique including Headtilt chin lift, C-E technique for Mask holding, jaw thrust. All basic, supra glottis, advanced airways and surgical airway procedures |  |
| 4. The manikin should have CPR feedback as per 2015 guidelines. The manikin should have feedback on depth chest compression, chest recoil, rate of compression, interruption time, ventilation volume and should also provide score on over all CPR skills. |  |
| 5. Should allow to perform needle decompression and chest drain Should have IV arm with veins that IV accesses as a part of drug therapy. |  |
| 6. The manikin should generate all ACLS normal and abnormal ECG rhythms. |  |
| 7. The manikin should be able to receive real Defib shock of 360/200 Joules. |  |
| 8. The manikin should have automatic carotid, brachial and radial pulse, the pulse should have synchronized with ECG and blood pressure. |  |
| 9. The manikin should have facility to have vocal sounds for patient to participant communication. |  |
| 10. The manikin should have visible chest rise coordinating with the set respiratory rate. The lung sound, chest rise, and RR should coordinate with each other. |  |
| 11. The manikin should be controlled by a wireless, touch screen control unit. |  |

Department ; PHYSICAL MEDICINE & REHABILITATION

Skeleton Traction Set

Skeleton Traction Set :

1. Should be able to deliver decompression therapy.

2. Should have a touch screen interactive display for easy treatment set-ups and easy angle selection and must come with treatment protocol manual; Must provide along with the package Angle reference chart.

3. Should be able to automatically calculate and digitally display the rope pull angle for decompression and traction as per the treatment protocol.

4. Should provide protocol manual for light therapy and lumbar and cervical protocol manual Must be a build in computerized software package and protection against accidental setting of force-must have a safety switch for emergency shut off. Hold time: Ito 99 Sec with Digital Display

5. Rest time: I to 99 Sec with Digital Display

6. Traction Force: 4-45 Kg (With Double 90 Kg)

7. Cervical: 4-15 Kg (Each 1 kg step)

8. Lumbar 23 to 45 Kg (Each 2 kg Steps)

9. Safety Switch.

10. Should come with Flexion stool, Knee bolsters, Cervical pillow, Ankle bolsters, and decompression belts thoracic and pelvic.

11. Should provide with a 4-section motorized table hi/low with clamps, frame attachments for connecting the traction unit.

12. Should operate from AC 220 VAC ± 10 %, 50 Hz & with Suitable capacity Servo Stabilizer.

13. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; PHYSICAL MEDICINE & REHABILITATION

Inter Ferential Therapy Unit

Interferential Therapy Unit :

1. Dual output Channels and isolated between channels

2. Should have 0-30 operation programs

3. Symmetrical Balanced Sine Wave

4. Output Current:0-100 mA

5. Interferance Frequency 2-160 Hz

6. Output Frequency 4000Hz (with ±1% tolerance) fixed on Channel 1

7. Modulating Frequency 4002 ― 4160Hz (with ±1% tolerance) adjustable on Channel 2

8. Treatment Timer Continuous, 15, 30, 45 or 60 minutes

9. 2pole/4pole multi vector mode

10. Patient Safety Fuse

11. Accessories, -

One set Patient wire IFT

Two set Fixation straps

One jell bottle

One Power cable

One operating manual

Big and Small rubber electrode

12. Should operate from AC 220 VAC ± 10 %, 50 Hz & with Suitable capacity Servo Stabilizer.

13. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; PHYSICAL MEDICINE & REHABILITATION

Short Wave Diathermy

Short Wave Diathermy :

1. The unit should offer minimum 20 preset therapeutic protocols with electrode placement images to make operation of the simple and convenient.

2. Should have an output of up to 500 W in continuous mode.

3. 800 to 1100 w in pulse mode.

4. Pulse repetition frequency of 20 to 200 Hz adjustable in 10 steps

5. LCD Screen display of parameter Treatment timer with all standard.

6. 30 minutes treatment timer & tuner control .

7. Accessories

Condenser pad with cable

Disc electrodes with arms and cables.

Patient safety switch

8. Should operate from AC 220 VAC ± 10 %, 50 Hz & with Suitable capacity Servo Stabilizer.

9. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; PHYSICAL MEDICINE & REHABILITATION

Hotpacks & Hydrocollator

Hotpacks & Hydrocollator :

1. The Unit should be provided with thermostat temperature control.

2. The unit should have detachable insert rack to hold and suspend packs for heating.

3. The Unit should made of stainless steel with rubber wheels for mobility.

4. Tank Capacity 30-50 Liters approx.

5. Filled Weight 50 kg - 80 kg approx.

6. Temperature Range upto 90° C)

7. Thermal Cut-out Temp (88° C ± 8° C)

8. Heating Up Time Time to (70° C) - 3 Hours

9. Cool Down Time Time from 160° F (70° C) - 2 Hours

10. Device should be Safety Class Type B and Safety Tests UL 544

11. Heating Indicator

12. It Should have Drainage valve .

13. Accessories

1. Includes two Oversize HotPacs,

2. Three Standard size HotPacs

3. Neck Contour HotPac.

4. Forceps and Tongue

14. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; PHYSICAL MEDICINE & REHABILITATION

Exercise Table

Exercise Table :

1. Made up of solid wood, Should have 4 legs

2. Plinth Size: High and Low

3. Top 19 mm thickness ply.

4. 4 inch cushioned with rexin cover legs cross section 8 x 10 cm.

Department ; PHYSICAL MEDICINE & REHABILITATION

Static Cycle

Static Cycle :

1. Accurate instrumentation to measure heart rate, speed, distance, time and energy.

2. Should have digital display showing speed, time, distance and calorie used.

3. Body made up of stainless steel stationary exercise Cycle.

4. Should have comfortable saddle and foam fitted handle.

5. Should have adjustable design to fit all heights and weights.

6. Should have maximum user weight of 100kg.

7. Buidin hand grip pulse sensor.

8. Should have Resistance system with manual control.

9. Should have large adjustable softer HR seat

10. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department : PHYSIOLOGY Polygraph (Digital)

|  | The system should be able to record and analyse. |  |
| --- | --- | --- |
|  | Pulse PPG, blood pressure, ECG, Pulse Transit time, Heart Rate variability (HRV) |  |
|  | ECG Multi leads configurations for real time cardiac axis & vector analysis. |  |
|  | Digital microphone for heart & Lung sounds studies. |  |
|  | Pulse Wave velocity, Deep breathing test, Valsalva Menuever. |  |
|  | SPO2 Acceleration (3 Axis), Activity, GSR & Temperature, EEG, EMG NCV etc. |  |
|  | Dynamometer to study Hand Grip Test. |  |
|  | Four channel balance board for static posturogtraphy or body Sway studies. |  |
|  | Hardware Specifications: |  |
|  | Simultaneous recording for at least 12 and more parameters. |  |
|  | High sampling rate of 40 KHz or more. |  |
|  | Should have Fully Isolated stimulator for constant current 7 voltage stimulation output rang (0-20 mA) with pulse duration: 50-200 us. |  |
|  | Must have preferred parameters for wireless monitoring ECG, SPO2, PPG, Registration, GSR and Temp etc. |  |
|  | The software should have step by step instructions, protocol and experimental design for performing various experiments in physiology teaching applications. Also should have sample data for animal experiments for demonstrating to the students. |  |
|  | Must be supplies with ECG, EMG, EEG, NCV, Isometric Force and reflective drop counter transducer for animal studies. |  |
|  | Facility for automatic analysis in while data recording and after recording for ECG, Heart Rate Variability, Blood Pressure, Peak analysis, spectrum analysis etc. |  |
|  | Should have option for Mathematical function and statistical analysis and export options to other software like MATLAB, Excel, Quick Time, Wav, Text etc for desired interpretation of the data. |  |
|  | The software should provide an easy file sharing option to a distant user with-out involving any cost with a 10 year of free updates and upgrade. |  |
|  | The system should have upgrade Options for: |  |
|  | o | Non Invasive Beat to Beat Blood Pressure monitor for tilt table testing. |
|  | o | Non Invasive had held Tonometer for vascular function testing. |
|  | o | Non Invasive skin blood flowmeter. |
|  | Outside Proper demonstration to be carried out before finalizing. |  |
|  | Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO. |  |

Department : PHYSIOLOGY Bicycle Ergometer

| a. Bicycle ergometer Work Load Range: 10 – 600Watt Independent Rpm: 20 – 130 rpm. |  |
| --- | --- |
| b. Work Load range Dependent RPM Lap Time: 9h 59min 59sec |  |
| c. | Load Steps : 1 – 99 Watt |
| d. Time Steps : 1 – 99 Min |  |
| e. Load Programs: Manual 1 Load/Step PWC 170, 150, 130 HR Steady State Safety Protection. |  |
| f. | Automatic Load: Heart rate |
| g. | Reduction: Blood pressure Blood Pressure D |
| h. Acoustic signal: ECG Alarm Signal Manual : By Stop |  |
| i. | Button Check: Calibration Dynamic Static Adjustments (Height): 83cm – 110cm Saddle: 68cm – 108mm Handle bars: free rotation Accuracy of load: 2% or 3 Watt Input/output |
| j. | External Control: RS – 232 |
| k. | Inputs: Analogue control Start signal Heart rate Blood pressure s Blood pressure D ECG Alarm Outputs: work load Pedaling Speed Step Marker |
| l. | Power supply: 220v – 240V, |
| m. Weight: approx 46Kg |  |
| n. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO. |  |

Department : PHYSIOLOGY

Ophthalmoscope

Ophthalmoscope (Direct)

1. Available with LED/Halogen light source.

2. Magnification up to x15 from direct vision to maximum magnification.

3. Red-free, blue and polarization filters and Anti-reflection lens.

4. Should have small and large spot sizes, fixation targets, slit aperture, hemispot and cobalt blue filter.

5. Should be rechargeable battery with Charger / battery/ mains operated.

6. At least 3 apertures and fixation star.

7. Range of lenses not smaller than - 30D to +20D with steps not greater than 1D.

8. Dust free sealed optics and aspherical optical system.

9. Internal batteries, rechargeable preferred compatible with both 2.5 V and 3.5 V batteries or handles provided; Led display indicating the charging status.

10. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department : PHYSIOLOGY

Otoscope

5. 2.5 V (White Led Light) Pocket Scope

6. Auriscope With Aa Battery Handle

7. Wide - Angle Viewing Lens Allows For Instrumentation Under Magnification.

8. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department : PHYSIOLOGY

Multi channel Physiograph 3 Channels,

complete with accessories

Technical Specifications of Physiograph Machine with transducers & other relevant access.

1. Should be supplies as a standalone recorder to monitor & record the data on three channels colored TFT Monitor.

2. System should have built in stimulator.

3. System should have facility to store the recording and review the recorded data on inbuilt TFT screen without need of any computer.

4. A software should be provided free of cost along with the system to review an printing the recorded data from PC whenever required, (No need to quote cost of computer)

5. Should be supplied with Forcer Transducer, Isotonic Transducer & Volume transduced.

6. For human experiment, transducers / electrodes to perform following tests i.e. ECG, EEG, EMG, GSR, Pulse, Respiration, Hand dynamometer Phonocardiogram & Temperature should be supplied.

7. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO

Department : PHYSIOLOGY

Centrifuge, high speed with tachometer

Centrifuge, high speed with tachometer

1. Should have a maximum speed of 5000 RPM with steeples regulator

2. Should be supplied with safety lid and lock. (Brushless Motor)

3. Should have digital speedometer and timer.

4. Should have imbalance detector and automatic cut off.

5.   Standard mark on quality and safety is Compulsory.

6.   Should work on a 220VAC +/- 10 %, 50Hz AC Supply.

7.   Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; PSYCHIATRY

Electro Convulsive Therapy (E.C.T.) Machine Preferably with ECG & EEG Monitoring

Electro Convulsive Therapy (E.C.T.) Machine Preferably with ECG & EEG Monitoring :

1. Should deliver current of bi-directional square wave ultra-brief pulse of 0.3ms.

2. Should have facility of Right Unilateral (0.3ms) ultra-brief stimulation.

3. Should have 4 channels EEG, 1 channel ECG and optical motion sensor to monitor the movement during seizure for providing assessing seizure efficacy.

4. Should have four different stimulus parameter knobs to vary pulse width, frequency, duration & current.

5. Should have facility to monitor real time dynamic impedance during procedure & also static impedance.

6. Should have LCD with Touch screen display with impedance display.

7. Should have protection against paddle to paddle short circuit or open circuit conditions.

8. The Stimulus Control push button to be hinged for prevention of accidental delivery of stimulus and should have visual indication of the status of Stimulus enable, Delivery or fault

9. Should have stimulus current from 500-900mA, Frequency 20-120Hz, Pulse width – 0.3-1msec, Stimulus duration of 0.5-8sec.

10. Minimum Power: 0.3 Joules for 220 ohm patient impedance

11. Maximum Power: 200 Joules for 220 ohm patient impedance.

12. Charge: 4.0 – 1152m Coulombs.

13. Should have 2 channel thermal chart recorder with gain knobs for higher resolution printing.

14. Should have facility to connect system to any External PC and will be provided with monitoring software to view physiological monitoring. The traces should be available in real time throughout the treatment. The data can be stored with all the treatment parameters on the PC or can be converted in to text format. Software should have features:

15. View up to six traces of real-time monitoring on a PC monitor: 4 EEG, 1 ECG, 1 Optical Motion Sensor

16. Equipment Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; PSYCHIATRY

ECT machine without monitor

ECT Machine with Monitor :

1. Should have constant current bi-directional square wave Brief Pulses.

2. Parameter display on LCD as well as on monitor screen.

3. Should be able to deliver ECT from voltage 50-400 volts.

4. Should have protection against paddle –to- paddles short circuit or open circuit conditions.

5. Should have stimulus current 500-800 MA Frequency 20-120 Hz, Pulse Width 0.5-1.5 m.sec stimulation duration of 0.1 - 5.9 Sec. - Minimum Power – 0.6 Joules for 220-ohm Patient Impedance. – Maximum Power 205.8 Joules for 220- ohm Patient Impedance. - Charge : 5.0 – 1152 milicimlumba in both manual and timer mode.

6. Should be provided with optical motion sensor for monitoring motor movement during seizure.

7. Should have facility upgrade to 24-32 Channels Digital EEG Systems.

8. Should have provision of monitoring EEG, EMG, ECG, Stimulus and Movement with optical motion sensor for providing assessing seizures efficacy.

9. Should be provided with monitoring software to view physiological monitoring of upto 4 traces. The trace should be available in real time through our the treatment.

10. Should have facility for the data to be stored with all the treatment parameter on the PC Hard disc or can be transfer to CD.

Should having a comprehensive database to store the complete patient information and can be configured according to user needs.

Output should displays in joules as well as in millicoulombs.

ECT module can be used in stand – alone mode also. COMPUTER SECTION (HARDWARE)

System should have following accessories:- Spring loaded ECT Headband, Instruction Manual, EEG, EMG & ECG Electrodes, OSM sensor, Bite Block, Earthingwire, Conductive Jelly, Paper Rim(A4 Size), Rubber Strap Electrodes.

Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; PSYCHIATRY

Lithium Analyzer

Lithium Analyzer :

1. Assay type 1-Point

2. Reaction time / Assay points 10 / 7

3. Wavelength (sub/main) 480/505 nm

4. Unit mmol/L (mg/dL)

5. Traceability: The lithium calibrator C.f.a.s. is traceable against AAS

6. Calibration mode Linear

7. Measuring range: 0.05-3.00 mmol/L (0.03-2.08 mg/dL)

8. Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; RESPIRATORY MEDICINE

Rigid Bronchoscope

Rigid Bronchoscope :

Digital Single Chip Endo vision Camera: one

(compatible with telescopes of Bronchoscope, Oesophagoscope , and Cystoscope

Color system PAL / NTSC , power supply: 100 – 240 VAC , 50/60 Hz compatible with Pal, NTSC Automatic white balance with control on base unit and also on camera.

Integrated zoom lens for manual and automatic control for exposure of fog.

Compatible with VHS and Comp and DVI.

Minimum sensitivity 3 lux.

Instrument coupling for all rigid endoscope.

Long camera cable 300 cm.

Two pre set function keys on camera for control of camera function , printer, computer , VCR and other peripherals

Medical Color Monitor 19” (fully flat screen): one

Special 19 “monitor, medical grade, Trintron colour tube, PAL , NTSC & SECAM color system. For Bronchoscopy, Oesophagoscopy and Cystoscopy

Telescope

Straight forward telescope , 0 degree , diameter 2.9 mm , length 36 cm , autoclavable, fiber optic light transmission incorporated( one)

Straight forward telescope 0 degree , diameter 2.9 mm , length 30 cm , autoclavable, fiber optic light transmission incorporated ( one)

Straight forward telescope 0 degree , diameter 2.7mm , length 18cm autoclavable, fiber optic light transmission incorporated ( one)

Bronchoscope Sheath Bronchoscope sheath , size 4.5 outer diameter 7.3 mm ,inner diameter 6.6 mm , length :30 cm ( one

Bronchoscope sheath , size 4 outer diameter 6.7mm ,inner diameter 6mm , length :30 cm ( one )

Bronchoscope sheath , size 3.7 outer diameter 6.4 mm ,inner diameter 5.7 mm , length :30 cm ( one

Bronchoscope sheath , size 3.5 outer diameter 6.4 mm ,inner diameter 5.7 mm , length :30 cm ( one

Bronchoscope sheath , size 4.0 outer diameter 6.7 mm ,inner diameter 6.4 mm , length :26 cm ( one

Bronchoscope sheath, size 3.7outer diameter 6.4 mm ,inner diameter 6.4 mm , length :26 cm ( one

Bronchoscope sheath , size 3.5 outer diameter 5.7 mm ,inner diameter 5.0 mm , length :26 cm ( one

Bronchoscope sheath , size 3.0 outer diameter 5.0 mm ,inner diameter 4.3 mm , length :26 cm ( one

Bronchoscope sheath , size 3.5 outer diameter 5.7 mm ,inner diameter 5.0 mm , length :18.5cm( one )

Bronchoscope sheath , size 3.0outer diameter 5.0 mm ,inner diameter 4.3 mm , length :18.5m ( one

Bronchoscope sheath , size 2.5 outer diameter 4.2 mm ,inner diameter 3.5 mm , length :18.5m ( one )

Optical forceps (compatible with above mentioned telescopes and bronchoscope sheaths

Optical alligator forceps , 2x2 teeth , with spring action handle , for controlled removal of flat foreign bodies ( such as coins) ;(two)

Optical forceps, with spring action handle for controlled removal of soft foreign bodies ( such as peanuts) : ( two)

Optical alligator forceps, with forced controlled handle , for removal of hard foreign bodies ( two

Forceps (one each)

Forceps alligator, grasping , double action jaw , sheath diameter 1.5 mm , working length: 35 cm

Forceps alligator, grasping, double action jaw, sheath diameter 1.5 mm , working length: 35 cm , pointed serrated for coins and foreign bodies.

Forceps alligator , grasping double action jaw, sheath diameter 1.5 mm , working length 35 cm , pointed for peanuts and soft foreign bodies.

Biopsy Forceps 35 Cm

Sponge holder:

Spring handle, working length 35 cm ( one)

Rigid suction tube with rubber tip( one each)

Rigid suction tube, straight length 35 cm, diameter: 3 mm

Rigid suction tube, straight length 25 cm. diameter: 3 mm

Foreign body basket with rigid handle, working length 35 cm

Telescope bridge for fixed position between telescope and bronchoscope.

Compatible with above mentioned telescopes and bronchoscope sheaths ( Three)

Prismatic light deflection with connection for fiber optic light cable, autoclave (Two)

Glass Window Plug (Two)

Rubber Telescope Guide for use with telescopes or optical forceps (Two)

FLUVOG Adaptor with sliding glass window plug, sealing cap, notched lens and keyhole opening , moveable (Two)

Injection canula with Leur lock outer diameter – 3.5 mm, for use with bronchoscope tubes (for positive pressure assisted ventilation (Two)

Injection canula with Leur lock outer diameter – 2.7 mm, for use with bronchoscope tubes (for positive pressure assisted ventilation (Two)

Instrument guide for suction catheter (Two)

Adaptor from bronchoscope to any type of pediatric respiration equipment (Two)

Sealing plug for ventilation attachment of bronchoscopes (Two)

Adjustable magnifier, swing – away type, autoclave (One)

Fiber optic light cable size 3.5mm, length 180 cm (Two)

Cold light fountain twin bulb Halogen 150 watts Power supply 220 VAC, 50 Hz. (One)

Facility to capture images and record videos of the procedure should be available. Storage facility should be adequate for at least 100 videos.

Should be USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO.

Department ; SKILLS LABS

Basic Life Support (BLS), CPR (Cardiopulmonary Resuscitation) Mannequin

Basic Life Support (BLS), CPR (Cardiopulmonary Resuscitation) Mannequin :

1. The material of mannequin should be of polyvinyl or silicone rubber, free from any hazardous materials.

2. The texture of the mannequin should be as close to the feel of the baby/adult skin as relevant.

3. The Internal parts of the mannequin must be realistically sculpted, anatomically accurate and feel must be smooth/resilient/bony as relevent and suitable for simulation.

4. It should have features to demonstrate opening of airway, head tilt/chin tilt and jaw thrust techniques.

5. Adult CPR Mannequin should have disposable airways.

6. Adult CPR Mannequins should have removable, reusable faces.

7. Adult CPR mannequin should have an indicator which confirms correct chest compression technique.

8. It should have compression spring for consistent resistance.

Department ; SKILLS LABS

Various Types of Injections- Subcutaneous, Intra-muscular, Intra-venous

Various Types of Injections- Subcutaneous, Intra-muscular, Intra-venous :

1. The material of mannequin should be of polyvinyl or silicone rubber, free from any hazardous materials.

2. The texture of the mannequin should be as close to the feel of the baby/adult skin as relevant.

3. The Internal parts of the mannequin must be realistically sculpted, anatomically accurate and feel must be smooth/resilient/bony as relevent and suitable for simulation.

4. Injection trainer should provide clear differenation of interdermal, intra muscular , and subcutaneous layers

Department ; SKILLS LABS

Gynecological Examination Mode / Mannequin including ICUD (Intra

Uterine Contraceptive Device) Training Model

Gynecological Examination Mode / Mannequin including ICUD (Intra Uterine Contraceptive Device) Training Model :

1. Should be made of washable an non -tearable material

2. It should support training in postpartum IUD vaginal examinations, IUD insertion, and removal, and for inspecting anteverted and retroverted positions of the uterus, family planning , menstrual hygiene counseling and uterine balloon tamponade insertions and other postpartum uterus interventions.

3. . Should have realistic anatomy representing the uterus after birth ,between pregnancies and immediately postabortion.

4. Should be used on a tabletop or attached to the manikin for practicing and demonstartion of IUCD insertion

5. Should have uterus, ovaries and fimbria.

6. Should have closed and transparent view.

Department ; SKILLS LABS

Neonatal & Pediatric Resuscitation Mannequins

Neonatal & Pediatric Resuscitation Mannequins :

1. The material of mannequin should be of polyvinyl and silicone rubber, free from any hazardous material.

2. The texture of the mannequin should be close to the feel of the baby/adult skin as relevant.

3. The Internal parts of the mannequin must be realistically sculpted, anatomically accurate and feel must be smooth/resilient/bony as relevent and suitable for simulation.

4. Newborn mannequin should have features for training essential newborn care (ENBC) and newborn resuscitation.

5. Newborn Mannequin should facilitate effective bag and mask ventilation, chest must rise only with correct technique.

6. The newborn mannequin should include the following: Squeeze bulbs for simulation of cord pulsation, spontaneous breathing and auscultation of heart sound.

7. The new born mannequin should demonstrate clearing of airways, perform suction; monitoring of ventilation and pulsation.

Department ; SKILLS LABS

Whole Body Mannequins

| Whole Body Mannequins : |  |
| --- | --- |
| 1. The Manikin should be a full body manikin designed to run advance life support course. |  |
| 2. The manikin should allow to teach patient assessment and physical examination skills including |  |
| a. Head to toe examination |  |
| b. Pupillary reaction with multiple scenarios |  |
| c. | Tongue swelling (angioedema) |
| d. Carotid pulse |  |
| e. Visible chest rise |  |
| f. | Lung auscultation including normal and abnormal breath sounds |
| g. | Pneumothorax |
| h. Heart sound including normal and abnormal |  |
| i. | Blood pressure measurement |
| j. | ECG monitoring |
| 3. The manikin should allow to teach airway technique including Headtilt chin lift, C-E technique for Mask holding, jaw thrust. All basic, supra glottis, advanced airways and surgical airway procedures |  |
| 4. The manikin should have CPR feedback as per 2015 guidelines. The manikin should have feedback on depth chest compression, chest recoil, rate of compression, interruption time, ventilation volume and should also provide score on over all CPR skills. |  |
| 5. Should allow to perform needle decompression and chest drain Should have IV arm with veins that IV accesses as a part of drug therapy. |  |
| 6. The manikin should generate all ACLS normal and abnormal ECG rhythms. |  |
| 7. The manikin should be able to receive real Defib shock of 360/200 Joules. |  |
| 8. The manikin should have automatic carotid, brachial and radial pulse, the pulse should have synchronized with ECG and blood pressure. |  |
| 9. The manikin should have facility to have vocal sounds for patient to participant communication. |  |
| 10. The manikin should have visible chest rise coordinating with the set respiratory rate. The lung sound, chest rise, and RR should coordinate with each other. |  |
| 11. The manikin should be controlled by a wireless, touch screen control unit. |  |

Department ; SKILLS LABS

Trauma Mannequin

Trauma Mannequin :

1. Mounted to a base but can be easily transfer to adults mannequin for use in full body trauma scenarios.

2. Recognition and assessment of the following

1. Palpable fractures.

2. Open depressed skull fracture.

3. Le Fort I & III

4. Nasal fracture,

5. Mandibular fracture ( Left)

6. Fracture of C-6 Vertebra.

7. Unequal pupils,

8. Hemotympanum.

SECTION-VI (SAMPLE FORMS)

1.BID FORM

(Note: -This Annexure must be sworn before First Class Magistrate/Notary)

Date: …………….. 2024                       [insert:  date of bid]

[Purchaser specify: “IFB No.:  BMSIC/2024-25/ME-382”]

[Insert: Procurement and Contracting of Medical Equipment for Medical Colleges and Hospitals of Bihar]

To:

Managing Director,

Bihar Medical Services and Medical Services Corporation,

3rd Floor, Swasthya Bhawan, Behind IGIMS, Sheikhpura, Adjacent to State Health Society, Patna- 800014, Bihar

Dear Sir or Madam:

Having examined the Bidding Documents, including Amendment and all corrigendum, the receipt of which is hereby acknowledged, we, the undersigned, offer to supply and deliver the Goods under the above-named Contract in full conformity with the said Bidding Documents for the sum of Rs. 11,800/-(hereinafter called “the Total Bid Price”) or such other sums as may be determined in accordance with the terms and conditions of the Contract. The above amounts are in accordance with the Price Schedules attached herewith and are made part of this bid.

We undertake, if our bid is accepted, to deliver the Goods in accordance with the delivery schedule specified in the Schedule of Requirements.

If our bid is accepted, we undertake to provide an advance payment security and a performance security in the form, in the amounts, and within the times specified in the Bidding Documents.

We agree to abide by this bid, for the Bid Validity Period specified in Clause 18 of the ITB and it shall remain binding upon us and may be accepted by you at any time before the expiration of that period.

Until the formal final Contract is prepared and executed between us, this bid, together with your written acceptance of the bid and your notification of award, shall constitute a binding Contract between us. We understand that you are not bound to accept the lowest or any bid you may receive.

We undertake that, in competing for (and, if the award is made to us, in executing) the above contract, we

will strictly observe the laws against fraud and corruption in force in India namely “Prevention of Corruption Act

1988”.

We confirm that we comply with the eligibility requirements as per ITB Clause 3 of the bidding documents.

We understand that you are not bound to accept the lowest or any bid you may receive. Dated this [insert: number] day of [insert: month], [insert: year].

Signed:   	 Date:

In the capacity of [insert: title or position]

Duly authorized to sign this bid for and on behalf of [insert: name of Bidder]

| ## 2.PRICE SCHEDULE |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| ##### 1 | ##### 2 | ##### 3 | ##### 4 | ##### 5 |  |  |  |  |  | ##### 6 | ##### 7 | ##### 8 |
| ##### S. N. | ##### Item Description | ##### Country of origin | ##### Quantity | ##### Ex-factory Ex- warehouse ex- Showroom off-shelf (A) | ##### Excise duty if any (B) | ##### Packing & Forwarding (C) | ##### Inland transport, Insurance & Incidental costs incidental to delivery (D) | ##### Incidental services as listed in GCC (E) | ##### Customs Duty (F) | ##### Unit Price A+B+C+D+E+F | ##### Total Price per schedule for delivery at final destination (4X6) | ##### GST payable if contract is awarded |

##### Unit Price (6) ( Rs. In words)

AMC Charges (Labour only)

| Equipment name | AMC CHARGES |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  | 4TH YEAR | 5TH YEAR | 6TH YEAR | 7TH YEAR | 8TH YEAR | 9TH YEAR | 10TH YEAR |
| TOTAL |  |  |  |  |  |  |  |

CMC CHARGES

| Equipment name | CMC CHARGES |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  | 4TH YEAR | 5TH YEAR | 6TH YEAR | 7TH YEAR | 8TH YEAR | 9TH YEAR | 10TH YEAR |
| TOTAL |  |  |  |  |  |  |  |

Note:

                - In case id discrepancy between unit price & total price Unit price shall prevail.

Place												Signature of Bidder/Authorized Signatory………………………

Date												Name          ………………………

(Should be submitted in format as available in e-mode only)

FORM – 3 FORM OF CONTRACT AGREEMENT

THIS CONTRACT AGREEMENT is made the               day of     	 [month and year purchase] and between the Bihar Medical Services And Infrastructure Corporation Limited, Patna [Name of Purchaser] on behalf of Governor of Bihar (hereinafter referred to as the ‘Purchaser’) and

[ Name of Supplier], having            its           principal           place           of           business           at

_                                                                                   [   address   of   Supplier ]

(hereinafter referred to as the “Supplier) on the other part.

WHEREAS the Purchaser invited bids for certain goods and ancillary services, viz., [insert: brief description of goods and services] and has accepted a bid by the Supplier for the supply of those goods and services in the sum of [insert: contract price in words and figures] (hereinafter called “the Contract Price”).

NOW THIS AGREEMENT WITNESSETH AS FOLLOWS:

1. In this Agreement words and expressions shall have the same meanings as are respectively assigned to them in the Conditions of Contract referred to.

2. The following documents shall constitute the Contract between the Purchaser and the Supplier, and each shall be read and construed as an integral part of the Contract:

(a)     This Contract Agreement

(b)     General Conditions of Contract. (c)     Special Conditions of Contract

(d)     Technical    Requirements    (including    Functional

Requirements and Implementation Schedule).

(e)     The  Supplier’s  original  Techno-commercial  and

Price bid

(f)     The Schedule of Requirements.

(g)     The Purchaser’s Notification of Award

(h)     [Add here: any other documents]

3.	In consideration of the payments to be made by the Purchaser to the Supplier as hereinafter mentioned, the Supplier hereby covenants with the Purchaser to provide the Goods and Services and to remedy defects therein in conformity in all respects with the provisions of the Contract.

4.	The Purchaser hereby covenants to pay the Supplier in consideration of the provision of the Goods and Services and the remedying of defects therein, the Contract Price or such other sum as may become payable under the provisions of the Contract at the times and in the manner prescribed by the Contract.

Brief  particulars  of  the  goods  and  services  which  shall  be supplied/provided by the Supplier are as under:

| SL | Brief Description of goods | Unit Price | Quantity to be supplied | Total price |
| --- | --- | --- | --- | --- |

TOTAL VALUE: Delivery Schedule:

For and on behalf of the Purchaser

Signed:

in the capacity of [ insert:   title or other appropriate designation ]

in                                          the                                          presence of

For and on behalf of the Supplier

Signed:

in the capacity of [ insert:   title or other appropriate designation ]

in                           the                           presence                           of

CONTRACT AGREEMENT

dated the [ insert:  number ] day of [ insert:  month ],

[ insert: year ]

BETWEEN

Bihar  Medical  Services  And  Infrastructure  Corporation

Limited, “the Purchaser”

and

[ insert:  name of Supplier ], “the Supplier”

To:

Managing Director,

4. PERFORMANCE SECURITY BANK GUARANTEE

(Unconditional)

Date:  [insert:  date]

IFB:  [insert:  name or number of IFB]

Contract:  [insert:  name or number of Contract]

Bihar Medical Services And Infrastructure Corporation Limited, Patna

Dear Sir or Madam:

We refer to the Contract Agreement (“the Contract” ) signed on [insert:date] between you and

[insert:name of Supplier] (“the Supplier”) concerning the supply and deliver y of [ insert:abrief description of the Goods]. By this letter we, the undersigned,[insert: name of bank], a bank(or company) organized under the laws of insert:country of bank] and having its registered/principal office at [insert:address of bank], (hereinafter, “the Bank”) do hereby jointly and s everally with the Supplier irrevocably guarantee payment owed to you by the Supplier, pursuant to the Contract, up to the sum of [ insert:amount in numbers and words]. This guarantee shall be reduced or expire as provided for by GCC clause 5.

We undertake to make payment under this Letter of Guarantee upon receipt by us of your first written demand signed by your duly authorized officer declaring the Supplier to be in default under the Contract and without cavil or argument any sum or sums within the above-named limits, without your need to prove or show grounds or reasons for your demand and without the right of the Supplier to dispute or question such demand. Our liability under this Letter of Guarantee shall be to pay to you whichever is the lesser of the sum so requested or the amount then guaranteed under this Letter in respect of any demand duly made under this Letter prior to expiry of this Letter of Guarantee, without being entitled to inquire whether or not this payment is lawfully demanded.

This Letter of Guarantee shall be valid from the date of issue until the date of expiration of the guarantee, as governed by the Contract. Except for the documents herein specified, no other documents or other action shall be required, notwithstanding any applicable law or regulation. Our liability under this Letter of Guarantee shall become null and void immediately upon its expiry, whether it is returned or not, and no claim may be made under this Letter after such expiry or after the aggregate of the sums paid by us to you shall equal the sums guaranteed under this Letter, whichever is the earlier. All notices to be given under this Letter shall be given by registered (airmail) post to the addressee at the address herein set out or as otherwise advised by and between the parties hereto.

We hereby agree that any part of the Contract may be amended, renewed, extended, modified, compromised, released, or discharged by mutual agreement between you and the Supplier, and this security may be exchanged or surrendered without in any way impairing or affecting our liabilities hereunder without notice to us and without the necessity for any additional endorsement, consent, or guarantee by us, provided, however, that the sum guaranteed shall not be increased or decreased.

No action, event, or condition that by any applicable law should operate to discharge us from liability hereunder shall have any effect, and we hereby waive any right we may have to apply such law, so that in all respects our liability hereunder shall be irrevocable and, except as stated herein, unconditional in all respects.

For and on behalf of the Bank

Signed:

Date:

in the capacity of:  [ insert: title or other appropriate designation] Common Seal of the Bank

5. MANUFACTURER’S AUTHORIZATION FORM

(Manufacturer’s / Producer’s letterhead)

To:

Managing Director,

Bihar Medical Services and Infrastructure Corporation Limited,

Patna

WHEREAS [name of the manufacturer / producer] (hereinafter referred to as “we” or “us”) who are established and reputed manufacturers / producers of [name and/or description of the Goods requiring this authorization ](hereinafter referred to as  “Goods”) having manufacturing / production facilities  at[ insert: address of factory ]do hereby authorize[ name and address of Bidder ](hereinafter, the “Bidder”) to submit bid, and sign the Contract with you against IFB [title and reference number of the Invitation for Bids]including the above Goods produced by us.

We hereby extend our full guarantee and warranty for the above specified Goods against these Bidding Documents and also confirm full functionality of the said equipment during the entire post warranty Comprehensive Maintenance Contract Period as agreed.

For and on behalf of the Manufacturer or Producer

Signed:

Date:

In the capacity of [title, position, or other appropriate designation] and duly authorize to sign this

Authorization on behalf of [name of manufacturer or producer]

Note: This letter of authority should be on the letter head of the manufacturers and should be signed by a person competent and having the power of attorney to legally bind the manufacturer. This should be included by the bidder in it’s bid.

6. PROFORMA FOR PERFORMANCE STATEMENT

Bid No:

Date of Opening:

Time:                            Hours

Name of the Firm:

| Order Placed By (Full address of Purchaser) | Order No and Date | Description of works /project | Date of supply of work/projects |  | Was supply of works/projects was Satisfactory? | (Attach a certificate from the Purchaser/consignee (as applicable) |
| --- | --- | --- | --- | --- | --- | --- |
|  |  |  | As per contract | Actual |  |  |
|  |  |  |  |  | Yes/No |  |

(Signature and seal of the Bidder/Authorised Signatory)

7. CONSIGNEE RECEIPT CERTIFICATE/ INSTALLATION REPORT/ CERTIFICATE

(To be given by consignee / end user of the project)

The following equipment has / have been received in good condition:

| Name of Project |  |
| --- | --- |
| Name of the Supplier / Manufacturer |  |
| Quantity supplied / cost of the project |  |
| Purchase Order reference no. |  |
| Detailed item list of the project |  |
| Place of the project executed |  |
| Name and Address of the Consignee along with tel. no. and fax no. |  |
| Date of receipt by the Consignee |  |
| Date of completion |  |
| Signature of Authorized Representative of Consignee with date |  |
| Name and designation of the authorized representative |  |
| Seal of the consignee |  |

Note: In case of Hospital the In charge of the hospital concerned would be treated as consignee. In case of office (other than hospital), the office in charge of the office would be treated as consignee.

(Hospital / Office In charge)                                                                        (User Department)

8. STATEMENT FOR TECHNICAL DEVIATION:

| Sr. No | Specifications desired by BMSICL | Bidders’ specifications | Bidders Deviation if Any |
| --- | --- | --- | --- |

(Signature of Bidder/ Authorized Signatory)

9. FORMAT FOR WARRANTY CERTIFICATE

(To be submitted on Firms Letterhead)

Warranty Certificate

Date:

We the Undersigned ……………………………………………………………………………...hereby guarantee satisfactory operation of …………………………………………………. supplied by us to you against your purchase order No………………………………………….for a period of………. calendar months from the date of commissioning and shall be responsible for failure of the equipment to conform to the standard of performance, proficiency, production and / or out-turn stipulated or implicit in the order and for any defects that may develop under proper use arising from the use of faulty materials, design or workmanship in the supply made and shall remedy such defects at our cost.

Consignee Name-

Serial number of Equipment-

For………………………………

Station: (Signature with Name and Designation) Date:

Company Seal

10. NON-CONVICTION DECLARATION (DULY NOTARIZED)

From:- M/s.....................

..........................

......................................................

...................................................... To

Managing Director

BMSICL, Patna

1.  I,

Son         /         Daughter         /         Wife         of

Shri

Proprietor/Director authorized signatory of the agency/Firm, mentioned above, is competent to sign this declaration and execute this tender document;

2. I have carefully read and understood all the terms and conditions of the tender and undertake to abide by them;

3. The information / documents furnished along with the above application are true and authentic to the best of my knowledge and belief. I / we, am / are well aware of the fact that furnishing of any false information / fabricated document would lead to rejection of my tender at any stage besides liabilities towards prosecution under appropriate law.

4. I/We further undertake that none of the Proprietor/Partners/Directors of the Agency/agency was or is Proprietor or Partner or Director of any Agency with whom the Government have banned /suspended business dealings. I/We further undertake to report to the Managing Director, BMSICL, Patna immediately after we are informed but in any case not later 15 days, if any Agency in which Proprietor/Partners/Directors are Proprietor or Partner or Director of such a Agency which is banned/suspended in future during the currency of the Contract with you.

Yours faithfully,

(Authorized Signatory/Signature of the Bidder)

Date:                                                                                                                          Name: Place:                                                                                                                         Designation Seal of the Agency                                                                                                    Address:

(Note: - This annexure must be sworn before First Class Magistrate/Notary)

11.BANK GUARANTEE FORM FOR EARNEST MONEY DEPOSIT (EMD)

Date: [ insert: date ]

IFB: [ insert: name and number of IFB ]

Contract: [ insert: name and number of Contract ]

To:

Managing Director,

Bihar Medical Services And Infrastructure Corporation Limited, Patna

WHEREAS [ insert: name of Bidder] (hereinafter called “the Bidder”) has submitted its bid dated [ insert:

date of bid] for the performance of the above-named Contract (hereinafter called “the Bid”)

KNOW ALL PERSONS by these present that WE [ insert name of bank] of [ insert: address of bank]

(hereinafter called “the Bank”) are bound unto [ insert name of Purchaser] (hereinafter called “the Purchaser”) in the sum of: [ insert: amount], for which payment well and truly to be made to the said Purchaser, the Bank binds itself, its successors and assigns by these presents.

Sealed with the Common Seal of the said Bank this [ insert: number] day of [ insert: month], [ insert: year]. THE CONDITIONS of this obligation are the following:

1. If, after the bid submission deadline

(a)     the Bidder withdraws or amends its tender or impairs or derogates from the tender in any respect before signing of the agreement or

(b)     does not accept the Purchaser’s corrections of arithmetic errors in accordance with the Instructions

to Bidders; or

(c)     within the period of validity of its tender or if it comes to notice that the information/ documents furnished in its tender is incorrect, false, misleading or forged or

(d)     engages in a corrupt practice, fraudulent practice, coercive practice, undesirable practice or restrictive practice

2.      If the Bidder, having been notified of the acceptance of its bid by the Purchaser during the period of bid

validity

(a)     fails or refuses to sign the Contract Agreement when required; or

(b)     fails or refuses to issue the performance security in accordance with the Instructions to Bidders.

We undertake to pay to the Purchaser up to the above amount upon receipt of its first written demand, without the Purchaser having to substantiate its demand, provided that in its demand the Purchaser will note that the amount claimed by it is due it, owing to the occurrence of any one of the two above-named CONDITIONS, and specifying the occurred condition or conditions.

This guarantee will remain in full force up to and including [ insert:the date that is 30 days after the period of bid validity ].

For and on behalf of the Bank

Signed:

Date:

in the capacity of: [ insert: title or other appropriate designation]

12. POWER OF ATTORNEY (Notarized)

I/ We…………………………………………………. (name and address of the registered office) do hereby constitute, appoint and authorise Sri/Smt ………………………………………………..…(name and address) who is presently employed with us and holding the position of ……………………………………………. As our attorney, to act and sign on my/our behalf to participate in the tender no…………………………………… for

…………………………………… (Project Name).

I/ We hereby also undertake that I/we will be responsible for all action of Sri/Smt.………………………………………... undertaken by him/her during the tender process and thereafter on award of the contract. His / her signature is attested below

Dated this the …. day of ... 2024_

Accepted _                                                                                                                 For

(Signature) (Name, Title and Address of the Attorney)                                 (Name, Designation and Address)

Date: _

| 13. CHECK LIST |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Name of the Tenderer |  |  |  |  |  |  |  |  |  |
| SL. No. | Item |  |  |  |  |  |  | Whether | Page No. |
|  |  |  |  |  |  |  |  | Included |  |
|  |  |  |  |  |  |  |  | Yes/No |  |
| A. Tender Fee, EMD |  |  |  |  |  |  |  |  |  |
| 1. | Tender Fee (Online mode) – Rs.11,800/- |  |  |  |  |  |  |  |  |
| 2. | EMD (Only offline Mode in the form of Bank Guarantee as per annexure-11). |  |  |  |  |  |  |  |  |
| B. Check list & Registration. |  |  |  |  |  |  |  |  |  |
| 1. | Make & Model Quoted items in the project. (Bidder can offer |  |  |  |  |  |  |  |  |
|  | Two alternate makes & Models.) |  |  |  |  |  |  |  |  |
| 2. |  | Document | claiming | the Registration for Trading/ |  |  |  |  |  |
|  |  | Manufacturing |  |  |  |  |  |  |  |
| 3. | Certificate of Incorporation and Articles of Memorandum of Association/Partnership Deed (As applicable) |  |  |  |  |  |  |  |  |
| 4. | Copy of certificate from Central Excise and Trades Tax/ Sales Tax |  |  |  |  |  |  |  |  |
| 5. | Copy of average Turn over certificate for last five consecutive Assessment years issued by Chartered Accountant. |  |  |  |  |  |  |  |  |
| 6. | Copy of certificate of Balance Sheet of last five consecutive Assessment years under the stamp and signature of Chartered Accountant. |  |  |  |  |  |  |  |  |
| 7. | Copy of certificate of P&L Statement of last five consecutive Assessment years under the stamp and signature of Chartered Accountant. |  |  |  |  |  |  |  |  |
| 8. | Copy of self-attested IT Returns for any three of last four consecutive Assessment years |  |  |  |  |  |  |  |  |
| 9. | Non-Conviction Declaration (Sworn before First Class Magistrate/Notary) as per Annexure 10 |  |  |  |  |  |  |  |  |
| 10. | Submission of Manufacturer's Authorization (if quoted by bidder other than manufacturer) as per Annexure 5 at the time |  |  |  |  |  |  |  |  |
|  | of supply is mandatory. |  |  |  |  |  |  |  |  |
| 11. | Bid Form (Sworn before First Class Magistrate/Notary) as per Annexure 1 |  |  |  |  |  |  |  |  |
| 12. | Supply/Purchase order issued by user institution to comply the criteria mentioned in ITB clause 16 |  |  |  |  |  |  |  |  |
| 13 | Performance Statement as per Annexure- 6 |  |  |  |  |  |  |  |  |
| 14 | Certificate from end user(s) indicating the Purchase order(s) as |  |  |  |  |  |  |  |  |
|  | submitted by the Bidder, date of supply, |  |  |  |  |  |  |  |  |
|  | installation/commissioning of the Equipment. |  |  |  |  |  |  |  |  |
| 15. | Technical Data Sheet/Brochure/Catalogue of the model of |  |  |  |  |  |  |  |  |
|  | quoted items. |  |  |  |  |  |  |  |  |
| 16. | Technical Deviation Compliance for every quoted item in project as per annexure 8 |  |  |  |  |  |  |  |  |
| 17. | Power of Attorney for the Signatory to the Bid as per annexure 12, duly notarized. |  |  |  |  |  |  |  |  |
| 18. | Quality Standard Certification (USFDA/ European CE (from 4 digit notified body)/ CE/ BIS / ISO) in accordance with technical specification in this |  |  |  |  |  |  |  |  |
|  | bid document must be furnished in technical bid documents |  |  |  |  |  |  |  |  |

| 19. | Notary attested declaration if exempted in EMD Fee, Technical Qualification as per Sankalp 675 (1), Dated 09/09/2013 of Govt. of Bihar as mentioned in special condition of contract. |  |  |
| --- | --- | --- | --- |
| 20. | Approval from Reserve Bank of India in case of Foreign Collaboration |  |  |
| 21. | IEC Certificate if the bidder is an importer of the quoted items in project. |  |  |