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Corrigendum-I

Bihar Medical Services and Infrastructure Corporation Limited (BMSICL) had invited E-Bids from the interested parties for Procurement, Rate contract, Supply, Installation of Medical Equipment vide Tender No.-BMSICL/2024-25/ME-396. During Pre-bid meeting various suggestions were received from different prospective bidders regarding amendment in technical specification of equipment which were discussed and deliberated on by the experts. On the basis of their recommendations certain amendments in the technical specification of the equipment have been made which are annexed as Annexure-I of this corrigendum. In order to facilitate maximum participation of bidders tender schedule is being revised as following:-
	Tender Reference No. 
	BMSICL/2024-25/ME-396

	Last date and time of submission of online bids
	17st February 2025 till 17:00 Hrs. 

	Last date and time of submission of original documents of EMD, Tender Fee and Document.
	18th February 2025 till 14:00 Hrs.  

	Date, Time and Place of opening of Technical Bid
	18th February 2025 (at 15:00 Hrs.) on the website of https://eproc2.bihar.gov in the office of BMSICL 

	Date and time of opening of financial Bids 
	To be announced later on https://eproc2.bihar.gov



	
                                                                                                                                           Sd/-
 GM (Procurement)











Annexure-I


	Name of Equipment:- Fibroscan Machine


	SI. No
	Technical Specification as per tender
	 Final Amendment

	1
	The equipment should be able to measure spleen and liver Fibrosis with dedicated Probes by generating vibrated transient elastic shear waves.

	No Change

	2
	It should be non-invasive painless, procedure and should provide immediate measurement of liver stiffness.

	No Change

	3
	 It should provide non-invasive assessment and quantification of fatty liver / steatosis.

	No Change

	4
	 It should have continuous controlled attenuation parameter CAP for steatosis quantification

	No Change

	5
	 The liver stiffness and steatosis measurements should be done at a frequency of 50Hz

	No Change

	6
	 It should be able to estimate spleen fibrosis /spleen stiffness measurements.

	No Change

	7
	 The spleen stiffness measurements should be done at a frequency of 100 Hz.

	No Change

	8
	Should be supplied with a single-element ultrasound transducer with a central frequency of 3.5 MHz

	No Change

	9
	 This transducer should be mounted on the axis of a mechanical vibrator.

	No Change

	10
	 The ultrasound should be able to generate a completely painless vibration (frequency 50- 100 Hz and amplitude 2 mm peak to peak)

	No Change

	11
	 The ultrasound transducer should be able to perform a series of ultrasound acquisitions with a repeal frequency of 4 KHz

	No Change

	12
	 Should be supplied with three probes with different frequencies for each category of patients adults, pediatrics and obese.

	No Change

	13
	 Adult probe 3.5 Mhz. with depth measurement 25-65 mm, transducer diameter less than 8mm.

	No Change

	14
	 Pediatric probe should be 5Mh, 15-50 mm, diameter less than 6mm.

	No Change

	15
	 Obese probe should be of frequency 2.5 Mhz. with depth measurement 35 to 75mm, and transducer less than 12mm

	No Change

	16
	The equipment should be a standalone unit with minimum 19 inches LED screen with wide viewing angle.

	No Change

	17
	 The system should be supported on inbuilt castors for all side stability.

	No Change

	18
	The equipment should have a minimum of 5 years warranty.
	The equipment should have 03 years warranty.

	19
	 5 Years CMC after completion of 5 years comprehensive warranty.
	07 Years CMC after completion of 3 years comprehensive warranty.

	20
	The equipment should be supplied with the following accessories:

	No Change

	(i)
	Two Examination Couch

	No Change

	(ii)
	Two Height adjustable Stools

	No Change

	(iii)
	One online UPS (1KVA) and at least three Extension Cords with surge protection

	No Change

	(iv)
	One Computer with 15” TFT Monitor and CPU and One Laser Colour Printer, with table and chair set for the system.

	No Change

	(v)
	 A minimum of 36 inch LED smart TV screen for visualization, teaching and training purpose.
	No Change

	(vi)
	 100 packets of photo glossy paper 90GSM (50 sheets each)
	No Change

	
	Added:-
	The firm has to provide training session to the concerned institute whenever required after the installation of machine
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