|  | Bihar Medical Services & Infrastructure Corporation Limited 2nd & 3rd Floor, Swasthya Bhawan, Behind IGIMS, Sheikhpura, Adjacent to State Health Society, Patna- 800014, Bihar, Phone/Fax: +91612 2283287, 91612 2283288 |
| --- | --- |

Corrigendum-I

Bihar Medical Services and Infrastructure Corporation Limited (BMSICL) had invited E-Bids from the interested parties for Procurement, Rate contract, Supply, Installation of Medical Equipment vide Tender No.-BMSICL/2024-25/ME-396. During Pre-bid meeting various suggestions were received from different prospective bidders regarding amendment in technical specification of equipment which were discussed and deliberated on by the experts. On the basis of their recommendations certain amendments in the technical specification of the equipment have been made which are annexed as Annexure-I of this corrigendum. In order to facilitate maximum participation of bidders tender schedule is being revised as following:-

| Tender Reference No. | BMSICL/2024-25/ME-396 |
| --- | --- |
| Last date and time of submission of online bids | 17st February 2025 till 17:00 Hrs. |
| Last date and time of submission of original documents of EMD, Tender Fee and Document. | 18th February 2025 till 14:00 Hrs. |
| Date, Time and Place of opening of Technical Bid | 18th February 2025 (at 15:00 Hrs.) on the website of https://eproc2.bihar.gov in the office of BMSICL |
| Date and time of opening of financial Bids | To be announced later on https://eproc2.bihar.gov |

Sd/-

GM (Procurement)

Annexure-I

| Name of Equipment:- Fibroscan Machine |  |  |
| --- | --- | --- |
| SI. No | Technical Specification as per tender | Final Amendment |
| 1 | The equipment should be able to measure spleen and liver Fibrosis with dedicated Probes by generating vibrated transient elastic shear waves. | No Change |
| 2 | It should be non-invasive painless, procedure and should provide immediate measurement of liver stiffness. | No Change |
| 3 | It should provide non-invasive assessment and quantification of fatty liver / steatosis. | No Change |
| 4 | It should have continuous controlled attenuation parameter CAP for steatosis quantification | No Change |
| 5 | The liver stiffness and steatosis measurements should be done at a frequency of 50Hz | No Change |
| 6 | It should be able to estimate spleen fibrosis /spleen stiffness measurements. | No Change |
| 7 | The spleen stiffness measurements should be done at a frequency of 100 Hz. | No Change |
| 8 | Should be supplied with a single-element ultrasound transducer with a central frequency of 3.5 MHz | No Change |
| 9 | This transducer should be mounted on the axis of a mechanical vibrator. | No Change |
| 10 | The ultrasound should be able to generate a completely painless vibration (frequency 50- 100 Hz and amplitude 2 mm peak to peak) | No Change |
| 11 | The ultrasound transducer should be able to perform a series of ultrasound acquisitions with a repeal frequency of 4 KHz | No Change |
| 12 | Should be supplied with three probes with different frequencies for each category of patients adults, pediatrics and obese. | No Change |
| 13 | Adult probe 3.5 Mhz. with depth measurement 25-65 mm, transducer diameter less than 8mm. | No Change |
| 14 | Pediatric probe should be 5Mh, 15-50 mm, diameter less than 6mm. | No Change |
| 15 | Obese probe should be of frequency 2.5 Mhz. with depth measurement 35 to 75mm, and transducer less than 12mm | No Change |
| 16 | The equipment should be a standalone unit with minimum 19 inches LED screen with wide viewing angle. | No Change |
| 17 | The system should be supported on inbuilt castors for all side stability. | No Change |
| 18 | The equipment should have a minimum of 5 years warranty. | The equipment should have 03 years warranty. |
| 19 | 5 Years CMC after completion of 5 years comprehensive warranty. | 07 Years CMC after completion of 3 years comprehensive warranty. |
| 20 | The equipment should be supplied with the following accessories: | No Change |
| (i) | Two Examination Couch | No Change |
| (ii) | Two Height adjustable Stools | No Change |
| (iii) | One online UPS (1KVA) and at least three Extension Cords with surge protection | No Change |
| (iv) | One Computer with 15” TFT Monitor and CPU and One Laser Colour Printer, with table and chair set for the system. | No Change |
| (v) | A minimum of 36 inch LED smart TV screen for visualization, teaching and training purpose. | No Change |
| (vi) | 100 packets of photo glossy paper 90GSM (50 sheets each) | No Change |
|  | Added:- | The firm has to provide training session to the concerned institute whenever required after the installation of machine |