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Corrigendum-I

Bihar Medical Services and Infrastructure Corporation Limited (BMSICL) had invited E-Bids from the interested parties for Procurement, Rate contract, Supply, Installation of Medical Equipment vide Tender No.-BMSICL/2025-26/ME-438. During Pre-bid meeting various suggestions were received from different prospective bidders regarding amendment in technical specification of equipment which were discussed and deliberated on by the experts. On the basis of their recommendations certain amendments in the technical specification of the equipment have been made which are annexed as Annexure-I of this corrigendum. In order to facilitate maximum participation of bidders tender schedule is being revised as following:-
	Tender Reference No. 
	BMSICL/2025-26/ME-438

	Last date and time of submission of online bids
	08th Sept.  2025 till 17:00 Hrs. 

	Last date and time of submission of original documents of EMD, Tender Fee and Document.
	09th Sept.  2025 till 14:00 Hrs.  

	Date, Time and Place of opening of Technical Bid
	[bookmark: _GoBack]09th Sept.  2025 (at 15:00 Hrs.) on the website of https://eproc2.bihar.gov in the office of BMSICL 

	Date and time of opening of financial Bids 
	To be announced later on https://eproc2.bihar.gov



                                                                                                                                         
 GM (Procurement)
 BMSICL













Annexure-I

	Name of Equipment:- Apheresis Machine

	SI. No
	Technical Specification as per tender
	 Final Amendment

	1
	Continuous/Intermittent flow Blood Cell separator.
	No Change

	2
	Dual/single Needle operation.
	No Change

	3
	Following protocols should be available
	No Change

	a
	Leucodepleted Platelets + Plasma
	No Change

	b
	 Therapeutic Plasma Exchange
	No Change

	c
	Single/double RBC collection or RBC Exchange
	No Change

	d
	Peripheral/continious  Blood Stem Cell Collections
	No Change

	e
	Granulocyte Collection.
	No Change

	f
	Leukoreduced platelet collection or plateleapheresis
	No Change

	4
	Extra Corporeal Volume less than 200ml
	Extra Corporeal Volume less than 250ml

	5
	 Inlet and return pressure monitor
	No Change

	6
	Built in air detectors to prevent air embolism
	No Change

	7
	Centrifuge leak detector, automatically detected before priming. 500ml ACD should be provided with plasma exchange / platelet apheresis kits
	No Change

	8
	Audio visual alarms.

	No Change

	9
	(i) USFDA 510(K) / European CE (issued by notified body) certified approved  (ii)  CDSCO (Central Drugs Standard Control Organization) certification 
	No Change

	10
	Power Supply 220VAC +/- 10 %, 50Hz.
	No Change

	11
	Should have inbuilt battery backup or on-line UPS of suitable capacity with 3 hours backup
	No Change

	12
	The rate for the following consumables should be provided in the price bid and the rates shall be freezed for three years from the date of price bid opening. The rate offered for consumables will be taken for L1 evaluation.                                                     
	No Change

	
	a. Platelet kit                                                                        
	No Change

	
	b. Plasma Exchange kit                                                                   
	No Change

	
	c. RBC Exchange kit                                                                                      
	No Change

	
	d. Stem cell kit
	No Change

	13
	20 nos of platelet apheresis kits / plasma exchange kits shall be supplied free of cost.
	10 nos. of platelet apheresis kits / plasma exchange kits shall be supplied free of cost.

	14
	Citrate control and monitoring facilit
	No Change

	15
	Platelet should be collected with plasma, no manual mixing needed
	No Change

	16
	Inbuilt quality control
	No Change

	17
	Programmable Haematocrit and Plasma Volume
	No Change

	18
	System should recover from exact point whenever paused
	No Change

	19
	ACD drip/flow monitoring with minimum ACD use
	No Change

	20
	Haemolysis monitoring for plasma line
	No Change

	21
	Yield of 3.3x1011 even at pre-count of 2.65 lakhs per micro liter
	No Change

	22
	 Inbuilt automatic cuff/ Pressure monitoring
	Deleted
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