|  | Bihar Medical Services & Infrastructure Corporation Limited 2nd & 3rd Floor, Swasthya Bhawan, Behind IGIMS, Sheikhpura, Adjacent to State Health Society, Patna- 800014, Bihar, Phone/Fax: +91612 2283287, 91612 2283288 |
| --- | --- |

Corrigendum-I

Bihar Medical Services and Infrastructure Corporation Limited (BMSICL) had invited E-Bids from the interested parties for Procurement, Rate contract, Supply, Installation of Medical Equipment vide Tender No.-BMSICL/2025-26/ME-438. During Pre-bid meeting various suggestions were received from different prospective bidders regarding amendment in technical specification of equipment which were discussed and deliberated on by the experts. On the basis of their recommendations certain amendments in the technical specification of the equipment have been made which are annexed as Annexure-I of this corrigendum. In order to facilitate maximum participation of bidders tender schedule is being revised as following:-

| Tender Reference No. | BMSICL/2025-26/ME-438 |
| --- | --- |
| Last date and time of submission of online bids | 08th Sept. 2025 till 17:00 Hrs. |
| Last date and time of submission of original documents of EMD, Tender Fee and Document. | 09th Sept. 2025 till 14:00 Hrs. |
| Date, Time and Place of opening of Technical Bid | 09th Sept. 2025 (at 15:00 Hrs.) on the website of https://eproc2.bihar.gov in the office of BMSICL |
| Date and time of opening of financial Bids | To be announced later on https://eproc2.bihar.gov |

GM (Procurement)

BMSICL

Annexure-I

| Name of Equipment:- Apheresis Machine |  |  |
| --- | --- | --- |
| SI. No | Technical Specification as per tender | Final Amendment |
| 1 | Continuous/Intermittent flow Blood Cell separator. | No Change |
| 2 | Dual/single Needle operation. | No Change |
| 3 | Following protocols should be available | No Change |
| a | Leucodepleted Platelets + Plasma | No Change |
| b | Therapeutic Plasma Exchange | No Change |
| c | Single/double RBC collection or RBC Exchange | No Change |
| d | Peripheral/continious Blood Stem Cell Collections | No Change |
| e | Granulocyte Collection. | No Change |
| f | Leukoreduced platelet collection or plateleapheresis | No Change |
| 4 | Extra Corporeal Volume less than 200ml | Extra Corporeal Volume less than 250ml |
| 5 | Inlet and return pressure monitor | No Change |
| 6 | Built in air detectors to prevent air embolism | No Change |
| 7 | Centrifuge leak detector, automatically detected before priming. 500ml ACD should be provided with plasma exchange / platelet apheresis kits | No Change |
| 8 | Audio visual alarms. | No Change |
| 9 | (i) USFDA 510(K) / European CE (issued by notified body) certified approved (ii) CDSCO (Central Drugs Standard Control Organization) certification | No Change |
| 10 | Power Supply 220VAC +/- 10 %, 50Hz. | No Change |
| 11 | Should have inbuilt battery backup or on-line UPS of suitable capacity with 3 hours backup | No Change |
| 12 | The rate for the following consumables should be provided in the price bid and the rates shall be freezed for three years from the date of price bid opening. The rate offered for consumables will be taken for L1 evaluation. | No Change |
|  | a. Platelet kit | No Change |
|  | b. Plasma Exchange kit | No Change |
|  | c. RBC Exchange kit | No Change |
|  | d. Stem cell kit | No Change |
| 13 | 20 nos of platelet apheresis kits / plasma exchange kits shall be supplied free of cost. | 10 nos. of platelet apheresis kits / plasma exchange kits shall be supplied free of cost. |
| 14 | Citrate control and monitoring facilit | No Change |
| 15 | Platelet should be collected with plasma, no manual mixing needed | No Change |
| 16 | Inbuilt quality control | No Change |
| 17 | Programmable Haematocrit and Plasma Volume | No Change |
| 18 | System should recover from exact point whenever paused | No Change |
| 19 | ACD drip/flow monitoring with minimum ACD use | No Change |
| 20 | Haemolysis monitoring for plasma line | No Change |
| 21 | Yield of 3.3x1011 even at pre-count of 2.65 lakhs per micro liter | No Change |
| 22 | Inbuilt automatic cuff/ Pressure monitoring | Deleted |