|  | Bihar Medical Services & Infrastructure Corporation Limited, 2nd & 3rd Floor, Swasthya Bhawan, Behind IGIMS, Sheikhpura, Adjacent to State Health Society, Patna 800023, Phone/Fax: +91612 2283287,+ 91612 2283288 |
| --- | --- |

Corrigendum-I

Bihar Medical Services and Infrastructure Corporation Limited (BMSICL) had invited E-Bids from the interested parties for Procurement, Rate contract, Supply, Installation of Medical Equipment vide Tender No.-BMSICL/2025-26/ME-434. During Pre-bid meeting various suggestions were received from different prospective bidders regarding amendment in technical specification of equipment which were discussed and deliberated on by the experts. On the basis of their recommendations certain amendments in the technical specification of the equipment have been made which are annexed as Annexure-I of this corrigendum.

Note:-

- Bidders are advised to refer to the Annexure-I of this corrigendum before submission of bid.

- Those who have submitted their bids are requested to re-submit their bids in accordance with this corrigendum.

Annexed:- as above

GM (Procurement)

BMSICL

| Annexure-I |  |  |
| --- | --- | --- |
| Name of Equipment :- MORTUARY CHAMBER (TWO BODY) |  |  |
| SI. No | Technical Specification as per tender | Final Amendment |
| 1 | A two-body unit exclusively for keeping decomposed bodies | No Change |
| 2 | Each unit shall have their individual, single cooling system to ensure the requisite temperature | No Change |
| 3 | All the units shall have a proper design ensuring best hygiene | No Change |
| 4 | All units shall be highly energy efficient | No Change |
| 5 | All units shall have a double wall with the gaps to be filled with high grade polyurethane insulation ensuring maximum thermal efficiency. | No Change |
| 6 | The puff used shall have a density of at least 40kg/m3 | No Change |
| 7 | The external and internal body of the mortuary units shall be corrosion free | No Change |
| 8 | The inner and outer walls shall be made of SS-304 grade sheets | No Change |
| 9 | The doors of the unit shall also be made of stainless steel. | No Change |
| 10 | The doors shall be connected by sturdy heavy-duty chrome plate hinges and fitted with hard chrome plated lubricated latches for opening the door. | No Change |
| 11 | The doors shall have a heated perimeter to eliminate surface condensation | No Change |
| 12 | There shall be Individual standard key lock for each chamber | No Change |
| 13 | All the doors to be fitted with high quality triple point neoprene rubber gaskets for air tight fittings and lock. | No Change |
| 14 | The units shall have washable interiors with channel for water outlet that can be plugged with rodent resistant material. | No Change |
| 15 | The units shall have a water proof CFL lamp inside | No Change |
| 16 | The body trays used shall have telescopic sturdy castors and castor locks to prevent rolling out of the trays | No Change |
|  | The trays shall also be made of SS-304. | No Change |
|  | The units shall have a Microprocessor based temperature control. | No Change |
|  | The temperature range of the units shall be +2 to +8°C | No Change |
|  | The units shall have a digital LED display indicating the temperature of the unit. | No Change |
|  | There shall be audio-visual alarms in case of low/high temperatures, system malfunctions etc. An external temperature display and hooter outside the mortuary room shall be provided | No Change |
|  | The units shall have a touchpad for adjusting temperature and alarm settings. | No Change |
|  | The cooling units used shall be ISI certified high end ultra-mute CFC-free hermetically sealed compressors, conforming to latest international standards and guidelines. | No Change |
|  | The condenser units shall have an evaporating system so as to avoid condensation and formation of ice. The components of the refrigeration unit shall be suitably placed outside the chamber. There shall be no serviceable components inside the chamber. | No Change |
|  | Shall have a forced air circulation system | No Change |
|  | Shall have an automatic defrosting system | No Change |
|  | Each individual unit shall have a noise level less than 60dB | No Change |
|  | Power Requirements: |  |
|  | 1. Shall work on single phase 220-230V AC, 50-60 Hz | No Change |
|  | 2. Shall be supplied with a voltage stabilizer of suitable rating for each unit/entire assembly | No Change |
|  | Protection: |  |
|  | 1. Shall have resettable overcurrent breakers of suitable rating for electrical protection | No Change |
|  | Accessories & Spares: - |  |
|  | A voltage stabilizer of suitable rating for individual units or for the entire unit | No Change |
|  | Certifications: - |  |
|  | 1. Manufacturer should have ISO 45001:2018 certification with reference to manufacturing of air conditioning and refrigeration products | US FDA/European CE/BIS/CE/ISO Approved model should be offered, The ISO certifications must be issued by bodies registered with NABCB under the Medical Devices Quality Management System. |