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Corrigendum-I      
[bookmark: _GoBack]Bihar Medical Services and Infrastructure Corporation Limited (BMSICL) had invited E-Bids from the interested parties for Procurement, Rate contract, Supply, Installation of Medical Equipment vide Tender No.-BMSICL/2025-26/ME-434. During Pre-bid meeting various suggestions were received from different prospective bidders regarding amendment in technical specification of equipment which were discussed and deliberated on by the experts. On the basis of their recommendations certain amendments in the technical specification of the equipment have been made which are annexed as Annexure-I of this corrigendum. 
Note:-
1. Bidders are advised to refer to the Annexure-I of this corrigendum before submission of bid. 
2. Those who have submitted their bids are requested to re-submit their bids in accordance with this corrigendum. 

Annexed:- as above

											
  GM (Procurement)                         
                                                                                                                       BMSICL











	Annexure-I

	Name of Equipment :-  MORTUARY CHAMBER (TWO BODY)                      

	SI. No
	Technical Specification as per tender
	Final Amendment

	1
	A two-body unit exclusively for keeping decomposed bodies
	No Change

	2
	Each unit shall have their individual, single cooling system to ensure the requisite temperature
	No Change

	3
	All the units shall have a proper design ensuring best hygiene
	No Change

	4
	All units shall be highly energy efficient
	No Change

	5
	All units shall have a double wall with the gaps to be filled with high grade polyurethane insulation ensuring maximum thermal efficiency.
	No Change

	6
	The puff used shall have a density of at least 40kg/m3
	No Change

	7
	The external and internal body of the mortuary units shall be corrosion free
	No Change

	8
	The inner and outer walls shall be made of SS-304 grade sheets
	No Change

	9
	The doors of the unit shall also be made of stainless steel.
	No Change

	10
	The doors shall be connected by sturdy heavy-duty chrome plate hinges and fitted with hard chrome plated lubricated latches for opening the door.
	No Change

	11
	The doors shall have a heated perimeter to eliminate surface condensation
	No Change

	12
	There shall be Individual standard key lock for each chamber
	No Change

	13
	All the doors to be fitted with high quality triple point neoprene rubber gaskets for air tight fittings and lock.
	No Change

	14
	The units shall have washable interiors with channel for water outlet that can be plugged with rodent resistant material.
	No Change

	15
	The units shall have a water proof CFL lamp inside
	No Change

	16
	The body trays used shall have telescopic sturdy castors and castor locks to prevent rolling out of the trays
	No Change

	 
	The trays shall also be made of SS-304.
	No Change

	 
	The units shall have a Microprocessor based temperature control.
	No Change

	 
	The temperature range of the units shall be +2 to +8°C
	No Change

	 
	The units shall have a digital LED display indicating the temperature of the unit.
	No Change

	 
	There shall be audio-visual alarms in case of low/high temperatures, system malfunctions etc. An external temperature display and hooter outside the mortuary room shall be provided
	No Change

	 
	The units shall have a touchpad for adjusting temperature and alarm settings.
	No Change

	 
	The cooling units used shall be ISI certified high end ultra-mute CFC-free hermetically sealed compressors, conforming to latest international standards and guidelines.
	No Change

	 
	The condenser units shall have an evaporating system so as to avoid condensation and formation of ice. The components of the refrigeration unit shall be suitably placed outside the chamber. There shall be no serviceable components inside the chamber.
	No Change

	 
	Shall have a forced air circulation system
	No Change

	 
	Shall have an automatic defrosting system
	No Change

	 
	Each individual unit shall have a noise level less than 60dB
	No Change

	 
	Power Requirements:
	 

	 
	1. Shall work on single phase 220-230V AC, 50-60 Hz
	No Change

	 
	 2. Shall be supplied with a voltage stabilizer of suitable rating for each unit/entire assembly
	No Change

	 
	Protection:
	 

	 
	1. Shall have resettable overcurrent breakers of suitable rating for electrical protection
	No Change

	 
	Accessories & Spares: -
	 

	 
	 A voltage stabilizer of suitable rating for individual units or for the entire unit
	No Change

	 
	Certifications: -
	 

	 
	1. Manufacturer should have ISO 45001:2018 certification with reference to manufacturing of air conditioning and refrigeration products
	US FDA/European CE/BIS/CE/ISO Approved model should be offered, The ISO certifications must be issued by bodies registered with NABCB under the Medical Devices Quality Management System.
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