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Corrigendum-I      
Bihar Medical Services and Infrastructure Corporation Limited (BMSICL) had invited E-Bids from the interested parties for the procurement, rate contract and the supply of medical equipment for different Govt. Institutions of Bihar vide Notice Inviting Tender No.- BMSICL/2025-26/ME-445. During and after Pre-bid meeting various suggestions were received from different prospective bidders regarding amendment in technical specification of equipment which were discussed and deliberated on by the experts, who after due deliberation recommended certain amendments in the technical specification of the equipment, which are annexed as Annexure-I of this corrigendum. In order to facilitate maximum participation of bidders the tender schedule is being revised as follows:-
	Tender Reference No. 
	BMSICL/2025-26/ME-445

	Last date and time of submission of online bids
	09th  October 2025 till 17:00 Hrs. 

	Last date and time of submission of original documents of EMD, Tender Fee and Document
	10th  October 2025 till 14:00 Hrs. 

	Date, Time and Place of opening of Technical Bid
	[bookmark: _GoBack]10th October 2025 (at 15:00 Hrs.) on the website of https:/eproc2.bihar.gov.in in the office of BMSICL 

	Date and time of opening of financial Bids
	To be announced later on https:/eproc2.bihar.gov.in



Note:- 
        1. Bidders are advised to refer to the Annexure-I of this corrigendum before submission of bid.
 2. Those who have submitted their bids are requested to re-submit their bids in accordance  with this corrigendum.                                                                

Annexed:- as above

											
  GM (Procurement)                         
                                                                                                                        BMSICL





	Annexure-I

	Name of Equipment: - Neonatal Ventilator

	SI. No
	Technical Specification as per tender
	Final Amendment

	1
	NIV (Non Invasive Ventilation) to be possible in all modes of ventilation available.
	No Change

	2
	Microprocessor controlled system with individual selection of various ventilation parameters & PEEP.
	No Change

	3
	Touch screen monitor with minimum 12’’ LED display to display real-time display of 2 wave forms & 2 loops with monitoring
	No Change

	4
	Ventilator, Humidifier & Compressor Should be US FDA/European CE (issued by notified body) approved Model should be offered.
	(i) Ventilator, Humidifier & Compressor Should be US FDA/European CE (issued by notified body) approved Model should be offered. 
(ii) CDSCO Certificate should be submitted.

	5
	Should have battery backup at least 90 mins.
	No Change

	6
	It should allow the user to deliver conventional ventilation with proximal flow sensor as well as HFOV& compatible to deliver inhaled Nitric Oxide   (I NO)
	It should allow the user to deliver conventional ventilation with proximal flow sensor/ Inbuild mass flow sensor as well as HFOV& compatible to deliver inhaled Nitric Oxide (I NO).

	7
	Should have the following modes of ventilation:  
	No Change

	
	a) Assist/ Control,
	No Change

	
	b) Pressure control,    
	No Change

	
	c) Pressure support, 
	No Change

	
	d) SIMV with pressure support (Pressure and volume control),                                
	No Change

	
	e) PEEP,
	No Change

	
	f) Noninvasive ventilation-BIPAP, CPAP/NIV/High flow,
	No Change

	
	g) Apnea ventilation, user selectable, volume & pressure 
	No Change

	
	h) HFOV
	No Change

	8
	Should have facility to measure and display of the following parameters:                                                                                                                            
	No Change

	
	a. Airway Pressure (Peak & Mean)
	No Change

	
	b. Tidal volume (Inspired & Expired)    
	No Change

	
	c. Minute volume (Inspired & Expired) 
	No Change

	
	d. Respiratory mechanics                                 
	No Change

	
	e. Spontaneous Minute Volume
	No Change

	
	f. Total Frequency,     
	No Change

	
	g. FiO2      
	No Change

	
	h. PEEP, 
	No Change

	
	i. Plateau Pressure, 
	No Change

	
	j. Use selector Alarms for all measured & monitored parameters                                                      
	No Change

	
	k. Occlusion Pressure, 
	No Change

	
	L.  Pressure Flow & Volume curves,
	No Change

	
	m. Leak%               
	No Change

	
	n. Leak compensated Tidal Volume
	No Change

	9
	Should have facility for auto compensation for ET Tube. automatic compliance & Leakage compensation for circuit
	No Change

	10
	Conventional ventilation & HFO Ventilation Mode Parameters:                                                                                                                       
	No Change

	
	a) BPM: 4to120,  
	No Change

	
	b) Inspiratory Time: 0.1 to 2.0 second,
	No Change

	
	c) CPAP Pressure: 2 to 25 mbar       
	No Change

	
	d) Inspiratory Pressure: 10 to 65 mbar                   
	No Change

	
	e) FIO2: 21% to 100%                
	No Change

	
	f) Tidal Volume 2-200 ml with Volume Guarantee    
	No Change

	
	g) I: E Ratio: 1:1, 1:2, 1:3,   
	No Change

	
	h) HFO Mode parameters        
	No Change

	
	i) HFO Frequency should be wide range with 5 to 20 Hz
	No Change

	11
	Alarm:
a. Adjustable Alarm. - Low/high minute volume, low/high pressure, low/high tidal volume, low/high rate, apnea time, low/high oxygen. 

	No Change

	
	b. Special alarm - O2 cell Failure, flow sensor, battery, power supply, gas supply, oxygen concentration,
	No Change

	12
	Should be supplied with one ultrasonic nebulizer (inbuilt or external type) with particle size of less than 3 micron Along with each Ventilator.
	No Change

	
	(i) Ventilator, Compressor & Humidifier should be Same Trolley/cart mounting for easy transportation.
	No Change

	
	(ii)Humidifier
a. Servo controlled heated Respiratory Humidifier. 
	No Change

	
	b. Display should be of LED /LCD.      
	No Change

	
	c. Temperature control settings & Temperature range: 28-40 deg.         
	No Change

	
	d. Temperature should be adjustable.   
	No Change

	
	e. Jar should be autoclavable
	No Change

	12
	Standard Accessories/Spare & Consumable
	No Change

	
	a Silicon breathing circuit circuit (Neonatal reusable) - 5 complete set.     
	No Change

	
	b. Nebulization assembly compatible circuit-5 complete set
	No Change

	
	c. Humidifier - 2 No. 
	No Change

	
	d. Hose for O2 connection with connector - 5 mts.    
	No Change

	
	e. Hose for compressed air with connector - 5 mts.    
	No Change

	
	 f. Test lung - 1 No.  
	No Change

	
	g. Inbuilt / integrated nebulizer-1 N0.    
	No Change

	
	h. All sensors and other non-consumable items (other than reusable silicon ventilator circuits) should be free of cost during 3 years of warranty and 
	No Change

	
	i. Silicon breathing circuit circuit (Neonatal reusable) - 5 complete set.    
	No Change

	
	j. Nebulization assembly compatible circuit-5 complete set.
	No Change

	
	k. Humidifier - 2 No.  
	No Change

	
	l. Hose for O2 connection with connector - 5 mts. 
	No Change

	
	 m. Hose for compressed air with connector - 5 mts. 
	No Change

	
	 n. Test lung - 1 No.  
	No Change

	
	o. Inbuilt / integrated nebulizer-1 N0.    
	No Change

	
	p. All sensors and other non-consumable items (other than reusable silicon ventilator circuits) should be free of cost during 3 years of warranty and 7 years of CMC.     
	No Change

	
	q. All attachment regarding non-invasive ventilation should be provided in all sizes minimum of 10 in number (additional).
	No Change

	13
	Ventilator, Humidifier & Compressor Power Supply Input to be 200-240V AC,50 HZ Fitted with Indian Condition Plug.
	No Change

	14
	Suitable online UPS with commensurate capacity for all ventilators including compressor & Humidifier with maintenance free batteries for minimum one hours back-up should be supplied.
	No Change

	15
	Ventilator, Humidifier & Compressor should be US FDA/European CE (issued by notified body) approved Model
 NOTE: 
1) Reusable consumables (other than reusable silicon ventilator circuits) should last during the warranty period. 2) Ventilator & Humidifier any additional reusable consumables are required during the warranty period those will be supplied free of charge by the supplier.
3) The life expectancy of the reusable consumable is expected to be of at least one year from the date of installation of the same. The reusable consumables will be procured at the prices accepted as per the contract. 
4) The bidders should submit all reusable consumable items price & their authorized local office/distributor name in the financial bid.

	No Change

	Name of Equipment: -. Oxygen Hood

	SI. No
	Technical Specification as per tender
	Final Amendment

	1
	Appropriate to comfortably fit all size babies up to 5 years of age. (Small and medium size)
	No Change

	2
	Transparent Polycarbonate unbreakable single moulded
	No Change

	3
	Silicon rubber Neck Port adjustment enabled to minimize the wastage of oxygen. Silicone rubber Neck port adjustment to ensures use in Neonate/Infant/Paediatric patients
	No Change

	4
	Oxygen inlet Port
	No Change

	5
	Oxygen tube of 2 m length must be provided.
	No Change

	6
	Should have USFDA/ European CE (Issued by Notified Body) / BIS/ISO 13485 certified.
	No Change

	7
	Supplied with 2 spare set of oxygen tubing set each size of oxygen Hood.
	No Change

	Name of Equipment: - Flame Photometer

	SI. No
	Technical Specification as per tender
	Final Amendment

	1
	Microprocessor Control with Filters, Na, K, Li, Ca,-Double display Filters-Narrow band-pass interference filters for sodium, Potassium, Calcium and Lithium available.
	No Change

	2
	Power supply – 220 V ± 10 V 50 Hz  
	No Change

	3
	Display- large, Bright Seven Segment LED /LCD Panel meter to display 0-100 scale with negative and over range indication.
	No Change

	4
	Compressor- with built in regulator and Air Filter to deliver stable and moisture/oil free air supplier.
	No Change

	5
	Ignition System- Auto/Manual ignition.
	No Change

	6
	Atomiser- Non-corroding, easily accessible and can be cleaned without putting out the flame.
	No Change

	7
	Sample Take up rate- 5ml-6ml/min approximately.
	No Change

	8
	Should have ISO 13485 (NABCB Accredited)

	No Change
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