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Corrigendum-I

Bihar Medical Services and Infrastructure Corporation Limited (BMSICL) had invited E-Bids from the interested parties for the procurement, rate contract and the supply of medical equipment for different Govt. Institutions of Bihar vide Notice Inviting Tender No.- BMSICL/2025-26/ME-446. During and after Pre-bid meeting various suggestions were received from different prospective bidders regarding amendment in technical specification of equipment which were discussed and deliberated on by the experts, who after due deliberation recommended certain amendments in the technical specification of the equipment, which are annexed as Annexure-I of this corrigendum. In order to facilitate maximum participation of bidders the tender schedule is being revised as follows:-
	Tender Reference No. 
	BMSICL/2025-26/ME-446

	Last date and time of submission of online bids
	09th  October 2025 till 17:00 Hrs. 

	Last date and time of submission of original documents of EMD, Tender Fee and Document
	10th  October 2025 till 14:00 Hrs. 

	Date, Time and Place of opening of Technical Bid
	10th October 2025 (at 15:00 Hrs.) on the website of https:/eproc2.bihar.gov.in in the office of BMSICL 

	Date and time of opening of financial Bids
	To be announced later on https:/eproc2.bihar.gov.in



Note:- 
        1. Bidders are advised to refer to the Annexure-I of this corrigendum before submission of bid.
 2. Those who have submitted their bids are requested to re-submit their bids in accordance   with this corrigendum.                                                                

[bookmark: _GoBack]Annexed:- as above

											
  GM (Procurement)                         
                                                                                                                       BMSICL





	Annexure-I

	Name of Equipment :- Polymerase Chain Reaction (PCR) Analyzer

	SI. No
	Technical Specification as per tender
	Final Amendment

	1
	Molecular system should be based on Nested multiplex PCR technology.
	The System should be an automated & integrated system for both real-time PCR and post-PCR (end-point) analysis.

	2
	The system should support applications including absolute and relative quantitation, multiplex-PCR allelic discrimination (SNP), melt curve analysis, pathogen detection and plus/ minus assays using an internal positive control.
	No Change

	3
	Should have minimum manual hands-on time less than 5 min and results should be available within 70 mins.
	Results should be available within 90 mints depending on selected protocol.

	4
	The system should have temperature range of 4-100 degrees C with peak block ramp rate for heating as well as cooling exceeding 4degrees C /second.
	No Change

	5
	Reaction volume should be approximately in the range of 10-30 µL and can run the templates from different sources simultaneously.
	No Change

	6
	The System should support both 0.1 and 0.2-ml tubes/strips/plates in single block.
	The system should support either 0.1 ml or 0.2 ml tubes/strips /plates in a single block.

	7
	The System should provide Touch Screen LCD Feature or Desktop/Laptop for System Control, operation, analysis and net-working of multiple system.
	No Change

	8
	Instrument should be supplied with a separate Desktop.
	No Change

	9
	Remote monitoring to analyse data by online web-browser based software or cloud-based data access browser should be available.
	Remote monitoring to analyse data by online web-browser based software or cloud-based data access browser should be available. (Optional)

	10
	USB port for data export to Power point, Excel or JPEG formats should be present.
	No Change

	11
	The system should be complete with licensed software’s for designing probes and primers.
	The system should be complete with licensed software’s for designing probes and primers. (Optional)

	12
	System should be capable of operating in test development and IVD modes with enhanced security to enable compliance with regulations.
	No Change

	13
	It should be an open system to use reagents from any manufacturer.
	No Change

	14
	System should be calibrated for common dyes of broader wavelength excitation and emission spectrum.
	No Change

	15
	Excitation Light source –Laser / LED.
	No Change

	16
	Change of position, block or chemistry should not require calibration or tools.
	No Change

	17
	The system should be capable of operating at ambient temperatures of 20-30 degrees C and relative humidity of 80%.
	No Change

	18
	Power input to be 220-240VAC, 50Hz fitted with Indian plug.
	No Change

	19
	Compatible online UPS with maintenance free batteries for minimum one-hour back-up should be supplied with the system.
1. System should have US FDA/ EUCE (Issued by notified body)
2. CE/IVD (optional)

	No Change

	Name of Equipment :- Electric Needle Destroyer

	SI. No
	Technical Specification as per tender
	Final Amendment

	1
	Should be Portable and electrical type.
	No Change

	2
	Principle of operation should be Electro-melting type.
	No Change

	3
	Material of the Housing /enclosure should be ABS Plastic.
	No Change

	4
	Housing/enclosure shall be molded type and shock proof.
	No Change

	5
	Provision to burn the needles and to cut the syringe tips shall be provided in the unit.
	No Change

	6
	Number of needles of 1mm dia and 80mm length that can be destroyed in continuous operation of 5 minutes
	No Change

	7
	Transformer winding should be copper.
	No Change

	8
	Sizes of injection needles of all kinds which can be destroyed (Dia Ranging from 0.4mm to 1.6mm (26 SWG to 14 SWG) with length12.5mm to 80mm).
	No Change

	9
	Power supply should be 207V to 253V,50Hz AC supply.
	No Change

	10
	Power ON/OFF switch with indicator shall be provided.
	No Change

	11
	Rated power (Watts) must be 60 watt.
	No Change

	12
	Shall be provided with fuse and power cord of min 3m length and earthing point.
	No Change

	13
	Certification:-
1. The quoted model should be either USFDA approved/European CE certified/ equivalent BIS/ISO 13485 certificate issued from certification bodies accredited by NABCB.
2. The unit should comply with relevant IEC safety standard and should submit the IEC certificate.

	No Change
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