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Corrigendum-I     
Bihar Medical Services and Infrastructure Corporation Limited (BMSICL) had invited E-Bids from the interested parties for the procurement, rate contract and the supply of medical equipment for different Govt. Institutions of Bihar vide Notice Inviting Tender No.- BMSICL/2025-26/ME-449. During and after Pre-bid meeting various suggestions were received from different prospective bidders regarding amendment in technical specification of equipment which were discussed and deliberated on by the experts, who after due deliberation recommended certain amendments in the technical specification of the equipment, which are annexed as Annexure-I of this corrigendum. In order to facilitate maximum participation of bidders the tender schedule is being revised as follows:-
	Tender Reference No. 
	BMSICL/2025-26/ME-449

	Last date and time of submission of online bids
	20th November 2025 till 17:00 Hrs. 

	Last date and time of submission of original documents of EMD, Tender Fee and Document
	[bookmark: _GoBack]21st November 2025 till 14:00 Hrs. 

	Date, Time and Place of opening of Technical Bid
	21st  November 2025 (at 15:00 Hrs.) on the website of https:/eproc2.bihar.gov.in in the office of BMSICL 

	Date and time of opening of financial Bids
	To be announced later on https:/eproc2.bihar.gov.in



Note:- 
        1. Bidders are advised to refer to the Annexure-I of this corrigendum before submission of bid.
 2. Those who have submitted their bids are requested to re-submit their bids in accordance 
     With this corrigendum.                                                                

Annexed: - as above

											
  GM (Procurement)                         
										                 BMSICL





	Annexure- I

	Name of Equipment:- Arthroscopy specialty Centre & Sports Physiotherapy

	SI. No
	Technical Specification as per tender
	 Final Amendment

	 
	Name of Equipment :- (1) High End Arthroscopy System  (For Medical College & Hospital and Super Specialty)
	No Change

	 
	Equipment
	Technical Specification as per tender
	No Change

	1
	Arthroscopic Instruments: Meniscal Punches, graspers, scissors, shavers, probes
	a. Probe Hook with Metal Handle 5-5.5 mm tip
	No Change

	
	
	b. Straight Punch Bakset Type
	No Change

	
	
	c. Left Punch Basket Type
	No Change

	
	
	d. Right Punch Basket Type
	No Change

	
	
	e. Grasper:  All grasper should have an infinite position sliding lock mechanism that hold tissue firmly without tearing and slipping – even in the tightest area.            · Pitbull Loose body Grasper with sliding lock mechanism
	No Change

	
	
	f. Low Profile Suture Retriever Straight Jaw
	No Change

	
	
	g.Arthroscopic Scissors 3.4mm Diameter                                                         · Scissor Punches should be straight, loop handle
	No Change

	
	
	· Scissor Punches Should be 20deg. Hooked Left
	No Change

	
	
	· Scissor Punches should be 20deg. Hooked Right
	No Change

	
	
	h. Knot Pusher Closed End 
	No Change

	
	
	i. Suture Cutter for #2-0, #2, #5 sutures
	No Change

	
	
	Meniscal Stitcher Set:A comprehensive system suitable for inside-out technique Comprising of – 1. Double Lumen Cannula,curved up-right 2. Double Lumen Cannula,curved up-left 3. Double lumen cannula ,straight 4. Double Lumen cannula left/right 5. Double lumen cannula,curved up/down 6. Single cannula,2 included with set 7. Thimble 8. Posterior Access Cannula 9. Bending Tool 10. Sterilization Tray for Mensical Stitcher Set 11. Needles straight
	No Change

	2
	ACL, PCL, Knee, Shoulder Instrumentation set
	ACL, PCL and Knee Instrumentation System
	No Change

	
	
	1. ACL Transportal offset
	No Change

	
	
	a. 6mm, 7mm, 8mm
	No Change

	
	
	b. Handle should be colour coded for easy identification
	No Change

	
	
	c. Markings to be graduated on shaft of Offset for evaluation under scope
	No Change

	
	
	d. Finger point type design should be there for easier reach
	No Change

	
	
	2. Tendon Stripper Closed
	No Change

	
	
	a. Should be at least 300mm in length
	No Change

	
	
	b. 7mm Dia
	No Change

	
	
	c. Markings to be graduated at 10 mm interval
	No Change

	
	
	3. Tendon Stripper Open
	No Change

	
	
	a. Standard Design
	No Change

	
	
	b. Markings to be graduated at 10 mm interval
	No Change

	
	
	c. Should be 300mm in length
	No Change

	
	
	4. 4.5 Endoscopic Drill
	No Change

	
	
	a. Markings to be Graduated at every 10 mm
	No Change

	
	
	b. 2.4mm cannulation
	No Change

	
	
	5. Cannulated Headed Reamer (Flower tip)
	No Change

	
	
	a. 6mm, 7mm, 8mm, 9mm, 10mm, 11mm
	No Change

	
	
	b. Marking to be Graduated every 5mm
	No Change

	
	
	c. 2.4mm cannulation
	No Change

	
	
	6. Tibial Reamer
	No Change

	
	
	a. 6mm, 7mm, 8mm, 9mm, 10mm, 11mm
	No Change

	
	
	b. Markings to be Graduated every 10mm
	No Change

	
	
	c. 2.4mm cannulation
	No Change

	
	
	7. Tibial Zig and Femoral zig with Marking Hook and Bullet
	No Change

	
	
	a. Should have side Release Zig Handle
	No Change

	
	
	b. Tibial Marking Zig should Have Finger Tip Design
	No Change

	
	
	c. 5mm marking for angle calculation on ACL marking Hook
	No Change

	
	
	d. 2.4mm cannulation in bullet with Ratcheting Sleeve
	No Change

	
	
	e. 10 mm markings lo be graduated for the Identification of tibial length
	No Change

	
	
	f. Femoral Zig with at least 4mm increments! marking on tip and 5mm on handle
	No Change

	
	
	8. Screw Driver for interference 
	No Change

	
	
	9. Graft Preparation System
	No Change

	
	
	a. Graft Master Board with multiple attachments compatibility
	No Change

	
	
	b. Master board with Scale Function
	No Change

	
	
	c. Graft Sizing Block with 4.5mm-12mm holes in 0.5mm increments
	No Change

	
	
	10. PCL Zig handle
	No Change

	
	
	a. Femoral PCL zig handle with 2mm markings graduated on Tip
	No Change

	
	
	b. 5mm markings graduated on Handle
	No Change

	
	
	c. Finger Tip Design With Pin for Stabile placement
	No Change

	
	
	d. Tibial PCL Zig handle with 2nm markings graduated on Tip
	No Change

	
	
	e. 5mm markings graduated on Handle
	No Change

	
	
	f. Finger Tip Design With Pin for Stable Placement
	No Change

	
	
	11. PCL Instruments
	No Change

	
	
	a. PCL Suture Pusher
	No Change

	
	
	b. PCL Rasp with Curved Shaft
	No Change

	
	
	c. PCL Poplileteal Curette / retractor
	No Change

	
	
	12. Instrument Case/Box  for Safe Storage
	No Change

	
	
	Shoulder Instruments Set
	No Change

	
	
	1. Cannula Instrument
	No Change

	
	
	a. Short and Thin Orbutrator for 6-8mm Cannula
	No Change

	
	
	b.Long and Thin Orbutrator for 6-8mm Cannula
	No Change

	
	
	c. Short and Thick Orbutrator for 6-8mm Cannula
	No Change

	
	
	d. Long and Thick Orbutrator for 6-8mm Cannula
	No Change

	
	
	e. Extra Long Switching Stick, 4 mm x 12 in - 2 Pc
	No Change

	
	
	f. 8 mm silicon/ Inserter Cannula
	8 mm silicon/ Inserter Cannula 10 nos.

	
	
	2. Suture management Instruments
	No Change

	
	
	a. Probe, Hook 5-5.5 mm, Tip w/5 mm Markings
	No Change

	
	
	b. Suture Retriever, ø3.4 mm, Straight
	No Change

	
	
	c. BirdBeak, 2.75 mm, 22° Up Tip
	No Change

	
	
	d. BirdBeak, 2.75 mm, 45° Up Tip
	No Change

	
	
	e. Suture/Crochet Hook
	No Change

	
	
	f. Knot Pusher, closed end
	No Change

	
	
	g. Suture Passer Reusable Type with Disposable Needle
	No Change

	
	
	j. Penetrator Suture Retriever, 15°
	No Change

	
	
	k. Suture Cutter, 4.2(+/- 10%) mm, Straight, closed end
	No Change

	
	
	l. FiberTape Suture Cutter
	No Change

	
	
	m. grasper and retriver
	No Change

	
	
	3. Tissue Management Instruments
	No Change

	
	
	a. Rotator Cuff Grasper, 4.2 mm Straight Shaft w/SR Handle
	Rotator Cuff Grasper, 3.4 to 4.5 mm Straight Shaft w/SR Handle

	
	
	b. Suture Retriever/Tissue
	No Change

	
	
	Grasper
	No Change

	
	
	c. Curette
	No Change

	
	
	d. Shoulder Debridement Rasp UP/Down
	No Change

	
	
	g. Shoulder eleveter UP/Down
	No Change

	
	
	5. Instrument Case or Box for safe storage
	No Change

	3
	Arthroscopic Fluid Management System
	Pump System should have intuitive touchscreen control 
	No Change

	
	
	Should have Presets for various joints like (Shoulder, Knee, Small Joint, Hip) 
	No Change

	
	
	The Flowrate: ≥ 1500 ml/min should be automatically adjusted 
	No Change

	
	
	The max flowrate setting should be adjustable from 50% to 100% (increments of 10%) 
	No Change

	
	
	Flow rate should be change as per operating cannula connection
	No Change

	
	
	The pressure setting should be adjustable between 10 and 120. (Increments of 5). 
	No Change

	
	
	Should have automatic Joint pressure maintenance up-to 120 mmHg
	No Change

	
	
	Should have the real time flow rate monitoring
	No Change

	
	
	The estimated fluid usage information should be available.
	No Change

	
	
	Should have interface with Shaver System if required
	No Change

	
	
	Should have shaver boost functionality: automated pressure increases when shaver is utilized; 0 to 50% (10% increment steps) 
	No Change

	
	
	Should have Lavage mode – increases pressure when utilized 
	No Change

	
	
	should be multiple tubing option
	No Change

	
	
	The overpressure control should be 300 mmHg ± 5
	No Change

	4
	Arthroscopic leg holder and traction system
	
	Arthroscopic leg holder and Shoulder Distraction System

	
	
	a. 3-Point Shoulder Distraction System 
	No Change

	
	
	b. Atraumatic Hand Held System
	No Change

	
	
	Arthroscopy Leg Holder, including release clamp 25 mm, neoprene cushion and closure stripe: 1 no. each
	No Change

	
	
	 All-round padding. 360-degree swivel and rotating clamp. Height-adjustable with a clamp to fit the operating table side rail. 
	No Change

	5
	Surgical Power Tools: Electric drills and saws and Reamer Attachment
	Power System GENERAL INFORMATION Intended Use
	No Change

	
	
	The intended use of the Power System is for cutting, shaping, and preparing bone and bone-related tissues in a variety of surgical procedures. The system is usable for performing numerous operations such as sawing, reaming and drilling as well as driving wires, pins and various instruments
	No Change

	
	
	PRODUCT CHARACTERISTICS
	No Change

	
	
	Double Trigger Handpiece
	No Change

	
	
	Type: Double Trigger Handpiece
	No Change

	
	
	Max. drill speed: 1200 rpm or more
	No Change

	
	
	Max. saw speed: 12000 cpm or more
	No Change

	
	
	Max. torque: 3.5 Nm or more
	No Change

	
	
	Mode: Forward/Reverse/Oscillating
	No Change

	
	
	Speed Trigger: Variable
	No Change

	
	
	Toggle Switch: On/Off
	No Change

	
	
	Cannulation: 4.0 mm
	No Change

	
	
	Dimensions (HxWxD): 152 x 127 x 52 mm± 10%
	No Change

	
	
	Weight: Should Not be more than 1250 gm(Without Battery)
	No Change

	
	
	Material: Aluminum/ Stainless Steel
	No Change

	
	
	Compatibility: Battery and AC Power Supply 
	Compatibility: Battery

	
	
	Drill Attachments
	No Change

	
	
	Max. Speed: 12000 rpm or more
	Max. Speed: 1200 rpm or more

	
	
	Max. Torque: 3.5 Nm
	No Change

	
	
	Shank Styles: Small AO, Hudson, Trinkle, Zimmer/ Hudson, 1/4in Key Hybrid Chuck, 5/32in Key Hybrid Chuck
	No Change

	
	
	Material: Stainless Steel
	No Change

	
	
	Standard Reamer Attachments
	No Change

	
	
	Max. Speed: 600 rpm
	No Change

	
	
	Max. Torque: 9 Nm or more 
	No Change

	
	
	Shank Styles: AO, Small AO, Hudson, Trinkle, Zimmer/ Hudson, Harris, DCS/DHS, 1/4in Key Hybrid Chuck
	No Change

	
	
	Material: Stainless Steel
	No Change

	
	
	High Torque Reamer Attachments
	No Change

	
	
	Max. Speed: 270 rpm
	No Change

	
	
	Max. Torque: 17.5 Nm or more
	No Change

	
	
	Shank Styles: AO, Hudson, Trinkle, Harris, Zimmer/Hudson, DCS/DHS, 1/4in Key Hybrid Chuck
	No Change

	
	
	Material: Stainless Steel
	No Change

	
	
	Pin Driver Max. Speed: 1200 rpm
	No Change

	
	
	Max. Torque: 5 Nm
	No Change

	
	
	Clamping Range: 2.0 – 4.0 mm
	No Change

	
	
	Material: Stainless Steel
	No Change

	
	
	Special Feature: No pin diameter adjustment needed
	Special Feature:  pin diameter adjustment needed

	
	
	Wire Driver
	No Change

	
	
	Max. Speed: 1200 rpm
	No Change

	
	
	Max. Torque: 5 Nm or more
	No Change

	
	
	Clamping Range: 0.8 – 2.0 mm
	No Change

	
	
	Material: Stainless Steel
	No Change

	
	
	Special Feature: No pin diameter adjustment needed
	Special Feature: pin diameter adjustment needed

	
	
	Sagittal Saw Attachment
	No Change

	
	
	Max. Speed: 12000 cpm
	No Change

	
	
	Arc of Deflection: 4.5° (+/- 2.25°)
	No Change

	
	
	Material: Stainless Steel
	No Change

	
	
	Special Feature: Single blade can be adjusted +/- 10mm from nominal position within the attachment
	Deleted

	
	
	Reciprocating Saw Attachment
	No Change

	
	
	Max. Speed: 12000 cpm
	No Change

	
	
	Stroke: 3.8 mm± 5%
	No Change

	
	
	Material: Stainless Steel
	No Change

	
	
	Special Feature: Reciprocating Saw attachment can be converted into a Sternum saw by connecting a Sternum Guard (Standard or Large)
	No Change

	
	
	Aseptic and Sterilisable Battery
	No Change

	
	
	Battery Type: Li-Ion, 14.4 VDC, 2000 mAh
	No Change

	
	
	Battery Capacity: 28.8 Wh±10 %
	No Change

	
	
	Charging type: max. 60min (in average: 40 min)
	No Change

	
	
	Special Feature: Equipped with real-time power gauge that indicates remaining battery capacity during the procedure
	No Change

	
	
	Battery Charger
	No Change

	
	
	Charging Bays: 4 or more
	No Change

	
	
	Electrical Requirements: 100–240 V ~ (AC), 50/60 Hz.
	No Change

	
	
	Special Feature: Visual indicator for real-time charging status and advanced battery error detection.
	No Change

	
	
	Sterilization Case
	No Change

	
	
	1 HP Basket with Housing Slot
	No Change

	Note-  All above Equipment/ Instrument should be USFDA/European CE Approved (Sl. No. 1 to 06)
	No Change

	 06
	External fixator (For District Hospital/SDH/CHC/PHC)
	No Change

	1
	Leg Set- consisting of commonly used tubular rod – 02 Nos., Schanz Screw 4.5mm -06 Nos and Polyaxial Clamps for above screw and rods – 12 Nos.   
	No Change

	2
	 Femur Set – consisting of commonly used Tubular Rods -02 Nos, Schanz Screw 4.5mm, 5mm -03 Nos. each and Polyaxial Clamps for above Screw and Rods -12 Nos.   
	No Change

	3
	 Forearm Set: Femur Set – consisting of commonly used solid rods for Forearm Fixator -02 Nos, Schanz Screw 3.5mm, 2.5mm - 03 Nos. each and Polyaxial Clamps for above screws and rods -12 Nos.
	No Change

	4
	USFDA/European CE (Issued by Notified body) /BIS Approved model should be offered.  
	No Change

	 07
	Pneumatic drill (For District Hospital/SDH/CHC/PHC)
	No Change

	1
	The Cannulated Pneumatic Drill handpiece - Compatible with existing attachments 
	No Change

	a.
	Cannulation with 3.2 mm diameter       
	No Change

	b.
	 Operating pressure : 6 - 7 bars (maximum 10 bars)      
	No Change

	c.
	Weight of handpiece 600-800grams without any attachments       
	No Change

	d.
	 Power 120 w       
	No Change

	e.
	Variable Speed from 0-900 rpm       
	No Change

	f.
	 Noise Level of max 75 db
	No Change

	g.
	Separate forward and reverse triggers       
	No Change

	h.
	Safety Device to cut off air supply to drill on handpiece       
	No Change

	i.
	 Handpiece is compatible with radiolucent drive       
	No Change

	j.
	Instant change between clockwise and counterclockwise rotation       
	No Change

	k.
	Offers reliable protection of soft tissues with oscillating drill attachment       
	No Change

	l.
	Fully Autoclavable       
	No Change

	m.
	 Fully machine washable       
	No Change

	n.
	 All Attachments can be fitted on single handpiece       
	No Change

	o.
	The reverse trigger automatically locks when the oscillating saw and the reduction drive attachments are attached to handpiece     
	No Change

	2
	 Adapter for Lubrication       
	No Change

	a.
	For oiling of hand piece &Autoclavable       
	No Change

	b.
	Should be made of Stainless Steel     
	No Change

	3
	Double Air Hose       
	No Change

	a.
	 Length 5 meters.       
	No Change

	b.
	It should be Autoclavable       
	No Change

	c.
	 Should have concentric inlet and outlet pipes     
	No Change

	4
	 Radiolucent Drive Precise aiming and drilling under image intensifier control for locking intramedullary nails      
	No Change

	 
	Drill Bit diameter 2.0 to 4.5 mm, 
	No Change

	 
	Length 100 to 150 mm, 
	No Change

	 
	Usable length 80 to 120 mm    
	No Change

	 
	 Reduced exposure to x-rays     
	No Change

	5
	Jacob's Chuck attachment      
	No Change

	 
	 Chuck capacity up to 0 to 6.5 mm  
	No Change

	 
	Cannulation of 3.2 mm diameter       
	No Change

	 
	Maximum Speed of 900rpm       
	No Change

	 
	Torque of 4-5 Nm     
	No Change

	6
	Quick Coupling attachment      
	No Change

	 
	 Cannulation 3.2 mm      
	No Change

	 
	 Maximum Speed: 900 rpm       
	No Change

	 
	Torque of 4-5 Nm    
	No Change

	7
	Reduction Drive for Intramedullary / Acetabular Reaming with reverse option       
	No Change

	 
	Reaming Speed of 300-350 rpm        
	No Change

	 
	Reaming Torque of 12-14 Nm       
	No Change

	 
	Option of attachment with reverse rotation    
	No Change

	8
	Quick Coupling for K-wire       
	No Change

	 
	Continous adjustment facility for wire diameter from 0.6 to 3.2 mm      
	No Change

	 
	 Speed up to 900 rpm   
	No Change

	9
	Oscillating Saw attachment with key      
	No Change

	 
	 It can operate on an oscillating frequency of 0 to 14,000 osc/min.       
	No Change

	 
	Attachment can be locked in 8 different positions    
	No Change

	10
	Quick Coupling for drill bits       
	No Change

	 
	Speed: 0–900 rpm      
	No Change

	 
	 Torque: 0–4.7 Nm      
	No Change

	 
	 Cannulation: 1.3 mm     
	No Change

	11
	Quick coupling for DHS / DCS triple reamers     
	No Change

	12
	Oscillating Saw Blades (All Sizes)      
	No Change

	 
	 For Trauma    6   For Joint replacement   
	No Change

	13
	6  Aluminium case for Pneumatic Drill system, Perforated, Autoclavable   1
	No Change

	14
	Aluminium Box for accessory attachment, Perforated, Autoclavable   1   Technical Specifications for Pneumatic Drill Set – 1. Company should have relevant experience in successful execution of similar work at least in five Institutes of national importance and central government Institutes 
	No Change

	3
	Bidder must enclose original literatures & technical data sheet in the support of the technical bid. 
	No Change

	4
	Physical demo may preferably be arranged at the time of requirement. 
	No Change

	5
	Instruments quality should meet the international standard. 
	No Change

	5
	USFDA/European CE (Issued by Notified body)  Approved model should be offered
	No Change

	 
	Reamer (For District Hospital/SDH/CHC/PHC)
	No Change

	1
	Should be supplied with a driving unit including motor, stand, foot control, flexible shaft, tool kit and container.
	No Change

	2
	Should have a cannulated drill hand piece of maximum speed 1200rpm with Non-corrosive Jacob chuck (0-1/4”). Autoclavable. 
	No Change

	3
	Should have a cannulated reaming hand piece with maximum speed 400rpm with AO type quick coupling. Autoclavable. 
	No Change

	4
	Should have a pistol grip sagital saw with a set of 5 blades. Autoclavable.
	No Change

	5
	Should have a flexible reamer shaft 8mm diameter fixed head. 
	No Change

	6
	Should have a flexible reamer shaft for detachable heads up to 12mm.
	No Change

	7
	 Should have reamer heads from 8mm to 12 mm (set of 9). 
	No Change

	8
	Should have reamer heads from 12.5mm to 15mm (set of 6). 
	No Change

	9
	Should have flexible reamer shaft for detachable heads up to 15mm. Autoclavable. 
	No Change

	10
	Should have an extra flexible shaft for driving unit autoclavable. 
	No Change

	11
	Should work with input 200 to 240Vac 50 Hz supply.
	No Change

	12
	USFDA/European CE (Issued by Notified body) /BIS Approved model should be offered.
	No Change

	8
	Arthroscopy Pump ( For High End Arthroscopy System-Medical College & Hospital)
	Deleted

	a
	It should have electronic digital suction irrigation pump with 5’’ or more colour touch LCD/LED screen display
	Deleted

	b
	It should have Precise presetting of flow and pressure of suction and irrigation parameters via. Touch Screen Display.
	Deleted

	C
	It should have built in low noise suction pump to allow simultaneous suction – irrigation process. 
	Deleted

	D
	It should have multiple utility mode, like Laparoscopy, Hysteroscopy, Arthroscopy, Urology etc. 
	Deleted

	E
	It should have irrigation pressure should be 10-400 mmHg or more.
	Deleted

	F
	It should have auto cut off facility for the preset amount of the fluid.
	Deleted

	G
	It should have preset multiple user setting for different surgeries. 
	Deleted

	H
	It should have flow rate 100-1000 ml/min.
	Deleted

	I
	It should have display of total volume consumed.
	Deleted

	J
	 It should have suction vacume bar 75Kpa or more.
	Deleted

	K
	It should have supplied with 2 ltr. Jar 02 nos.
	Deleted

	L
	It should have supplied with Y shaped irrigation tubing set.
	Deleted

	M
	It should have auto stop facility.
	Deleted

	N
	I- Power Supply 100-240 V, II -Frequency : 50 - 60 Hz
	Deleted




Page 2 of 2

image1.jpeg
%ﬂ"%amsmL




