|  | Bihar Medical Services & Infrastructure Corporation Limited, 2nd & 3rd Floor, Swasthya Bhawan, Behind IGIMS, Sheikhpura, Adjacent to State Health Society, Patna 800023, Phone/Fax: +91612 2283287,+ 91612 2283288 |
| --- | --- |

Corrigendum-I

Bihar Medical Services and Infrastructure Corporation Limited (BMSICL) had invited E-Bids from the interested parties for the procurement, rate contract and the supply of medical equipment for different Govt. Institutions of Bihar vide Notice Inviting Tender No.- BMSICL/2025-26/ME-449. During and after Pre-bid meeting various suggestions were received from different prospective bidders regarding amendment in technical specification of equipment which were discussed and deliberated on by the experts, who after due deliberation recommended certain amendments in the technical specification of the equipment, which are annexed as Annexure-I of this corrigendum. In order to facilitate maximum participation of bidders the tender schedule is being revised as follows:-

| Tender Reference No. | BMSICL/2025-26/ME-449 |
| --- | --- |
| Last date and time of submission of online bids | 20th November 2025 till 17:00 Hrs. |
| Last date and time of submission of original documents of EMD, Tender Fee and Document | 21st November 2025 till 14:00 Hrs. |
| Date, Time and Place of opening of Technical Bid | 21st November 2025 (at 15:00 Hrs.) on the website of https:/eproc2.bihar.gov.in in the office of BMSICL |
| Date and time of opening of financial Bids | To be announced later on https:/eproc2.bihar.gov.in |

Note:-

1. Bidders are advised to refer to the Annexure-I of this corrigendum before submission of bid.

2. Those who have submitted their bids are requested to re-submit their bids in accordance

With this corrigendum.

Annexed: - as above

GM (Procurement)

BMSICL

| Annexure- I |  |  |  |
| --- | --- | --- | --- |
| Name of Equipment:- Arthroscopy specialty Centre & Sports Physiotherapy |  |  |  |
| SI. No | Technical Specification as per tender |  | Final Amendment |
|  | Name of Equipment :- (1) High End Arthroscopy System (For Medical College & Hospital and Super Specialty) |  | No Change |
|  | Equipment | Technical Specification as per tender | No Change |
| 1 | Arthroscopic Instruments: Meniscal Punches, graspers, scissors, shavers, probes | a. Probe Hook with Metal Handle 5-5.5 mm tip | No Change |
|  |  | b. Straight Punch Bakset Type | No Change |
|  |  | c. Left Punch Basket Type | No Change |
|  |  | d. Right Punch Basket Type | No Change |
|  |  | e. Grasper: All grasper should have an infinite position sliding lock mechanism that hold tissue firmly without tearing and slipping – even in the tightest area. · Pitbull Loose body Grasper with sliding lock mechanism | No Change |
|  |  | f. Low Profile Suture Retriever Straight Jaw | No Change |
|  |  | g.Arthroscopic Scissors 3.4mm Diameter · Scissor Punches should be straight, loop handle | No Change |
|  |  | · Scissor Punches Should be 20deg. Hooked Left | No Change |
|  |  | · Scissor Punches should be 20deg. Hooked Right | No Change |
|  |  | h. Knot Pusher Closed End | No Change |
|  |  | i. Suture Cutter for #2-0, #2, #5 sutures | No Change |
|  |  | Meniscal Stitcher Set:A comprehensive system suitable for inside-out technique Comprising of – 1. Double Lumen Cannula,curved up-right 2. Double Lumen Cannula,curved up-left 3. Double lumen cannula ,straight 4. Double Lumen cannula left/right 5. Double lumen cannula,curved up/down 6. Single cannula,2 included with set 7. Thimble 8. Posterior Access Cannula 9. Bending Tool 10. Sterilization Tray for Mensical Stitcher Set 11. Needles straight | No Change |
| 2 | ACL, PCL, Knee, Shoulder Instrumentation set | ACL, PCL and Knee Instrumentation System | No Change |
|  |  | 1. ACL Transportal offset | No Change |
|  |  | a. 6mm, 7mm, 8mm | No Change |
|  |  | b. Handle should be colour coded for easy identification | No Change |
|  |  | c. Markings to be graduated on shaft of Offset for evaluation under scope | No Change |
|  |  | d. Finger point type design should be there for easier reach | No Change |
|  |  | 2. Tendon Stripper Closed | No Change |
|  |  | a. Should be at least 300mm in length | No Change |
|  |  | b. 7mm Dia | No Change |
|  |  | c. Markings to be graduated at 10 mm interval | No Change |
|  |  | 3. Tendon Stripper Open | No Change |
|  |  | a. Standard Design | No Change |
|  |  | b. Markings to be graduated at 10 mm interval | No Change |
|  |  | c. Should be 300mm in length | No Change |
|  |  | 4. 4.5 Endoscopic Drill | No Change |
|  |  | a. Markings to be Graduated at every 10 mm | No Change |
|  |  | b. 2.4mm cannulation | No Change |
|  |  | 5. Cannulated Headed Reamer (Flower tip) | No Change |
|  |  | a. 6mm, 7mm, 8mm, 9mm, 10mm, 11mm | No Change |
|  |  | b. Marking to be Graduated every 5mm | No Change |
|  |  | c. 2.4mm cannulation | No Change |
|  |  | 6. Tibial Reamer | No Change |
|  |  | a. 6mm, 7mm, 8mm, 9mm, 10mm, 11mm | No Change |
|  |  | b. Markings to be Graduated every 10mm | No Change |
|  |  | c. 2.4mm cannulation | No Change |
|  |  | 7. Tibial Zig and Femoral zig with Marking Hook and Bullet | No Change |
|  |  | a. Should have side Release Zig Handle | No Change |
|  |  | b. Tibial Marking Zig should Have Finger Tip Design | No Change |
|  |  | c. 5mm marking for angle calculation on ACL marking Hook | No Change |
|  |  | d. 2.4mm cannulation in bullet with Ratcheting Sleeve | No Change |
|  |  | e. 10 mm markings lo be graduated for the Identification of tibial length | No Change |
|  |  | f. Femoral Zig with at least 4mm increments! marking on tip and 5mm on handle | No Change |
|  |  | 8. Screw Driver for interference | No Change |
|  |  | 9. Graft Preparation System | No Change |
|  |  | a. Graft Master Board with multiple attachments compatibility | No Change |
|  |  | b. Master board with Scale Function | No Change |
|  |  | c. Graft Sizing Block with 4.5mm-12mm holes in 0.5mm increments | No Change |
|  |  | 10. PCL Zig handle | No Change |
|  |  | a. Femoral PCL zig handle with 2mm markings graduated on Tip | No Change |
|  |  | b. 5mm markings graduated on Handle | No Change |
|  |  | c. Finger Tip Design With Pin for Stabile placement | No Change |
|  |  | d. Tibial PCL Zig handle with 2nm markings graduated on Tip | No Change |
|  |  | e. 5mm markings graduated on Handle | No Change |
|  |  | f. Finger Tip Design With Pin for Stable Placement | No Change |
|  |  | 11. PCL Instruments | No Change |
|  |  | a. PCL Suture Pusher | No Change |
|  |  | b. PCL Rasp with Curved Shaft | No Change |
|  |  | c. PCL Poplileteal Curette / retractor | No Change |
|  |  | 12. Instrument Case/Box for Safe Storage | No Change |
|  |  | Shoulder Instruments Set | No Change |
|  |  | 1. Cannula Instrument | No Change |
|  |  | a. Short and Thin Orbutrator for 6-8mm Cannula | No Change |
|  |  | b.Long and Thin Orbutrator for 6-8mm Cannula | No Change |
|  |  | c. Short and Thick Orbutrator for 6-8mm Cannula | No Change |
|  |  | d. Long and Thick Orbutrator for 6-8mm Cannula | No Change |
|  |  | e. Extra Long Switching Stick, 4 mm x 12 in - 2 Pc | No Change |
|  |  | f. 8 mm silicon/ Inserter Cannula | 8 mm silicon/ Inserter Cannula 10 nos. |
|  |  | 2. Suture management Instruments | No Change |
|  |  | a. Probe, Hook 5-5.5 mm, Tip w/5 mm Markings | No Change |
|  |  | b. Suture Retriever, ø3.4 mm, Straight | No Change |
|  |  | c. BirdBeak, 2.75 mm, 22° Up Tip | No Change |
|  |  | d. BirdBeak, 2.75 mm, 45° Up Tip | No Change |
|  |  | e. Suture/Crochet Hook | No Change |
|  |  | f. Knot Pusher, closed end | No Change |
|  |  | g. Suture Passer Reusable Type with Disposable Needle | No Change |
|  |  | j. Penetrator Suture Retriever, 15° | No Change |
|  |  | k. Suture Cutter, 4.2(+/- 10%) mm, Straight, closed end | No Change |
|  |  | l. FiberTape Suture Cutter | No Change |
|  |  | m. grasper and retriver | No Change |
|  |  | 3. Tissue Management Instruments | No Change |
|  |  | a. Rotator Cuff Grasper, 4.2 mm Straight Shaft w/SR Handle | Rotator Cuff Grasper, 3.4 to 4.5 mm Straight Shaft w/SR Handle |
|  |  | b. Suture Retriever/Tissue | No Change |
|  |  | Grasper | No Change |
|  |  | c. Curette | No Change |
|  |  | d. Shoulder Debridement Rasp UP/Down | No Change |
|  |  | g. Shoulder eleveter UP/Down | No Change |
|  |  | 5. Instrument Case or Box for safe storage | No Change |
| 3 | Arthroscopic Fluid Management System | Pump System should have intuitive touchscreen control | No Change |
|  |  | Should have Presets for various joints like (Shoulder, Knee, Small Joint, Hip) | No Change |
|  |  | The Flowrate: ≥ 1500 ml/min should be automatically adjusted | No Change |
|  |  | The max flowrate setting should be adjustable from 50% to 100% (increments of 10%) | No Change |
|  |  | Flow rate should be change as per operating cannula connection | No Change |
|  |  | The pressure setting should be adjustable between 10 and 120. (Increments of 5). | No Change |
|  |  | Should have automatic Joint pressure maintenance up-to 120 mmHg | No Change |
|  |  | Should have the real time flow rate monitoring | No Change |
|  |  | The estimated fluid usage information should be available. | No Change |
|  |  | Should have interface with Shaver System if required | No Change |
|  |  | Should have shaver boost functionality: automated pressure increases when shaver is utilized; 0 to 50% (10% increment steps) | No Change |
|  |  | Should have Lavage mode – increases pressure when utilized | No Change |
|  |  | should be multiple tubing option | No Change |
|  |  | The overpressure control should be 300 mmHg ± 5 | No Change |
| 4 | Arthroscopic leg holder and traction system |  | Arthroscopic leg holder and Shoulder Distraction System |
|  |  | a. 3-Point Shoulder Distraction System | No Change |
|  |  | b. Atraumatic Hand Held System | No Change |
|  |  | Arthroscopy Leg Holder, including release clamp 25 mm, neoprene cushion and closure stripe: 1 no. each | No Change |
|  |  | All-round padding. 360-degree swivel and rotating clamp. Height-adjustable with a clamp to fit the operating table side rail. | No Change |
| 5 | Surgical Power Tools: Electric drills and saws and Reamer Attachment | Power System GENERAL INFORMATION Intended Use | No Change |
|  |  | The intended use of the Power System is for cutting, shaping, and preparing bone and bone-related tissues in a variety of surgical procedures. The system is usable for performing numerous operations such as sawing, reaming and drilling as well as driving wires, pins and various instruments | No Change |
|  |  | PRODUCT CHARACTERISTICS | No Change |
|  |  | Double Trigger Handpiece | No Change |
|  |  | Type: Double Trigger Handpiece | No Change |
|  |  | Max. drill speed: 1200 rpm or more | No Change |
|  |  | Max. saw speed: 12000 cpm or more | No Change |
|  |  | Max. torque: 3.5 Nm or more | No Change |
|  |  | Mode: Forward/Reverse/Oscillating | No Change |
|  |  | Speed Trigger: Variable | No Change |
|  |  | Toggle Switch: On/Off | No Change |
|  |  | Cannulation: 4.0 mm | No Change |
|  |  | Dimensions (HxWxD): 152 x 127 x 52 mm± 10% | No Change |
|  |  | Weight: Should Not be more than 1250 gm(Without Battery) | No Change |
|  |  | Material: Aluminum/ Stainless Steel | No Change |
|  |  | Compatibility: Battery and AC Power Supply | Compatibility: Battery |
|  |  | Drill Attachments | No Change |
|  |  | Max. Speed: 12000 rpm or more | Max. Speed: 1200 rpm or more |
|  |  | Max. Torque: 3.5 Nm | No Change |
|  |  | Shank Styles: Small AO, Hudson, Trinkle, Zimmer/ Hudson, 1/4in Key Hybrid Chuck, 5/32in Key Hybrid Chuck | No Change |
|  |  | Material: Stainless Steel | No Change |
|  |  | Standard Reamer Attachments | No Change |
|  |  | Max. Speed: 600 rpm | No Change |
|  |  | Max. Torque: 9 Nm or more | No Change |
|  |  | Shank Styles: AO, Small AO, Hudson, Trinkle, Zimmer/ Hudson, Harris, DCS/DHS, 1/4in Key Hybrid Chuck | No Change |
|  |  | Material: Stainless Steel | No Change |
|  |  | High Torque Reamer Attachments | No Change |
|  |  | Max. Speed: 270 rpm | No Change |
|  |  | Max. Torque: 17.5 Nm or more | No Change |
|  |  | Shank Styles: AO, Hudson, Trinkle, Harris, Zimmer/Hudson, DCS/DHS, 1/4in Key Hybrid Chuck | No Change |
|  |  | Material: Stainless Steel | No Change |
|  |  | Pin Driver Max. Speed: 1200 rpm | No Change |
|  |  | Max. Torque: 5 Nm | No Change |
|  |  | Clamping Range: 2.0 – 4.0 mm | No Change |
|  |  | Material: Stainless Steel | No Change |
|  |  | Special Feature: No pin diameter adjustment needed | Special Feature: pin diameter adjustment needed |
|  |  | Wire Driver | No Change |
|  |  | Max. Speed: 1200 rpm | No Change |
|  |  | Max. Torque: 5 Nm or more | No Change |
|  |  | Clamping Range: 0.8 – 2.0 mm | No Change |
|  |  | Material: Stainless Steel | No Change |
|  |  | Special Feature: No pin diameter adjustment needed | Special Feature: pin diameter adjustment needed |
|  |  | Sagittal Saw Attachment | No Change |
|  |  | Max. Speed: 12000 cpm | No Change |
|  |  | Arc of Deflection: 4.5° (+/- 2.25°) | No Change |
|  |  | Material: Stainless Steel | No Change |
|  |  | Special Feature: Single blade can be adjusted +/- 10mm from nominal position within the attachment | Deleted |
|  |  | Reciprocating Saw Attachment | No Change |
|  |  | Max. Speed: 12000 cpm | No Change |
|  |  | Stroke: 3.8 mm± 5% | No Change |
|  |  | Material: Stainless Steel | No Change |
|  |  | Special Feature: Reciprocating Saw attachment can be converted into a Sternum saw by connecting a Sternum Guard (Standard or Large) | No Change |
|  |  | Aseptic and Sterilisable Battery | No Change |
|  |  | Battery Type: Li-Ion, 14.4 VDC, 2000 mAh | No Change |
|  |  | Battery Capacity: 28.8 Wh±10 % | No Change |
|  |  | Charging type: max. 60min (in average: 40 min) | No Change |
|  |  | Special Feature: Equipped with real-time power gauge that indicates remaining battery capacity during the procedure | No Change |
|  |  | Battery Charger | No Change |
|  |  | Charging Bays: 4 or more | No Change |
|  |  | Electrical Requirements: 100–240 V ~ (AC), 50/60 Hz. | No Change |
|  |  | Special Feature: Visual indicator for real-time charging status and advanced battery error detection. | No Change |
|  |  | Sterilization Case | No Change |
|  |  | 1 HP Basket with Housing Slot | No Change |
| Note- All above Equipment/ Instrument should be USFDA/European CE Approved (Sl. No. 1 to 06) |  |  | No Change |
| 06 | External fixator (For District Hospital/SDH/CHC/PHC) |  | No Change |
| 1 | Leg Set- consisting of commonly used tubular rod – 02 Nos., Schanz Screw 4.5mm -06 Nos and Polyaxial Clamps for above screw and rods – 12 Nos. |  | No Change |
| 2 | Femur Set – consisting of commonly used Tubular Rods -02 Nos, Schanz Screw 4.5mm, 5mm -03 Nos. each and Polyaxial Clamps for above Screw and Rods -12 Nos. |  | No Change |
| 3 | Forearm Set: Femur Set – consisting of commonly used solid rods for Forearm Fixator -02 Nos, Schanz Screw 3.5mm, 2.5mm - 03 Nos. each and Polyaxial Clamps for above screws and rods -12 Nos. |  | No Change |
| 4 | USFDA/European CE (Issued by Notified body) /BIS Approved model should be offered. |  | No Change |
| 07 | Pneumatic drill (For District Hospital/SDH/CHC/PHC) |  | No Change |
| 1 | The Cannulated Pneumatic Drill handpiece - Compatible with existing attachments |  | No Change |
| a. | Cannulation with 3.2 mm diameter |  | No Change |
| b. | Operating pressure : 6 - 7 bars (maximum 10 bars) |  | No Change |
| c. | Weight of handpiece 600-800grams without any attachments |  | No Change |
| d. | Power 120 w |  | No Change |
| e. | Variable Speed from 0-900 rpm |  | No Change |
| f. | Noise Level of max 75 db |  | No Change |
| g. | Separate forward and reverse triggers |  | No Change |
| h. | Safety Device to cut off air supply to drill on handpiece |  | No Change |
| i. | Handpiece is compatible with radiolucent drive |  | No Change |
| j. | Instant change between clockwise and counterclockwise rotation |  | No Change |
| k. | Offers reliable protection of soft tissues with oscillating drill attachment |  | No Change |
| l. | Fully Autoclavable |  | No Change |
| m. | Fully machine washable |  | No Change |
| n. | All Attachments can be fitted on single handpiece |  | No Change |
| o. | The reverse trigger automatically locks when the oscillating saw and the reduction drive attachments are attached to handpiece |  | No Change |
| 2 | Adapter for Lubrication |  | No Change |
| a. | For oiling of hand piece &Autoclavable |  | No Change |
| b. | Should be made of Stainless Steel |  | No Change |
| 3 | Double Air Hose |  | No Change |
| a. | Length 5 meters. |  | No Change |
| b. | It should be Autoclavable |  | No Change |
| c. | Should have concentric inlet and outlet pipes |  | No Change |
| 4 | Radiolucent Drive Precise aiming and drilling under image intensifier control for locking intramedullary nails |  | No Change |
|  | Drill Bit diameter 2.0 to 4.5 mm, |  | No Change |
|  | Length 100 to 150 mm, |  | No Change |
|  | Usable length 80 to 120 mm |  | No Change |
|  | Reduced exposure to x-rays |  | No Change |
| 5 | Jacob's Chuck attachment |  | No Change |
|  | Chuck capacity up to 0 to 6.5 mm |  | No Change |
|  | Cannulation of 3.2 mm diameter |  | No Change |
|  | Maximum Speed of 900rpm |  | No Change |
|  | Torque of 4-5 Nm |  | No Change |
| 6 | Quick Coupling attachment |  | No Change |
|  | Cannulation 3.2 mm |  | No Change |
|  | Maximum Speed: 900 rpm |  | No Change |
|  | Torque of 4-5 Nm |  | No Change |
| 7 | Reduction Drive for Intramedullary / Acetabular Reaming with reverse option |  | No Change |
|  | Reaming Speed of 300-350 rpm |  | No Change |
|  | Reaming Torque of 12-14 Nm |  | No Change |
|  | Option of attachment with reverse rotation |  | No Change |
| 8 | Quick Coupling for K-wire |  | No Change |
|  | Continous adjustment facility for wire diameter from 0.6 to 3.2 mm |  | No Change |
|  | Speed up to 900 rpm |  | No Change |
| 9 | Oscillating Saw attachment with key |  | No Change |
|  | It can operate on an oscillating frequency of 0 to 14,000 osc/min. |  | No Change |
|  | Attachment can be locked in 8 different positions |  | No Change |
| 10 | Quick Coupling for drill bits |  | No Change |
|  | Speed: 0–900 rpm |  | No Change |
|  | Torque: 0–4.7 Nm |  | No Change |
|  | Cannulation: 1.3 mm |  | No Change |
| 11 | Quick coupling for DHS / DCS triple reamers |  | No Change |
| 12 | Oscillating Saw Blades (All Sizes) |  | No Change |
|  | For Trauma 6 For Joint replacement |  | No Change |
| 13 | 6 Aluminium case for Pneumatic Drill system, Perforated, Autoclavable 1 |  | No Change |
| 14 | Aluminium Box for accessory attachment, Perforated, Autoclavable 1 Technical Specifications for Pneumatic Drill Set – 1. Company should have relevant experience in successful execution of similar work at least in five Institutes of national importance and central government Institutes |  | No Change |
| 3 | Bidder must enclose original literatures & technical data sheet in the support of the technical bid. |  | No Change |
| 4 | Physical demo may preferably be arranged at the time of requirement. |  | No Change |
| 5 | Instruments quality should meet the international standard. |  | No Change |
| 5 | USFDA/European CE (Issued by Notified body) Approved model should be offered |  | No Change |
|  | Reamer (For District Hospital/SDH/CHC/PHC) |  | No Change |
| 1 | Should be supplied with a driving unit including motor, stand, foot control, flexible shaft, tool kit and container. |  | No Change |
| 2 | Should have a cannulated drill hand piece of maximum speed 1200rpm with Non-corrosive Jacob chuck (0-1/4”). Autoclavable. |  | No Change |
| 3 | Should have a cannulated reaming hand piece with maximum speed 400rpm with AO type quick coupling. Autoclavable. |  | No Change |
| 4 | Should have a pistol grip sagital saw with a set of 5 blades. Autoclavable. |  | No Change |
| 5 | Should have a flexible reamer shaft 8mm diameter fixed head. |  | No Change |
| 6 | Should have a flexible reamer shaft for detachable heads up to 12mm. |  | No Change |
| 7 | Should have reamer heads from 8mm to 12 mm (set of 9). |  | No Change |
| 8 | Should have reamer heads from 12.5mm to 15mm (set of 6). |  | No Change |
| 9 | Should have flexible reamer shaft for detachable heads up to 15mm. Autoclavable. |  | No Change |
| 10 | Should have an extra flexible shaft for driving unit autoclavable. |  | No Change |
| 11 | Should work with input 200 to 240Vac 50 Hz supply. |  | No Change |
| 12 | USFDA/European CE (Issued by Notified body) /BIS Approved model should be offered. |  | No Change |
| 8 | Arthroscopy Pump ( For High End Arthroscopy System-Medical College & Hospital) |  | Deleted |
| a | It should have electronic digital suction irrigation pump with 5’’ or more colour touch LCD/LED screen display |  | Deleted |
| b | It should have Precise presetting of flow and pressure of suction and irrigation parameters via. Touch Screen Display. |  | Deleted |
| C | It should have built in low noise suction pump to allow simultaneous suction – irrigation process. |  | Deleted |
| D | It should have multiple utility mode, like Laparoscopy, Hysteroscopy, Arthroscopy, Urology etc. |  | Deleted |
| E | It should have irrigation pressure should be 10-400 mmHg or more. |  | Deleted |
| F | It should have auto cut off facility for the preset amount of the fluid. |  | Deleted |
| G | It should have preset multiple user setting for different surgeries. |  | Deleted |
| H | It should have flow rate 100-1000 ml/min. |  | Deleted |
| I | It should have display of total volume consumed. |  | Deleted |
| J | It should have suction vacume bar 75Kpa or more. |  | Deleted |
| K | It should have supplied with 2 ltr. Jar 02 nos. |  | Deleted |
| L | It should have supplied with Y shaped irrigation tubing set. |  | Deleted |
| M | It should have auto stop facility. |  | Deleted |
| N | I- Power Supply 100-240 V, II -Frequency : 50 - 60 Hz |  | Deleted |