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Corrigendum-II      
Bihar Medical Services and Infrastructure Corporation Limited (BMSICL) had invited E-Bids from the interested parties for Procurement, Rate contract, Supply, Installation of Medical Equipment vide Tender No.-BMSICL/2025-26/ME-458. During Pre-bid meeting various suggestions were received from different prospective bidders regarding amendment in technical specification of equipment which were discussed and deliberated on by the experts. On the basis of their recommendations certain amendments in the technical specification of the equipment have been made which are annexed as Annexure-I and Annexure-II of this corrigendum. 
Note:-
1. Bidders are advised to refer to the Annexure-I and Annexure-II of this corrigendum before submission of bid. 
2. Those who have submitted their bids are requested to re-submit their bids in accordance with this corrigendum. 

Annexed:- as above

											Sd/-
  GM (Procurement)                         
[bookmark: _GoBack]                                                                                                                       BMSICL












	Annexure-I

	Name of Equipment: - Flow Cytometer

	SI. No
	Technical Specification as per tender
	Final Amendment

	1
	Pre-configured flow cytometer equipped with at least three lasers including blue (488nm) and red (637nm -640nm) and violet (405nm) lasers.
	No Change

	
	a) System should have minimum capability of at least 10 fluorescent colors and parameters.
	No Change

	
	b) System should have a provision of future upgradable for more colors.(optional).
	No Change

	
	c) For each parameter the flow cytometer should be capable of measuring area, height and width simultaneously.
	No Change

	2
	The excitation and collection optics of both lasers should be fixed requiring no alignment to be done by operator.
	No Change

	3
	Should have high quality quartz flow cell.
	No Change

	4
	Should have single tube sample loading mode, integrated and automated multi-tube loader with at least 24 tubes loading capacity/well plate loader.
	No Change

	5
	Should offer low, medium and high flow rates.
	No Change

	6
	Should be able to acquire at least up to 1,000-10,000 events per second or above.
	No Change

	7
	The sample carryover must be ≤0.1%.
	No Change

	8
	Minimum detectable particle size should be 0.5 Micrometer.
	No Change

	9
	Should have compensation capability between all fluorescence channels with online as well as post-acquisition manual and auto-compensation features.
	No Change

	10
	Should have digital signal processing with linear and log modes and dynamic range of at least 5 decades.
	No Change

	11
	Should be operable at 220-230V and 50Hz.
	No Change

	12
	Should be capable of online and offline analysis and capabilities of quality control and validation, cell counting, viability, apoptosis analysis, cytometric bead analysis, cell cycle analysis, surface marker studies and kinetic studies.
	No Change

	13
	The Cytometer should have bio-hazard containment system and proper waste collection and management system.
	No Change

	14
	Compatible computer system (branded only): PC workstation with at least Core i7 or higher, 2 TB hard drive or more, at least 23-inch LCD monitor.
	No Change

	15
	The computer system should have latest licensed windows (Professional) software and Microsoft office.
	No Change

	16
	An additional computer system with above configuration and software for offline analysis should be provided. It should be capable of connecting to the flow cytometer.
	No Change

	17
	One colour printer.
	No Change

	18
	Suitable branded online UPS of sufficient capacity with half hour backup for uninterrupted running of full equipment during power interruption.
	No Change

	19
	Provide with complete essential accessories including appropriate starter kits, QC beads, and maintenance kits.
	No Change

	20
	System should have US FDA/ EUCE (Issued by notified body)/ CE-IVD(Issued by notified body).
	No Change

	21
	Additional Point: The price list of all Mandatory and optional test kit/consumable/accessories/regents/QC shall be attached for future reference. The company shall arrange QC material time to consignee for a period up to 5 years
	No Change





















Annexure-II
Revised Earnest Money Deposit (EMD) in Indian Rupees

	Schedule No.
	Brief Description of Goods and Services
	Qty./ No.
	Delivery Schedule (in days)
	Revised Earnest Money Deposit (EMD) in Indian Rupees

	1
	Flow Cytometer
	28
	As per GCC Clause -8
	50,00,000/-
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