|  | Bihar Medical Services & Infrastructure Corporation Limited, 2nd & 3rd Floor, Swasthya Bhawan, Behind IGIMS, Sheikhpura, Adjacent to State Health Society, Patna 800023, Phone/Fax: +91612 2283287,+ 91612 2283288 |
| --- | --- |

Corrigendum-II

Bihar Medical Services and Infrastructure Corporation Limited (BMSICL) had invited E-Bids from the interested parties for the procurement, rate contract and the supply of medical equipment for different Govt. Institutions of Bihar vide Notice Inviting Re-Tender No.- BMSICL/2025-26/ME-471. During and after Pre-bid meeting various suggestions were received from different prospective bidders regarding amendment in technical specification of equipment which were discussed and deliberated on by the experts, who after due deliberation recommended certain amendments in the technical specification of the equipment, which are annexed as Annexure-I of this corrigendum. In order to facilitate maximum participation of bidders the tender schedule is being revised as follows: -

| Re-Tender Reference No. | BMSICL/2025-26/ME-471 |
| --- | --- |
| Last date and time of submission of online bids | 17th February 2026 till 17:00 Hrs. |
| Last date and time of submission of original documents of EMD, Tender Fee and Document | 18th February 2026 till 14:00 Hrs. |
| Date, Time and Place of opening of Technical Bid | 18th February 2026 (at 15:00 Hrs.) on the website of https:/eproc2.bihar.gov.in in the office of BMSICL |
| Date and time of opening of financial Bids | To be announced later on https:/eproc2.bihar.gov.in |

Note:-

1. Bidders are advised to refer to the Annexure-I of this corrigendum before submission of bid.

2. Those who have submitted their bids are requested to re-submit their bids in accordance with this corrigendum.

Annexed: - as above

GM (Procurement)

BMSICL

| Annexure-I |  |  |  |
| --- | --- | --- | --- |
| Name of Equipment: - Flame Photometer |  |  |  |
| SI. No | Technical Specification as per tender | Final Amendment |  |
| 1 | Microprocessor Control with Filters, Na, K, Li, Ca,-Double display Filters-Narrow band-pass interference filters for sodium, Potassium, Calcium and Lithium available. | No Change |  |
| 2 | Power supply – 220 V ± 10 V 50 Hz | No Change |  |
| 3 | Display- large, Bright Seven Segment LED /LCD Panel meter to display 0-100 scale with negative and over range indication | No Change |  |
| 4 | Compressor- with built in regulator and Air Filter to deliver stable and moisture/oil free air supplier. | No Change |  |
| 5 | Ignition System- Auto/Manual ignition. | No Change |  |
| 6 | Atomiser- Non-corroding, easily accessible and can be cleaned without putting out the flame. | No Change |  |
| 7 | Sample Take up rate- 5ml-6ml/min approximately. | No Change |  |
| 8 | Should have ISO 13485 (NABCB Accredited) | Should have BIS/ISO Certificate |  |
| Name of Equipment: - Chemical Balance |  |  |  |
| SI. No | Technical Specification as per tender | Final Amendment |  |
| 1 | Maximum Capacity – 200 gms or above. | No Change |  |
| 2 | Sensitivity – 10 mg. | No Change |  |
| 3 | Calibration certificates should be submitted for NABH accreditation every year in warranty period | No Change |  |
| 4 | Pan size should be minimum 100mm x 100mm or equivalent. | No Change |  |
| 5 | Stabilization time should be less than 3 seconds. | No Change |  |
|  |  | Should have BIS/ISO Certificate |  |
| Name of Equipment: - Vacuum with Vacuum Cup |  |  |  |
| SI. No | Technical Specification as per tender | Final Amendment |  |
| 1 | 1. Available in manual and electric modes. | No Change |  |
|  | Manual: | No Change |  |
|  | a) Silicon cups (40mm, 50mm and 60 mm without valves). | No Change |  |
|  | b) Autoclavable silicone pressure tube. | No Change |  |
|  | c) Traction handle. | No Change |  |
|  | d) Suction pump. | No Change |  |
|  | e) Fitted Fiber Jar/Polycarbonate Jar (500-600 ml). | No Change |  |
|  | Electric: | No Change |  |
|  | a) Housing ABS plastic molded. | Housing ABS plastic molded/Epoxy Powder Coated Mild Steel (Rust Free) |  |
|  | b) Capacity :- 600 mm hg + 10.45 Lt/min | No Change |  |
|  | c) Pump type: diaphragm with Flat spring Technology. | Pump type: diaphragm with Flat spring Technology/Piston Cylinder Type |  |
|  | d) 1 x 2 Liters Autoclavable polycarbonate jars having Autoclavable Lids with overflow safety. | No Change |  |
|  | e) Filter: Bacterial filter autoclavable / reusable | No Change |  |
|  | f) Quality certificate should be ISO 13485 NABCB Accredited | USFDA/EUCE/ BIS/ISO 13485 NABCB Accredited |  |
| Name of Equipment: - Interval Timer |  |  |  |
| SI. No | Technical Specification as per tender | Final Amendment |  |
| 1 | It should be Programmable. | No Change |  |
| 2 | 1t has maximum count up time & Down time of 23 hours 59 min 59 sec. | No Change |  |
| 3 | Should have large LCD display. | No Change |  |
| 4 | Should have easy setting. | No Change |  |
| 5 | Clock feature with 12/24-hour format. | No Change |  |
| 6 | Manufacturer should have ISO (NABCB accredited) certificate | Should have BIS/ISO Certificate |  |
| Name of Equipment: - Laminar Air -Flow Bench |  |  |  |
| SI. No | Technical Specification as per tender | Final Amendment |  |
| 1 | The Laminar Airflow UV Chamber when switched on, the blower unit should create a suction pressure through the primary filter (or Pre-filter), which removes dust particles of above 10- micron size in the first stage. Subsequently, the filtered air passed to the HEPA filters, where the particles or substances of 0.3 micron size and above are removed. Finally, the ultra-clean filtered air supplied to the working chamber as a uniform airflow to perform precision analysis activities. | No Change |  |
| 2 | The system should have: | No Change |  |
|  | a) Laminar Air Flow Cabinet should have fully enclosed bench design. | No Change |  |
|  | b) The Laminar flow bench should have Stainless Steel SS 304 table with MS coated tabular. | No Change |  |
|  | c) The frame and body should be made of CRCA sheet metal construction with powder coated finish. | No Change |  |
|  | d) Laminar flow unit should be with fixed stand with standard height with levseler and locking wheels. | No Change |  |
| 3 | The unit should have: | No Change |  |
|  | a) Should have LCD display to show measured parameters like exhaust air flow, flow velocity, cabinet temperature, UV/FL elapsed hour timer (Non resettable), UV/FL lamp on/off, HEPA filter life span(Non resettable). | No Change |  |
|  | b) Unit should have Differential pressure indicator. | No Change |  |
| 4 | The system should have CLASS 100 (ISO 5 for particle sizes 0.5µ < 3530 particles/M³ of air at both at Rest & Operation Condition asper ISO 14644 –1. | No Change |  |
| 5 | Working area Minimum 4 ft (W)x 2 ft (D) x 2ft (H). | No Change |  |
| 6 | Work table. | No Change |  |
|  | a) It should have SS 304 grade Stainless Steel with finish 4 polish surface with curved corners and of a single sheet front door. | No Change |  |
|  | b) 5 mm thick clear Acrylic Sheet - Vertical sliding. | No Change |  |
| 7 | Floor standing Base stand Have leveling feet or locking casters or motorized height adjustment. | No Change |  |
| 8 | Direction of flow Horizontal airflow. | No Change |  |
| 9 | Airflow Speed Filter face Velocity should have 100 Feet/Minute ± 20 (0.45 to 0.65 m/s). | No Change |  |
| 10 | Blower Assembly It should have one set blower system which consists of dynamically & statically balanced aluminium centrifugal impeller driven by 1/4 HP, single phase, 1200 to 1400 RPM motor, enclosed in an PU coated GI casing suitably suspended in a pair springs & connected to the filter chamber through flexible canvas duct. The Blower motor should be dynamically balanced, with low noise & vibration. | No Change |  |
| 11 | HEPA Filters | No Change |  |
|  | The filters should have | No Change |  |
|  | a) Size: 48” x 24” x 3” | No Change |  |
|  | b) Type: Separator less type, Mini-Pleats HEPA | No Change |  |
|  | c) Media: Ultra clean glass fiber paper | No Change |  |
|  | d) Initial Pressure: 16 mm WG | No Change |  |
|  | e) Grade: H13 rating | No Change |  |
|  | f) Filter class: H14 according to EN 1822:2009 | No Change |  |
|  | g) Retention: 0.3 Micron Efficiency: 99.997% or better (TEST Report from NABL accredited lab should be submitted) | No Change |  |
| 12 | Pre-Filters | No Change |  |
|  | a) Size: 600 x 300 x 65 mm | No Change |  |
|  | b) Media: Synthetic, non-woven polyester | No Change |  |
|  | c) Casing: Epoxy painted GI frame | No Change |  |
|  | d) Retention: 10 Micron & above | No Change |  |
|  | e) Efficiency: not less than 95% | No Change |  |
|  | f) Initial Pressure: 6 mm WG | No Change |  |
| 13 | Particle Retention 0.3 Micron | No Change |  |
| 14 | Visual & audio alarm For low exhaust flow, exhaust fan malfunctioning. | No Change |  |
| 15 | Noise level < 60 dBA±5 | No Change |  |
| 16 | Power Supply | No Change |  |
|  | a) Power supply should be 220-230 V, 50 Hz. | No Change |  |
|  | b) Should be supplied with suitable voltage corrector/stabilizer. | No Change |  |
| 17 | Illumination Externally mounted illuminating lamp with separate switch to illuminate the work area. | No Change |  |
| 18 | Light | No Change |  |
|  | a) High intensity, low wattage >800 lux. | No Change |  |
|  | b) It should be 15 Watts, 1½ Feet length– 1 No. each | No Change |  |
| 19 | Should have USFDA/ European CE (Issued by Notified Body) /BIS/ISO13485 Approved. | No Change |  |
| 20 | UV lamp pre- mounted germicidal UV lamp (30 W) with separate switch with UV light hours run indicator. For Safety, UV light should be closed when the viewing window is openand when the white light is on. Other accessories | No Change |  |
|  | a) Two gas outlets in the working area, one on each side wall Leveling Screws & Castor Wheels. | No Change |  |
|  | b) PAO (Poly Alpha Olefin) test port Easily changeable pre-filters. | No Change |  |
|  | c) Fitted with UV Germicidal lamp for sterilization. | No Change |  |
|  | d) Pre-installed pressure gauge for Measurement of HEPA Filters Choking system. | No Change |  |
|  | e) Ensure noiseless operation and anti-vibration construction provides efficient working environment. | No Change |  |
|  | f) Audible or highly visual alarm for filter replacement warning. | No Change |  |
| 21 | Electrical sockets or Pass Through Ports a) Side mounted switches for minimum three (15/5 amp) electrical sockets for ancillary equipment operation or b) Convenient rear-wall pass through ports for safe routing of instrument cords, cables and leads for 15/5 amps multiple c) Socket with switches on the wall. | No Change |  |
| Name of Equipment: - pH Meter |  |  |  |
| SI. No | Technical Specification as per tender |  | Final Amendment |
| 1 | Should have up to 5-point calibration with auto-buffer recognition. |  | No Change |
| 2 | Should be Quick, easy electrode diagnosis with multiple pH slopes and offset display. |  | No Change |
| 3 | Should have Non-volatile memory holds up to 500 data points – time and date – stamped for GLP compliance. |  | No Change |
| 4 | Should calibrate with up to 5 custom pH buffers – use any pH values that are >1.0 pH unit apart. |  | No Change |
| 5 | Should have Cal-due alarm – no more out – dated calibrations. |  | No Change |
| 6 | Should have auto-logging function for convenient continuous monitoring. |  | No Change |
| 7 | Should have password protection for setup and calibration. |  | No Change |
| 8 | Should have pH Range up to 14 pH. |  | No Change |
| 9 | pH meter should have resolution of at least 0.01 pH units. |  | No Change |
| 10 | Temp. readings should be accurate within ± 1℃ with a resolution of at least 0.1 ℃. |  | No Change |
| 11 | Should have ATC Probe. |  | No Change |
| 12 | Should have buffer sets each 20ml for all 5-point calibrations. NIST traceable buffer sets shall be provided along with individual certificate. |  | No Change |
| 13 | Should have power requirements of 230v, 50HZ ac supply. |  | No Change |
| 14 | Should have output Rs. 232C (via cable). |  | No Change |
|  | Should have European CE (issued by a notified body)/ USFDA / BIS/ISO 13485 certificate |  | Should have BIS/ISO Certificate |
| Name of Equipment: - Flow Cytometer |  |  |  |
| SI. No | Technical Specification as per tender | Final Amendment |  |
| 1 | 1. Pre-configured flow cytometer equipped with at least three lasers including blue (488nm) and red (637nm -640nm) and violet (405nm) lasers. | No Change |  |
|  | a) System should have minimum capability of at least 10 fluorescent colors and parameters. | No Change |  |
|  | b) System should have a provision of future upgradable for more colors. (optional). | No Change |  |
|  | c) For each parameter the flow cytometer should be capable of measuring area, height and width simultaneously. | For each parameter the flow cytometer should be capable of measuring area, height and width. |  |
| 2 | The excitation and collection optics of both lasers should be fixed requiring no alignment to be done by operator. | No Change |  |
| 3 | Should have high quality quartz flow cell. | No Change |  |
| 4 | Should have single tube sample loading mode, integrated and automated multi-tube loader with at least 24 tubes loading capacity/well plate loader. | No Change |  |
| 5 | Should offer low, medium and high flow rates. | No Change |  |
| 6 | Should be able to acquire at least up to 1,000-10,000 events per second or above. | No Change |  |
| 7 | The sample carryover must be ≤0.1%. | No Change |  |
| 8 | Minimum detectable particle size should be 0.5 Micrometer. | No Change |  |
| 9 | Should have compensation capability between all fluorescence channels with online as well as post-acquisition manual and auto-compensation features. | No Change |  |
| 10 | Should have digital signal processing with linear and log modes and dynamic range of at least 5 decades. | No Change |  |
| 11 | Should be operable at 220-230V and 50Hz. | No Change |  |
| 12 | Should be capable of online and offline analysis and capabilities of quality control and validation, cell counting, viability, apoptosis analysis, cytometric bead analysis, cell cycle analysis, surface marker studies and kinetic studies. | No Change |  |
| 13 | The Cytometer should have bio-hazard containment system and proper waste collection and management system. | No Change |  |
| 14 | Compatible computer system (branded only): PC workstation with at least Core i7 or higher, 2 TB hard drive or more, at least 23-inch LCD monitor. | No Change |  |
| 15 | The computer system should have latest licensed windows (Professional) software and Microsoft office. | No Change |  |
| 16 | An additional computer system with above configuration and software for offline analysis should be provided. It should be capable of connecting to the flow cytometer. | No Change |  |
| 17 | One colour printer. | No Change |  |
| 18 | Suitable branded online UPS of sufficient capacity with half hour backup for uninterrupted running of full equipment during power interruption. | No Change |  |
| 19 | Provide with complete essential accessories including appropriate starter kits, QC beads, and maintenance kits. | No Change |  |
| 20 | System should have US FDA/ EUCE (Issued by notified body)/ CE-IVD(Issued by notified body). | No Change |  |
| 21 | Additional Point: The price list of all Mandatory and optional test kit/consumable/accessories/regents/QC shall be attached for future reference. The company shall arrange QC material time to consignee for a period up to 5 years | No Change |  |